## Philips_EPIQAffiniti_V5.0_DICOM
Philips EPIQ/Affiniti V5.0 DICOM Export Specifications
Template | App | Label | ResultLabel | Type | internalId | internal Result Id | laterality | hashcode | Finding Site | Concept | Result | patientCharacteristics | Units | Target | Mode | View | Phase | Direction | Method | Topographical | anatomy | branch | derivation | laterality2 | respiratory | segment | traceMethod
Adult Echocardiography Procedure Report | Adult Echo | LV Endo Peak GLS | Strain | meas | m_TT_LV_Peak_GLS | r_Strain | unilateral | 125200\G-03A2\GLSendoPeak\Strain_TP\T-32600 | SNM3, T-32600, Left Ventricle | 99IMPROOF, GLSendoPeak, Endo Peak Global Longitudinal Strain |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99IMPROOF, Strain_TP, Strain Triplane
Adult Echocardiography Procedure Report | Adult Echo | LV A3C Endo Peak GLS | Strain | meas | m_TT_A3C_Peak_GLS | r_Strain | unilateral | 125200\C12226-04\G-03A2\GLSendoPeak\Strain_TP\T-32600 | SNM3, T-32600, Left Ventricle | 99IMPROOF, GLSendoPeak, Endo Peak Global Longitudinal Strain |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-04, Apical three chamber |  |  | 99IMPROOF, Strain_TP, Strain Triplane
Adult Echocardiography Procedure Report | Adult Echo | LV A2C Endo Peak GLS | Strain | meas | m_TT_A2C_Peak_GLS | r_Strain | unilateral | 125200\G-03A2\G-A19B\GLSendoPeak\Strain_TP\T-32600 | SNM3, T-32600, Left Ventricle | 99IMPROOF, GLSendoPeak, Endo Peak Global Longitudinal Strain |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99IMPROOF, Strain_TP, Strain Triplane
Adult Echocardiography Procedure Report | Adult Echo | LV A4C Endo Peak GLS | Strain | meas | m_TT_A4C_Peak_GLS | r_Strain | unilateral | 125200\G-03A2\G-A19C\GLSendoPeak\Strain_TP\T-32600 | SNM3, T-32600, Left Ventricle | 99IMPROOF, GLSendoPeak, Endo Peak Global Longitudinal Strain |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99IMPROOF, Strain_TP, Strain Triplane
Adult Echocardiography Procedure Report | Adult Echo | ANNULUS Measurements | MV Ann2DArea (MVA) | meas | m_MVA_ANNULUS_MEAS | r_TT_MVA_ANN2D_AREA | unilateral | 125200\C12207-52\C12228-14\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-52, Mitral Valve 2D Annular Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | COAPTATION | MV A ClLine 2D Length (MVA) | meas | m_MVA_COAPTATION | r_TT_MVA_A_CL_2D | unilateral | 125200\C12207-58\C12228-14\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-58, Mitral Valve Anterior Closure Line Length 2D |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | COAPTATION | MV P ClLine 2D Length (MVA) | meas | m_MVA_COAPTATION | r_TT_MVA_P_CL_2D | unilateral | 125200\C12207-60\C12228-14\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-60, Mitral Valve Posterior Closure Line Length 2D |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | DYNAMIC MVA | MV Annulus Area Fraction (MVA) | meas | m_MVA_DYNAMIC_MVA | r_TT_MVA_ANN_AREA_F | unilateral | 125200\C12207-67\C12228-14\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-67, Mitral Valve Annulus Area Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | LA ED Reservoir L Strain | Strain | meas | m_TT_LA_ED_Res_LS | r_Strain | unilateral | 125200\C12205-16\C12228-15\G-03A2\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-16, Left Atrium ED Reservoir Longitudinal Strain |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-15, AutoStrain LA
Adult Echocardiography Procedure Report | Adult Echo | LA ED Conduit L Strain | Strain | meas | m_TT_LA_ED_Cond_LS | r_Strain | unilateral | 125200\C12205-17\C12228-15\G-03A2\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-17, Left Atrium ED Conduit Longitudinal Strain |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-15, AutoStrain LA
Adult Echocardiography Procedure Report | Adult Echo | LA ED Contraction L Strain | Strain | meas | m_TT_LA_ED_Contr_LS | r_Strain | unilateral | 125200\C12205-18\C12228-15\G-03A2\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-18, Left Atrium ED Contraction Longitudinal Strain |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-15, AutoStrain LA
Adult Echocardiography Procedure Report | Adult Echo | LA PreA Reservoir L Strain | Strain | meas | m_TT_LA_PreA_Res_LS | r_Strain | unilateral | 125200\C12205-19\C12228-15\G-03A2\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-19, Left Atrium Pre-A Reservoir Longitudinal Strain |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-15, AutoStrain LA
Adult Echocardiography Procedure Report | Adult Echo | LA PreA Conduit L Strain | Strain | meas | m_TT_LA_PreA_Cond_LS | r_Strain | unilateral | 125200\C12205-20\C12228-15\G-03A2\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-20, Left Atrium Pre-A Conduit Longitudinal Strain |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-15, AutoStrain LA
Adult Echocardiography Procedure Report | Adult Echo | LA PreA Contraction L Strain | Strain | meas | m_TT_LA_PreA_Cont_LS | r_Strain | unilateral | 125200\C12205-21\C12228-15\G-03A2\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-21, Left Atrium Pre-A Contraction Longitudinal Strain |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-15, AutoStrain LA
Adult Echocardiography Procedure Report | Adult Echo | 2D | RV FAC | meas | m_RVHM_2D | r_TT_RV_FAC | unilateral | 125200\C12204-03\C12228-13\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-03, Right Ventricular Fractional Area Change |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-13, 3D AUTO RV
Adult Echocardiography Procedure Report | Adult Echo | 2D | RVD base ED | meas | m_RVHM_2D | r_TT_RV_RVDD_BASE_ED | unilateral | 125200\C12204-20\C12228-13\G-03A2\R-FAB5C\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-20, Right Ventricular Basal Diameter at End Diastole |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5C, End Diastole |  | 99PMSBLUS, C12228-13, 3D AUTO RV
Adult Echocardiography Procedure Report | Adult Echo | 2D | RVD mid cavity ED | meas | m_RVHM_2D | r_TT_RV_RVDD_MID_ED | unilateral | 125200\C12204-21\C12228-13\G-03A2\R-FAB5C\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-21, Right Ventricular Mid Cavity Diameter at End Diastole |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5C, End Diastole |  | 99PMSBLUS, C12228-13, 3D AUTO RV
Adult Echocardiography Procedure Report | Adult Echo | 2D | RV length ED | meas | m_RVHM_2D | r_TT_RV_RVLD_ED | unilateral | 125200\C12204-22\C12228-13\G-03A2\R-FAB5C\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-22, Right Ventricular Base to Apex Length at End Diastole |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5C, End Diastole |  | 99PMSBLUS, C12228-13, 3D AUTO RV
Adult Echocardiography Procedure Report | Adult Echo | 2D | RVLS Septum | meas | m_RVHM_2D | r_TT_RV_LS_SEPTUM | unilateral | 125200\C12204-23\C12228-13\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-23, Right Ventricular Septum Longitudinal Strain |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-13, 3D AUTO RV
Adult Echocardiography Procedure Report | Adult Echo | 2D | RVLS Freewall | meas | m_RVHM_2D | r_TT_RV_LS_FREEWALL | unilateral | 125200\C12204-24\C12228-13\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-24, Right Ventricular Freewall Longitudinal Strain |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-13, 3D AUTO RV
Adult Echocardiography Procedure Report | Adult Echo | RV Free wall L Strain | Strain | meas | m_TT_RV_Free_Wall_LS | r_Strain | unilateral | 125200\C12204-27\C12228-16\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-27, Right Ventricle Free wall Longitudinal Strain |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-16, AutoStrain RV
Adult Echocardiography Procedure Report | Adult Echo | RV L Strain | Strain | meas | m_TT_RV_4C_LS | r_Strain | unilateral | 125200\C12204-28\C12228-16\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-28, Right Ventricle Longitudinal Strain |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-16, AutoStrain RV
Adult Echocardiography Procedure Report | Adult Echo | 2D | TAPSE | meas | m_RVHM_2D | r_TT_RV_TAPSE | unilateral | 125200\C12208-11\C12228-13\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12208-11, Tricuspid Annular Plane Systolic Excursion |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-13, 3D AUTO RV
Adult Echocardiography Procedure Report | Adult Echo | ANNULUS Measurements | MV AP Diam (MVA) | meas | m_MVA_ANNULUS_MEAS | r_TT_MVA_AP_Diam | unilateral | 125200\C12207-07\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-07, Mitral Valve Anterior Posterior Diameter |  |  | UCUM, mm, mm |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | ANNULUS Measurements | MV AL-PM Diam (MVA) | meas | m_MVA_ANNULUS_MEAS | r_TT_MVA_AL_PM_DIAM | unilateral | 125200\C12207-08\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-08, Mitral Valve Anterolateral Posteromedial Diameter |  |  | UCUM, mm, mm |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | ANNULUS Measurements | MV AnnulusCirc (MVA) | meas | m_MVA_ANNULUS_MEAS | r_TT_MVA_ANN_CIRC | unilateral | 125200\C12207-09\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-09, Mitral Valve Annulus 3D Circumference |  |  | UCUM, mm, mm |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | LEAFLETS | MV A Leaf 3D Area (MVA) | meas | m_MVA_LEAFLETS | r_TT_MVA_A_LEAF3D_AR | unilateral | 125200\C12207-10\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-10, Mitral Valve Anterior Leaflet Area |  |  | UCUM, mm2, mm2 |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | LEAFLETS | MV P Leaf 3D Area (MVA) | meas | m_MVA_LEAFLETS | r_TT_MVA_P_LEAF3D_AR | unilateral | 125200\C12207-13\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-13, Mitral Valve Posterior Leaflets 3D Area |  |  | UCUM, mm2, mm2 |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | LEAFLETS | MV TentVol (MVA) | meas | m_MVA_LEAFLETS | r_TT_MVA_TENT_VOL | unilateral | 125200\C12207-14\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-14, Mitral Valve Tenting Volume |  |  | UCUM, mm3, mm3 |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | LEAFLETS | MV Dist A Leaf Angle (MVA) | meas | m_MVA_LEAFLETS | r_TT_MVA_DA_LEAF_ANG | unilateral | 125200\C12207-18\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-18, Mitral Valve Anterior Leaflet Angle |  |  | UCUM, deg, deg |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | LEAFLETS | MV Post Leaf Angle (MVA) | meas | m_MVA_LEAFLETS | r_TT_MVA_P_LEAF_ANG | unilateral | 125200\C12207-19\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-19, Mitral Valve Posterior Leaflet Angle |  |  | UCUM, deg, deg |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | ANNULUS Measurements | MV Nonplanar Angle (MVA) | meas | m_MVA_ANNULUS_MEAS | r_TT_MVA_NONPLAN_ANG | unilateral | 125200\C12207-20\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-20, Mitral Valve Nonplanar Angle |  |  | UCUM, deg, deg |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | MISC | MV Ao OR-Mitr Ang (MVA) | meas | m_MVA_MISC | r_TT_MVA_AO_OR_M_ANG | unilateral | 125200\C12207-21\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-21, Mitral Valve Aortic Orifice to Mitral Plane Angle |  |  | UCUM, deg, deg |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | LEAFLETS | MV TentHeight (MVA) | meas | m_MVA_LEAFLETS | r_TT_MVA_TENT_HEIGHT | unilateral | 125200\C12207-23\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-23, Mitral Valve Tenting Height |  |  | UCUM, mm, mm |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | ANNULUS Measurements | MV Ann3DArea (MVA) | meas | m_MVA_ANNULUS_MEAS | r_TT_MVA_ANN3D_AREA | unilateral | 125200\C12207-26\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-26, Mitral Valve 3D Annular Area |  |  | UCUM, mm2, mm2 |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | ANNULUS Calculation | MV Sphericity Index (MVA) | meas | m_MVA_ANNULUS_CALC | r_TT_MVA_SPHER_IDX | unilateral | 125200\C12207-51\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-51, Mitral Valve Spherical Index |  |  | UCUM, 1, no units |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | ANNULUS Measurements | MV Ann Height (MVA) | meas | m_MVA_ANNULUS_MEAS | r_TT_MVA_ANN_HEIGHT | unilateral | 125200\C12207-53\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-53, Mitral Valve Annular Height |  |  | UCUM, mm, mm |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | LEAFLETS | MV TentArea (MVA) | meas | m_MVA_LEAFLETS | r_TT_MVA_TENT_AREA | unilateral | 125200\C12207-54\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-54, Mitral Valve Tent Area |  |  | UCUM, mm2, mm2 |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | LEAFLETS | MV Comm Diam (MVA) | meas | m_MVA_LEAFLETS | r_TT_MVA_COMM_DIAM | unilateral | 125200\C12207-55\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-55, Mitral Valve Commissural Diameter |  |  | UCUM, mm, mm |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | LEAFLETS | MV A Leaf Length (MVA) | meas | m_MVA_LEAFLETS | r_TT_MVA_A_LEAF_LEN | unilateral | 125200\C12207-56\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-56, Mitral Valve Anterior Leaflet Length |  |  | UCUM, mm, mm |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | LEAFLETS | MV P Leaf Length (MVA) | meas | m_MVA_LEAFLETS | r_TT_MVA_P_LEAF_LEN | unilateral | 125200\C12207-57\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-57, Mitral Valve Posterior Leaflet Length |  |  | UCUM, mm, mm |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | COAPTATION | MV A ClLine 3D Length (MVA) | meas | m_MVA_COAPTATION | r_TT_MVA_A_CL_3D | unilateral | 125200\C12207-59\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-59, Mitral Valve Anterior Closure Line Length 3D |  |  | UCUM, mm, mm |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | COAPTATION | MV P ClLine 3D Length (MVA) | meas | m_MVA_COAPTATION | r_TT_MVA_P_CL_3D | unilateral | 125200\C12207-61\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-61, Mitral Valve Posterior Closure Line Length 3D |  |  | UCUM, mm, mm |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | MISC | MV Intertrigonal 3D Dist (MVA) | meas | m_MVA_MISC | r_TT_MVA_IT_3D_DIST | unilateral | 125200\C12207-62\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-62, Mitral Valve Intertrigonal Distance |  |  | UCUM, mm, mm |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | MISC | MV CShaped 3D Ann (MVA) | meas | m_MVA_MISC | r_TT_MVA_CS_3D_ANN | unilateral | 125200\C12207-63\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-63, Mitral Valve C_shaped Annulus |  |  | UCUM, mm, mm |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | DYNAMIC MVA | MV Max Ann Displacement (MVA) | meas | m_MVA_DYNAMIC_MVA | r_TT_MVA_MAX_A_DISP | unilateral | 125200\C12207-64\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-64, Mitral Valve Max Annulus Displacement |  |  | UCUM, mm, mm |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | DYNAMIC MVA | MV Max Ann Displacement Vel (MVA) | meas | m_MVA_DYNAMIC_MVA | r_TT_MVA_MAX_A_DVEL | unilateral | 125200\C12207-65\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-65, Mitral Valve Annulus Max Displacement Velocity |  |  | UCUM, mm/s, mm/s |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | DYNAMIC MVA | MV Tenting Volume Fraction (MVA) | meas | m_MVA_DYNAMIC_MVA | r_TT_MVA_TENT_VOL_F | unilateral | 125200\C12207-66\C12228-14\P0-02242\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-66, Mitral Valve Tenting Volume Fraction |  |  | UCUM, %, Percent |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-14, MVA
Adult Echocardiography Procedure Report | Adult Echo | Calculation(s) | RV EDVi | meas | m_RVHM_Calcs | r_TT_RV_EDVI | unilateral | 125200\C12204-18\C12228-13\P0-02242\R-FAB5C\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-18, Right Ventricular End Diastolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, P0-02242, 3D mode |  | SRT, R-FAB5C, End Diastole |  | 99PMSBLUS, C12228-13, 3D AUTO RV
Adult Echocardiography Procedure Report | Adult Echo | Calculation(s) | RV ESVi | meas | m_RVHM_Calcs | r_TT_RV_ESVI | unilateral | 125200\109070\C12204-19\C12228-13\P0-02242\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-19, Right Ventricular End Systolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, P0-02242, 3D mode |  | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-13, 3D AUTO RV
Adult Echocardiography Procedure Report | Adult Echo | Calculation(s) | RV EF | meas | m_RVHM_Calcs | r_TT_RV_EF | unilateral | 125200\C12204-25\C12228-13\P0-02242\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-25, Right Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-13, 3D AUTO RV
Adult Echocardiography Procedure Report | Adult Echo | Calculation(s) | RV SV | meas | m_RVHM_Calcs | r_TT_RV_SV | unilateral | 125200\C12204-26\C12228-13\P0-02242\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-26, Right Ventricular Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-13, 3D AUTO RV
Adult Echocardiography Procedure Report | Adult Echo | Volume(s) | RV EDV | meas | m_RVHM_Volume | r_TT_RV_EDV | unilateral | 125200\8822-3\C12228-13\P0-02242\R-FAB5C\T-32500 | SRT, T-32500, Right Ventricle | LN, 8822-3, Right Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, P0-02242, 3D mode |  | SRT, R-FAB5C, End Diastole |  | 99PMSBLUS, C12228-13, 3D AUTO RV
Adult Echocardiography Procedure Report | Adult Echo | Volume(s) | RV ESV | meas | m_RVHM_Volume | r_TT_RV_ESV | unilateral | 125200\109070\8824-5\C12228-13\P0-02242\T-32500 | SRT, T-32500, Right Ventricle | LN, 8824-5, Right Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, P0-02242, 3D mode |  | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-13, 3D AUTO RV
Adult Echocardiography Procedure Report | Adult Echo | MV Vmax-pr | Vmax | meas | m_MV_Vmax_pr | r_Vmax | unilateral | 125200\11726-7\C12207-46\R-42047 | 99PMSBLUS, C12207-46, Mitral Valve (Prosthetic) | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Vmax-pr | Max PG | meas | m_MV_Vmax_pr | r_MaxPG | unilateral | 125200\20247-3\C12207-46\R-42047 | 99PMSBLUS, C12207-46, Mitral Valve (Prosthetic) | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV VTI-pr | Mean PG | meas | m_MV_VTI_pr | r_CardPGmean | unilateral | 125200\20256-4\C12207-46\R-42047 | 99PMSBLUS, C12207-46, Mitral Valve (Prosthetic) | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV P1/2t-pr | P½t | meas | m_MV_PHalfT_pr | r_PHalftime | unilateral | 125200\20280-4\C12207-46 | 99PMSBLUS, C12207-46, Mitral Valve (Prosthetic) | LN, 20280-4, Pressure Half-Time |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | MV VTI-pr | Vmean | meas | m_MV_VTI_pr | r_CardTAPV | unilateral | 125200\20352-1\C12207-46\R-42047 | 99PMSBLUS, C12207-46, Mitral Valve (Prosthetic) | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV VTI-pr | VTI | meas | m_MV_VTI_pr | r_VTI | unilateral | 125200\20354-7\C12207-46\R-42047 | 99PMSBLUS, C12207-46, Mitral Valve (Prosthetic) | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AV Vmax-pr | Vmax | meas | m_AV_Vmax_pr | r_Vmax | unilateral | 125200\11726-7\C12211-15\R-42047 | 99PMSBLUS, C12211-15, Aortic Valve (Prosthetic) | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AV Vmax-pr | Max PG | meas | m_AV_Vmax_pr | r_MaxPG | unilateral | 125200\20247-3\C12211-15\R-42047 | 99PMSBLUS, C12211-15, Aortic Valve (Prosthetic) | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AV VTI-pr | Mean PG | meas | m_AV_VTI_pr | r_CardPGmean | unilateral | 125200\20256-4\C12211-15\R-42047 | 99PMSBLUS, C12211-15, Aortic Valve (Prosthetic) | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AV P1/2t-pr | P½t | meas | m_AV_PHalfT_pr | r_PHalftime | unilateral | 125200\20280-4\C12211-15 | 99PMSBLUS, C12211-15, Aortic Valve (Prosthetic) | LN, 20280-4, Pressure Half-Time |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | AV VTI-pr | Vmean | meas | m_AV_VTI_pr | r_CardTAPV | unilateral | 125200\20352-1\C12211-15\R-42047 | 99PMSBLUS, C12211-15, Aortic Valve (Prosthetic) | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AV VTI-pr | VTI | meas | m_AV_VTI_pr | r_VTI | unilateral | 125200\20354-7\C12211-15\R-42047 | 99PMSBLUS, C12211-15, Aortic Valve (Prosthetic) | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 1 | meas | m_A2Cs | r_Seg1 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 10 | meas | m_A2Cs | r_Seg10 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 11 | meas | m_A2Cs | r_Seg11 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 12 | meas | m_A2Cs | r_Seg12 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 13 | meas | m_A2Cs | r_Seg13 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 14 | meas | m_A2Cs | r_Seg14 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 15 | meas | m_A2Cs | r_Seg15 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 16 | meas | m_A2Cs | r_Seg16 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 17 | meas | m_A2Cs | r_Seg17 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 18 | meas | m_A2Cs | r_Seg18 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 19 | meas | m_A2Cs | r_Seg19 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 2 | meas | m_A2Cs | r_Seg2 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 20 | meas | m_A2Cs | r_Seg20 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 3 | meas | m_A2Cs | r_Seg3 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 4 | meas | m_A2Cs | r_Seg4 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 5 | meas | m_A2Cs | r_Seg5 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 6 | meas | m_A2Cs | r_Seg6 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 7 | meas | m_A2Cs | r_Seg7 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 8 | meas | m_A2Cs | r_Seg8 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | Seg 9 | meas | m_A2Cs | r_Seg9 | unilateral | 125200\109070\C12201-01\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 1 | meas | m_A2Cd | r_Seg1 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 10 | meas | m_A2Cd | r_Seg10 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 11 | meas | m_A2Cd | r_Seg11 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 12 | meas | m_A2Cd | r_Seg12 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 13 | meas | m_A2Cd | r_Seg13 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 14 | meas | m_A2Cd | r_Seg14 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 15 | meas | m_A2Cd | r_Seg15 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 16 | meas | m_A2Cd | r_Seg16 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 17 | meas | m_A2Cd | r_Seg17 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 18 | meas | m_A2Cd | r_Seg18 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 19 | meas | m_A2Cd | r_Seg19 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 2 | meas | m_A2Cd | r_Seg2 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 20 | meas | m_A2Cd | r_Seg20 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 3 | meas | m_A2Cd | r_Seg3 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 4 | meas | m_A2Cd | r_Seg4 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 5 | meas | m_A2Cd | r_Seg5 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 6 | meas | m_A2Cd | r_Seg6 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 7 | meas | m_A2Cd | r_Seg7 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 8 | meas | m_A2Cd | r_Seg8 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | Seg 9 | meas | m_A2Cd | r_Seg9 | unilateral | 125200\C12201-01\F-32011\G-A19B\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 1 | meas | m_A4Cs | r_Seg1 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 10 | meas | m_A4Cs | r_Seg10 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 11 | meas | m_A4Cs | r_Seg11 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 12 | meas | m_A4Cs | r_Seg12 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 13 | meas | m_A4Cs | r_Seg13 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 14 | meas | m_A4Cs | r_Seg14 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 15 | meas | m_A4Cs | r_Seg15 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 16 | meas | m_A4Cs | r_Seg16 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 17 | meas | m_A4Cs | r_Seg17 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 18 | meas | m_A4Cs | r_Seg18 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 19 | meas | m_A4Cs | r_Seg19 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 2 | meas | m_A4Cs | r_Seg2 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 20 | meas | m_A4Cs | r_Seg20 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 3 | meas | m_A4Cs | r_Seg3 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 4 | meas | m_A4Cs | r_Seg4 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 5 | meas | m_A4Cs | r_Seg5 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 6 | meas | m_A4Cs | r_Seg6 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 7 | meas | m_A4Cs | r_Seg7 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 8 | meas | m_A4Cs | r_Seg8 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | Seg 9 | meas | m_A4Cs | r_Seg9 | unilateral | 125200\109070\C12201-01\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 1 | meas | m_A4Cd | r_Seg1 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 10 | meas | m_A4Cd | r_Seg10 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 11 | meas | m_A4Cd | r_Seg11 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 12 | meas | m_A4Cd | r_Seg12 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 13 | meas | m_A4Cd | r_Seg13 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 14 | meas | m_A4Cd | r_Seg14 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 15 | meas | m_A4Cd | r_Seg15 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 16 | meas | m_A4Cd | r_Seg16 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 17 | meas | m_A4Cd | r_Seg17 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 18 | meas | m_A4Cd | r_Seg18 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 19 | meas | m_A4Cd | r_Seg19 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 2 | meas | m_A4Cd | r_Seg2 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 20 | meas | m_A4Cd | r_Seg20 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 3 | meas | m_A4Cd | r_Seg3 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 4 | meas | m_A4Cd | r_Seg4 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 5 | meas | m_A4Cd | r_Seg5 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 6 | meas | m_A4Cd | r_Seg6 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 7 | meas | m_A4Cd | r_Seg7 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 8 | meas | m_A4Cd | r_Seg8 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | Seg 9 | meas | m_A4Cd | r_Seg9 | unilateral | 125200\C12201-01\F-32011\G-A19C\T-32600\T5203-01 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-01, Left Ventricle MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | LV Length ES (HM) | meas | m_HM_LV_EF | r_HM_LVBAl_ES | unilateral | 125200\109070\C12201-03\C12228-11\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-03, LV Base to Apex Length |  |  | UCUM, mm, mm |  | 99PMSBLUS, T12224-01, 3D mode |  | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | LV Length ED (HM) | meas | m_HM_LV_EF | r_HM_LVBAl_ED | unilateral | 125200\C12201-03\C12228-11\F-32011\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12201-03, LV Base to Apex Length |  |  | UCUM, mm, mm |  | 99PMSBLUS, T12224-01, 3D mode |  | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 16-SD (ms) | Time | meas | m_Tmsv_16_SD | r_Time | unilateral | 125200\C12203-10\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-10, SD of Time to min sys vol for 16 wall segments |  |  | UCUM, ms, ms |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | avg EDV (HM) | meas | m_HM_LV_EF | r_HM_EDV_LV_AVG | unilateral | 125200\C12203-107\C12228-11\F-32011\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-107, Left Ventricular MultiBeat End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | 99PMSBLUS, T12224-01, 3D mode |  | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | avg ESV (HM) | meas | m_HM_LV_EF | r_HM_ESV_LV_AVG | unilateral | 125200\109070\C12203-108\C12228-11\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-108, Left Ventricular MultiBeat End Systolic Volume |  |  | UCUM, mm3, mm3 |  | 99PMSBLUS, T12224-01, 3D mode |  | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | avg LV Length ES (HM) | meas | m_HM_LV_EF | r_HM_LVBAl_ES_AVG | unilateral | 125200\109070\C12203-109\C12228-11\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-109, Left Ventricular MultiBeat Base to Apex Length |  |  | UCUM, mm, mm |  | 99PMSBLUS, T12224-01, 3D mode |  | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | avg LV Length ED (HM) | meas | m_HM_LV_EF | r_HM_LVBAl_ED_AVG | unilateral | 125200\C12203-109\C12228-11\F-32011\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-109, Left Ventricular MultiBeat Base to Apex Length |  |  | UCUM, mm, mm |  | 99PMSBLUS, T12224-01, 3D mode |  | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 12-SD (ms) | Time | meas | m_Tmsv_12_SD | r_Time | unilateral | 125200\C12203-11\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-11, SD of Time to min sys vol for 6 basal N 6 Mid segments |  |  | UCUM, ms, ms |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | avg SV (HM) | meas | m_HM_LV_EF | r_HM_SV_AVG | unilateral | 125200\C12203-110\C12228-11\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-110, MultiBeat Stroke Volume |  |  | UCUM, mm3, mm3 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | avg CI (HM) | meas | m_HM_LV_EF | r_HM_CI_AVG | unilateral | 125200\C12203-111\C12228-11\G-A19C\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-111, MultiBeat Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | 99PMSBLUS, T12224-01, 3D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | avg EF (HM) | meas | m_HM_LV_EF | r_HM_EF_AVG | unilateral | 125200\C12203-112\C12228-11\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-112, Left Ventricular MultiBeat Ejection Fraction |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | Other (HM) | avg HR (HM) | meas | m_HM_Other | r_HeartRate_AVG | unilateral | 125200\C12203-113\C12228-11\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-113, MultiBeat Heart rate |  |  | UCUM, {H.B.}/min, bpm |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | avg LV Mass (HM) | meas | m_HM_LV_EF | r_HM_MassED_AVG | unilateral | 125200\C12203-114\C12228-11\F-32011\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-114, Left Ventricular MultiBeat Mass End Diastole |  |  | UCUM, g, grams |  | 99PMSBLUS, T12224-01, 3D mode |  | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | avg ES Current (HM) | meas | m_HM_LV_EF | r_HM_ES_Border_Setting_Avg | unilateral | 125200\109070\C12203-115\C12228-11\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-115, LV MultiBeat Border Setting |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | avg ED Current (HM) | meas | m_HM_LV_EF | r_HM_ED_Border_Setting_Avg | unilateral | 125200\C12203-115\C12228-11\F-32011\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-115, LV MultiBeat Border Setting |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 6-SD (ms) | Time | meas | m_Tmsv_6_SD | r_Time | unilateral | 125200\C12203-12\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-12, SD of Time to min sys vol for 6 basal segments |  |  | UCUM, ms, ms |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 16-DIF (ms) | Time | meas | m_Tmsv_16_MxDiff | r_Time | unilateral | 125200\C12203-13\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-13, Max Dif of Time to min sys vol for 16 wall segments |  |  | UCUM, ms, ms |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 12-DIF (ms) | Time | meas | m_Tmsv_12_MxDiff | r_Time | unilateral | 125200\C12203-14\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-14, Max Dif of Time to min sys vol for 6 basal N 6 Mid segments |  |  | UCUM, ms, ms |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 6-DIF (ms) | Time | meas | m_Tmsv_6_MxDiff | r_Time | unilateral | 125200\C12203-15\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-15, Max Dif of Time to min sys vol for 6 basal segments |  |  | UCUM, ms, ms |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 3-5 (ms) | Time | meas | m_Tmsv_SP_MxDiff | r_Time | unilateral | 125200\C12203-16\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-16, Dif of Time to min sys vol btn basal sept N post segments |  |  | UCUM, ms, ms |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 3-6 (ms) | Time | meas | m_Tmsv_SL_MxDiff | r_Time | unilateral | 125200\C12203-17\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-17, Dif of Time to min sys vol btn basal sept N lat segments |  |  | UCUM, ms, ms |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 16-SD (%) | Percent | meas | m_nTmsv_16_SD | r_Percent | unilateral | 125200\C12203-18\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-18, Nrm. SD of Time to min sys vol for 16 wall segments |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 12-SD (%) | Percent | meas | m_nTmsv_12_SD | r_Percent | unilateral | 125200\C12203-19\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-19, Nrm. SD of Time to min sys vol for 6 basal N 6 Mid segments |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 6-SD (%) | Percent | meas | m_nTmsv_6_SD | r_Percent | unilateral | 125200\C12203-20\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-20, Nrm. SD of Time to min sys vol for 6 basal segments |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 16-DIF (%) | Percent | meas | m_nTmsv_16_MxDiff | r_Percent | unilateral | 125200\C12203-21\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-21, Nrm. Mx Dif of Time to min sys vol for 16 wall segments |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 12-DIF (%) | Percent | meas | m_nTmsv_12_MxDiff | r_Percent | unilateral | 125200\C12203-22\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-22, Nrm. Mx Dif of Time to min sys vol for 6 basal N 6 Mid segments |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 6-DIF (%) | Percent | meas | m_nTmsv_6_MxDiff | r_Percent | unilateral | 125200\C12203-23\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-23, Nrm. Mx Dif of Time to min sys vol for 6 basal segments |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 3-5 (%) | Percent | meas | m_nTmsv_SP_MxDiff | r_Percent | unilateral | 125200\C12203-24\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-24, Nrm. Diff of Time to min sys vol btn basal sept N post segments |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 3-6 (%) | Percent | meas | m_nTmsv_SL_MxDiff | r_Percent | unilateral | 125200\C12203-25\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-25, Nrm. Diff of Time to min sys vol btn basal sept N lat segments |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | R-R time | Time | meas | m_3DQA_RRTime | r_Time | unilateral | 125200\C12203-26\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-26, Time between R-R waves |  |  | UCUM, ms, ms |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | LV Excursion - Avg | Dist | meas | m_3DQA_LVExcMean | r_Distance | unilateral | 125200\C12203-27\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-27, The Avg Inward Motion of the Myocardium Across the LV Wall |  |  | UCUM, mm, mm |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | LV Excursion - SD | Dist | meas | m_3DQA_LVExcSD | r_Distance | unilateral | 125200\C12203-28\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-28, The SD Inward Motion of the Myocardium Across the LV Wall |  |  | UCUM, mm, mm |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | LV Excursion - Min | Dist | meas | m_3DQA_LVExcMin | r_Distance | unilateral | 125200\C12203-29\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-29, The Min Inward Motion of the Myocardium Across the LV Wall |  |  | UCUM, mm, mm |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | LV Excursion - Max | Dist | meas | m_3DQA_LVExcMax | r_Distance | unilateral | 125200\C12203-30\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-30, The Max Inward Motion of the Myocardium Across the LV Wall |  |  | UCUM, mm, mm |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | EF BP-3DQ | LV Mass ED | meas | m_3DQ_EF_BP | r_3DQLVMassED | unilateral | 125200\C12203-35\C12228-06\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-35, Left Ventricular Mass End Diastole |  |  | UCUM, g, grams |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-06, QLab 3DQ
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | LV Mass (HM) | meas | m_HM_LV_EF | r_HM_MassED | unilateral | 125200\C12203-35\C12228-11\F-32011\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-35, Left Ventricular Mass End Diastole |  |  | UCUM, g, grams |  | 99PMSBLUS, T12224-01, 3D mode |  | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | EF BP-3DQ | LV Mass ES | meas | m_3DQ_EF_BP | r_3DQLVMassES | unilateral | 125200\C12203-36\C12228-06\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-36, Left Ventricular Mass End Systole |  |  | UCUM, g, grams |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-06, QLab 3DQ
Adult Echocardiography Procedure Report | Adult Echo | LV Lat Wall | Dist | meas | m_LV_Lat_Wall | r_Distance | unilateral | 125200\C12203-67\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-67, Left Ventricle Lateral Wall Thickness |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LV Ant Wall | Dist | meas | m_LV_Ant_Wall | r_Distance | unilateral | 125200\C12203-68\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-68, Left Ventricle Anterior Wall Thickness |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LV Inf Wall | Dist | meas | m_LV_Inf_Wall | r_Distance | unilateral | 125200\C12203-69\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-69, Left Ventricle Inferior Wall Thickness |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Inf S Vel | Vel | meas | m_Inf_S_Vel | r_Vel | unilateral | 125200\C12203-71\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-71, Left Ventricle Inferior S Velocity |  |  | UCUM, mm/s, mm/s |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Inf E' Vel | Vel | meas | m_Inf_E_Vel | r_Vel | unilateral | 125200\C12203-72\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-72, Left Ventricle Inferior E Velocity |  |  | UCUM, mm/s, mm/s |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Inf A' Vel | Vel | meas | m_Inf_A_Vel | r_Vel | unilateral | 125200\C12203-73\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-73, Left Ventricle Inferior A Velocity |  |  | UCUM, mm/s, mm/s |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Time to Inf S | Time | meas | m_Time_to_Inf_S | r_DopTime | unilateral | 125200\C12203-74\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-74, Left Ventricle Time To Inferior S |  |  | UCUM, ms, ms |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Ant S Vel | Vel | meas | m_Ant_S_Vel | r_Vel | unilateral | 125200\C12203-75\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-75, Left Ventricle Anterior S Velocity |  |  | UCUM, mm/s, mm/s |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Ant E€™ Vel | Vel | meas | m_Ant_E_Vel | r_Vel | unilateral | 125200\C12203-76\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-76, Left Ventricle Anterior E Velocity |  |  | UCUM, mm/s, mm/s |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Ant A€™ Vel | Vel | meas | m_Ant_A_Vel | r_Vel | unilateral | 125200\C12203-77\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-77, Left Ventricle Anterior A Velocity |  |  | UCUM, mm/s, mm/s |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Time to Ant S | Time | meas | m_Time_to_Ant_S | r_DopTime | unilateral | 125200\C12203-78\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-78, Left Ventricle Time To Anterior S |  |  | UCUM, ms, ms |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Post S Vel | Vel | meas | m_Post_S_Vel | r_Vel | unilateral | 125200\C12203-79\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-79, Left Ventricle Posterior S Velocity |  |  | UCUM, mm/s, mm/s |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Post E€™ Vel | Vel | meas | m_Post_E_Vel | r_Vel | unilateral | 125200\C12203-80\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-80, Left Ventricle Posterior E Velocity |  |  | UCUM, mm/s, mm/s |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Post A€™ Vel | Vel | meas | m_Post_A_Vel | r_Vel | unilateral | 125200\C12203-81\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-81, Left Ventricle Posterior A Velocity |  |  | UCUM, mm/s, mm/s |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Time to Post S | Time | meas | m_Time_to_Post_S | r_DopTime | unilateral | 125200\C12203-82\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-82, Left Ventricle Time To Posterior S |  |  | UCUM, ms, ms |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | AS S Vel | Vel | meas | m_AS_S_Vel | r_Vel | unilateral | 125200\C12203-83\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-83, Left Ventricle Anteroseptal S velocity |  |  | UCUM, mm/s, mm/s |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | AS E€™ Vel | Vel | meas | m_AS_E_Vel | r_Vel | unilateral | 125200\C12203-84\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-84, Left Ventricle Anteroseptal E velocity |  |  | UCUM, mm/s, mm/s |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | AS A€™ Vel | Vel | meas | m_AS_A_Vel | r_Vel | unilateral | 125200\C12203-85\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-85, Left Ventricle Anteroseptal A velocity |  |  | UCUM, mm/s, mm/s |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Time to AS S | Time | meas | m_Time_To_AS_S | r_DopTime | unilateral | 125200\C12203-86\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-86, Left Ventricle Time To Anteroseptal S |  |  | UCUM, ms, ms |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | ES Default (HM) | meas | m_HM_LV_EF | r_HM_Default_ES_Border_Setting | unilateral | 125200\109070\C12203-87\C12228-11\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-87, Left Ventricle Default Border Setting |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | ED Default (HM) | meas | m_HM_LV_EF | r_HM_Default_ED_Border_Setting | unilateral | 125200\C12203-87\C12228-11\F-32011\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-87, Left Ventricle Default Border Setting |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | ES Current (HM) | meas | m_HM_LV_EF | r_HM_Current_ES_Border_Setting | unilateral | 125200\109070\C12203-88\C12228-11\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-88, Left Ventricle Border Setting |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | ED Current (HM) | meas | m_HM_LV_EF | r_HM_Current_ED_Border_Setting | unilateral | 125200\C12203-88\C12228-11\F-32011\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-88, Left Ventricle Border Setting |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | Other (HM) | Height (HM) | meas | m_HM_Other | r_HM_Height | unilateral | 125200\C12228-12\T-32600\T9910-218 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, T9910-218, Patient Height used in QLAB HM calculations |  |  | UCUM, mm, mm | [5 empty] | 99PMSBLUS, C12228-12, QLab
Adult Echocardiography Procedure Report | Adult Echo | Other (HM) | Weight (HM) | meas | m_HM_Other | r_HM_Weight | unilateral | 125200\C12228-12\T-32600\T9910-219 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, T9910-219, Patient Weight used in QLAB HM calculations |  |  | UCUM, g, grams | [5 empty] | 99PMSBLUS, C12228-12, QLab
Adult Echocardiography Procedure Report | Adult Echo | Other (HM) | BSA (HM) | meas | m_HM_Other | r_HM_BSA | unilateral | 125200\C12228-12\T-32600\T9910-220 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, T9910-220, Body Surface Area used in QLAB HM calculations |  |  | UCUM, mm2, mm2 | [5 empty] | 99PMSBLUS, C12228-12, QLab
Adult Echocardiography Procedure Report | Adult Echo | LVET (DOP) | Time | meas | m_LVET_DOP | r_DopTime | unilateral | 125200\122211\T-32600 | SNM3, T-32600, Left Ventricle | DCM, 122211, Left Ventricular ejection time |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | LVOT Vmax | Vmax | meas | m_LVOT_Vmax | r_Vmax | unilateral | 125200\11726-7\T-32600\T-32650 | SNM3, T-32600, Left Ventricle | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s | SNM3, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | EF BP-3DQ | EDV | meas | m_3DQ_EF_BP | r_3DQEDV | unilateral | 125200\18026-5\C12228-06\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-06, QLab 3DQ
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-3DQA | EDV | meas | m_3DQ_EF_VOL | r_3DQEDV | unilateral | 125200\18026-5\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | EDV (HM) | meas | m_HM_LV_EF | r_HM_EDV_LV | unilateral | 125200\18026-5\C12228-11\F-32011\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | 99PMSBLUS, T12224-01, 3D mode |  | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | LV EF BP-a2DQ | EDV | meas | m_2DQ_EF_BP | r_EDVol | unilateral | 125200\18026-5\C12228-09\C99-12228-05\F-32011\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LV EF BP-aCMQ | EDV | meas | m_CMQ_EF_BP | r_EDVol | unilateral | 125200\18026-5\C12228-10\C99-12228-07\F-32011\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | LV Vol | meas | m_A2Cd | r_SimpVolume | unilateral | 125200\125208\18026-5\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LV EF AP2-a2DQ | EDV | meas | m_2DQ_EF_Ap2 | r_EDVol | unilateral | 125200\18026-5\C12228-09\C99-12228-05\F-32011\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LV EF AP2-aCMQ | EDV | meas | m_CMQ_EF_Ap2 | r_EDVol | unilateral | 125200\18026-5\C12228-10\C99-12228-07\F-32011\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | LV Vol | meas | m_A4Cd | r_SimpVolume | unilateral | 125200\125208\18026-5\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LV EF AP4-a2DQ | EDV | meas | m_2DQ_EF_Ap4 | r_EDVol | unilateral | 125200\18026-5\C12228-09\C99-12228-05\F-32011\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LV EF AP4-aCMQ | EDV | meas | m_CMQ_EF_Ap4 | r_EDVol | unilateral | 125200\18026-5\C12228-10\C99-12228-07\F-32011\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | EF BP-3DQ | EF | meas | m_3DQ_EF_BP | r_3DQEF | unilateral | 125200\18043-0\C12228-06\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-06, QLab 3DQ
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-3DQA | EF | meas | m_3DQ_EF_VOL | r_3DQEF | unilateral | 125200\18043-0\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | EF (HM) | meas | m_HM_LV_EF | r_HM_EF | unilateral | 125200\18043-0\C12228-11\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | LV EF BP-a2DQ | EF | meas | m_2DQ_EF_BP | r_EF | unilateral | 125200\18043-0\C12228-09\C99-12228-05\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LV EF BP-aCMQ | EF | meas | m_CMQ_EF_BP | r_EF | unilateral | 125200\18043-0\C12228-10\C99-12228-07\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV EF AP2-a2DQ | EF | meas | m_2DQ_EF_Ap2 | r_EF | unilateral | 125200\18043-0\C12228-09\C99-12228-05\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LV EF AP2-aCMQ | EF | meas | m_CMQ_EF_Ap2 | r_EF | unilateral | 125200\18043-0\C12228-10\C99-12228-07\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV EF AP4-a2DQ | EF | meas | m_2DQ_EF_Ap4 | r_EF | unilateral | 125200\18043-0\C12228-09\C99-12228-05\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LV EF AP4-aCMQ | EF | meas | m_CMQ_EF_Ap4 | r_EF | unilateral | 125200\18043-0\C12228-10\C99-12228-07\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | IVRT | Time | meas | m_IVRT | r_DopTime | unilateral | 125200\18071-1\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18071-1, Left Ventricular Isovolumic Relaxation Time |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | Med IVRT | Time | meas | m_TDI_Med_IVRT | r_DopTime | unilateral | 125200\18071-1\G-0391\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18071-1, Left Ventricular Isovolumic Relaxation Time |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat IVRT | Time | meas | m_TDI_Lat_IVRT | r_DopTime | unilateral | 125200\18071-1\G-0392\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18071-1, Left Ventricular Isovolumic Relaxation Time |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | LVs (A/L) | Length | meas | m_LV_Vs_SPE | r_Length | unilateral | 125200\125226\18076-0\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18076-0, Left Ventricle systolic major axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | LV Length | meas | m_A2Cs | r_SimpDistance | unilateral | 125200\125208\18076-0\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18076-0, Left Ventricle systolic major axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | LV Length | meas | m_A4Cs | r_SimpDistance | unilateral | 125200\125208\18076-0\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18076-0, Left Ventricle systolic major axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LVLd Apical | Dist | meas | m_LV_LdApical | r_Distance | unilateral | 125200\18077-8\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18077-8, Left Ventricle diastolic major axis |  |  | UCUM, mm, mm
Adult Echocardiography Procedure Report | Adult Echo | LVd (A/L) | Length | meas | m_LV_Vd_SPE | r_Length | unilateral | 125200\125226\18077-8\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18077-8, Left Ventricle diastolic major axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | LV Length | meas | m_A2Cd | r_SimpDistance | unilateral | 125200\125208\18077-8\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18077-8, Left Ventricle diastolic major axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | LV Length | meas | m_A4Cd | r_SimpDistance | unilateral | 125200\125208\18077-8\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18077-8, Left Ventricle diastolic major axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | EF BP-3DQ | ESV | meas | m_3DQ_EF_BP | r_3DQESV | unilateral | 125200\18148-7\C12228-06\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-06, QLab 3DQ
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-3DQA | ESV | meas | m_3DQ_EF_VOL | r_3DQESV | unilateral | 125200\18148-7\C12228-07\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | ESV (HM) | meas | m_HM_LV_EF | r_HM_ESV_LV | unilateral | 125200\109070\18148-7\C12228-11\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | 99PMSBLUS, T12224-01, 3D mode |  | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | LV EF BP-a2DQ | ESV | meas | m_2DQ_EF_BP | r_ESVol | unilateral | 125200\109070\18148-7\C12228-09\C99-12228-05\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LV EF BP-aCMQ | ESV | meas | m_CMQ_EF_BP | r_ESVol | unilateral | 125200\109070\18148-7\C12228-10\C99-12228-07\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | LV Vol | meas | m_A2Cs | r_SimpVolume | unilateral | 125200\125208\18148-7\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LV EF AP2-a2DQ | ESV | meas | m_2DQ_EF_Ap2 | r_ESVol | unilateral | 125200\109070\18148-7\C12228-09\C99-12228-05\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LV EF AP2-aCMQ | ESV | meas | m_CMQ_EF_Ap2 | r_ESVol | unilateral | 125200\109070\18148-7\C12228-10\C99-12228-07\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | LV Vol | meas | m_A4Cs | r_SimpVolume | unilateral | 125200\125208\18148-7\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LV EF AP4-a2DQ | ESV | meas | m_2DQ_EF_Ap4 | r_ESVol | unilateral | 125200\109070\18148-7\C12228-09\C99-12228-05\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LV EF AP4-aCMQ | ESV | meas | m_CMQ_EF_Ap4 | r_ESVol | unilateral | 125200\109070\18148-7\C12228-10\C99-12228-07\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LVPWd (MM) | Dist | meas | m_LVPWd_MM | r_Distance | unilateral | 125200\18152-9\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18152-9, Left Ventricle Posterior Wall Diastolic Thickness |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LVPWd (2D) | Dist | meas | m_LVPWd_2D | r_Distance | unilateral | 125200\18152-9\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18152-9, Left Ventricle Posterior Wall Diastolic Thickness |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | IVSd (MM) | Dist | meas | m_IVSd_MM | r_Distance | unilateral | 125200\18154-5\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18154-5, Interventricular Septum Diastolic Thickness |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | IVSd (2D) | Dist | meas | m_IVSd_2D | r_Distance | unilateral | 125200\18154-5\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18154-5, Interventricular Septum Diastolic Thickness |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LVPWs (MM) | Dist | meas | m_LVPWs_MM | r_Distance | unilateral | 125200\18156-0\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18156-0, Left Ventricle Posterior Wall Systolic Thickness |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LVPWs (2D) | Dist | meas | m_LVPWs_2D | r_Distance | unilateral | 125200\18156-0\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18156-0, Left Ventricle Posterior Wall Systolic Thickness |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | IVSs (MM) | Dist | meas | m_IVSs_MM | r_Distance | unilateral | 125200\18158-6\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18158-6, Interventricular Septum Systolic Thickness |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | IVSs (2D) | Dist | meas | m_IVSs_2D | r_Distance | unilateral | 125200\18158-6\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18158-6, Interventricular Septum Systolic Thickness |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LVOT Accel Time | Slope | meas | m_LVOT_AccelTime | r_DopSlope | unilateral | 125200\20167-3\T-32600\T-32650 | SNM3, T-32600, Left Ventricle | LN, 20167-3, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 | SNM3, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | Med Accel Time | Time | meas | m_TDIMedAccelTime | r_DopTime | unilateral | 125200\20168-1\G-0391\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat Accel Time | Time | meas | m_TDILatAccelTime | r_DopTime | unilateral | 125200\20168-1\G-0392\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | LVOT Accel Time | Time | meas | m_LVOT_AccelTime | r_DopTime | unilateral | 125200\20168-1\T-32600\T-32650 | SNM3, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms | SNM3, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | Med Decel Time | Time | meas | m_TDIMedDecelTime | r_DopTime | unilateral | 125200\20217-6\G-0391\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat Decel Time | Time | meas | m_TDILatDecelTime | r_DopTime | unilateral | 125200\20217-6\G-0392\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | LV ET | Time | meas | m_LV_Eject_Tm | r_MMTime | unilateral | 125200\20222-6\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 20222-6, Ejection Time |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LVOT Vmax | Max PG | meas | m_LVOT_Vmax | r_MaxPG | unilateral | 125200\20247-3\T-32600\T-32650 | SNM3, T-32600, Left Ventricle | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SNM3, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | LVOT VTI | Mean PG | meas | m_LVOT_VTI | r_CardPGmean | unilateral | 125200\20256-4\T-32600\T-32650 | SNM3, T-32600, Left Ventricle | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg | SNM3, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | LVOT VTI | Vmean | meas | m_LVOT_VTI | r_CardTAPV | unilateral | 125200\20352-1\T-32600\T-32650 | SNM3, T-32600, Left Ventricle | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s | SNM3, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | LVOT VTI | VTI | meas | m_LVOT_VTI | r_VTI | unilateral | 125200\20354-7\T-32600\T-32650 | SNM3, T-32600, Left Ventricle | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm | SNM3, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | LVIDd (MM) | Dist | meas | m_LVIDd_MM | r_Distance | unilateral | 125200\29436-3\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 29436-3, Left Ventricle Internal End Diastolic Dimension |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LVIDd (2D) | Dist | meas | m_LVIDd_2D | r_Distance | unilateral | 125200\29436-3\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 29436-3, Left Ventricle Internal End Diastolic Dimension |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LVIDs (MM) | Dist | meas | m_LVIDs_MM | r_Distance | unilateral | 125200\29438-9\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 29438-9, Left Ventricle Internal Systolic Dimension |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LVIDs (2D) | Dist | meas | m_LVIDs_2D | r_Distance | unilateral | 125200\29438-9\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 29438-9, Left Ventricle Internal Systolic Dimension |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LV PEP | Time | meas | m_LV_PreEject_Tm | r_MMTime | unilateral | 125200\59085-1\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 59085-1, Pre-Ejection Period |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LV Post Wall | Dist | meas | m_LV_Post_Wall | r_Distance | unilateral | 125200\59089-3\G-03A2\R-42175\T-32600 | SNM3, T-32600, Left Ventricle | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm | SRT, R-42175, Posterior Wall | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Time to Med S | Time | meas | m_TimeMedPeakS | r_DopTime | unilateral | 125200\59095-0\G-0391\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | LN, 59095-0, Time to Left Ventricle S Tissue Velocity |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Time to Lat S | Time | meas | m_TimeLatPeakS | r_DopTime | unilateral | 125200\59095-0\G-0392\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | LN, 59095-0, Time to Left Ventricle S Tissue Velocity |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Time to Med E` | Time | meas | m_TimeMedPeakE | r_DopTime | unilateral | 125200\59096-8\G-0391\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | LN, 59096-8, Time to Left Ventricle E Tissue Velocity |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Time to Lat E` | Time | meas | m_TimeLatPeakE | r_DopTime | unilateral | 125200\59096-8\G-0392\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | LN, 59096-8, Time to Left Ventricle E Tissue Velocity |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Med E` Area | VTI | meas | m_AreaUnderMedPeakE | r_VTI | unilateral | 125200\59124-8\G-0391\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | LN, 59124-8, Tissue Velocity Time Integral (VTI) for the area under Left Ventricle E wave |  |  | UCUM, mm, mm | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat E` Area | VTI | meas | m_AreaUnderLatPeakE | r_VTI | unilateral | 125200\59124-8\G-0392\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | LN, 59124-8, Tissue Velocity Time Integral (VTI) for the area under Left Ventricle E wave |  |  | UCUM, mm, mm | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Med A` Area | VTI | meas | m_AreaUnderMedPeakA | r_VTI | unilateral | 125200\59125-5\G-0391\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | LN, 59125-5, Tissue Velocity Time Integral (VTI) for the area under Left Ventricle A wave |  |  | UCUM, mm, mm | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat A` Area | VTI | meas | m_AreaUnderLatPeakA | r_VTI | unilateral | 125200\59125-5\G-0392\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | LN, 59125-5, Tissue Velocity Time Integral (VTI) for the area under Left Ventricle A wave |  |  | UCUM, mm, mm | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | HR - LV | HR | meas | m_HR_LV | r_HeartRate | unilateral | 125200\8867-4\T-32600 | SNM3, T-32600, Left Ventricle | LN, 8867-4, Heart rate |  |  | UCUM, {H.B.}/min, bpm
Adult Echocardiography Procedure Report | Adult Echo | Other (HM) | HR | meas | m_HM_Other | r_HeartRate | unilateral | 125200\8867-4\C12228-11\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | LN, 8867-4, Heart rate |  |  | UCUM, {H.B.}/min, bpm |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | CI (HM) | meas | m_HM_LV_EF | r_HM_CI | unilateral | 125200\C12228-11\F-32110\G-A19C\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | 99PMSBLUS, T12224-01, 3D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-3DQA | SV | meas | m_3DQ_EF_VOL | r_3DQSV | unilateral | 125200\C12228-07\F-32120\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | EF VOL-HM | SV (HM) | meas | m_HM_LV_EF | r_HM_SV | unilateral | 125200\C12228-11\F-32120\T12224-01\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | LVs (A/L) | Area | meas | m_LV_Vs_SPE | r_Area | unilateral | 125200\125226\G-0374\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | LV ESA | Area | meas | m_LV_ESA | r_Area | unilateral | 125200\109070\G-0374\G-039B\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LV FAC SAX-a2DQ | ESA | meas | m_2DQ_FAC_Sax | r_ESA | unilateral | 125200\109070\C12226-06\C12228-09\C99-12228-05\G-0374\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LV FAC SAX-aCMQ | ESA | meas | m_CMQ_FAC_Sax | r_ESA | unilateral | 125200\109070\C12226-06\C12228-10\C99-12228-07\G-0374\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | A2Cs | LV Area | meas | m_A2Cs | r_SimpArea | unilateral | 125200\125208\G-0374\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LV FAC AP2-a2DQ | ESA | meas | m_2DQ_FAC_Ap2 | r_ESA | unilateral | 125200\109070\C12228-09\C99-12228-05\G-0374\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | A4Cs | LV Area | meas | m_A4Cs | r_SimpArea | unilateral | 125200\125208\G-0374\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LV FAC AP4-a2DQ | ESA | meas | m_2DQ_FAC_Ap4 | r_ESA | unilateral | 125200\109070\C12228-09\C99-12228-05\G-0374\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LVd (A/L) | Area | meas | m_LV_Vd_SPE | r_Area | unilateral | 125200\125226\G-0375\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | LV EDA | Area | meas | m_LV_EDA | r_Area | unilateral | 125200\F-32011\G-0375\G-039B\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | LV FAC SAX-a2DQ | EDA | meas | m_2DQ_FAC_Sax | r_EDA | unilateral | 125200\C12226-06\C12228-09\C99-12228-05\F-32011\G-0375\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LV FAC SAX-aCMQ | EDA | meas | m_CMQ_FAC_Sax | r_EDA | unilateral | 125200\C12226-06\C12228-10\C99-12228-07\F-32011\G-0375\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | A2Cd | LV Area | meas | m_A2Cd | r_SimpArea | unilateral | 125200\125208\G-0375\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LV FAC AP2-a2DQ | EDA | meas | m_2DQ_FAC_Ap2 | r_EDA | unilateral | 125200\C12228-09\C99-12228-05\F-32011\G-0375\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | A4Cd | LV Area | meas | m_A4Cd | r_SimpArea | unilateral | 125200\125208\G-0375\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LV FAC AP4-a2DQ | EDA | meas | m_2DQ_FAC_Ap4 | r_EDA | unilateral | 125200\C12228-09\C99-12228-05\F-32011\G-0375\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LVAd Sax Endo | Area | meas | m_LV_AdSaxEndo | r_Area | unilateral | 125200\G-0375\G-039B\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  |  | UCUM, mm2, mm2 |  |  | SRT, G-039B, Parasternal short axis at the Papillary Muscle level
Adult Echocardiography Procedure Report | Adult Echo | LV FAC SAX-a2DQ | FAC | meas | m_2DQ_FAC_Sax | r_FAC | unilateral | 125200\C12226-06\C12228-09\C99-12228-05\G-0376\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0376, Left Ventricular Fractional Area Change |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LV FAC SAX-aCMQ | FAC | meas | m_CMQ_FAC_Sax | r_FAC | unilateral | 125200\C12226-06\C12228-10\C99-12228-07\G-0376\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0376, Left Ventricular Fractional Area Change |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV FAC AP2-a2DQ | FAC | meas | m_2DQ_FAC_Ap2 | r_FAC | unilateral | 125200\C12228-09\C99-12228-05\G-0376\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0376, Left Ventricular Fractional Area Change |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LV FAC AP4-a2DQ | FAC | meas | m_2DQ_FAC_Ap4 | r_FAC | unilateral | 125200\C12228-09\C99-12228-05\G-0376\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0376, Left Ventricular Fractional Area Change |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LVAd Sax Epi | Area | meas | m_LV_AdSaxEpi | r_Area | unilateral | 125200\G-0379\G-039B\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0379, Left Ventricle Epicardial Diastolic Area, psax pap view |  |  | UCUM, mm2, mm2 |  |  | SRT, G-039B, Parasternal short axis at the Papillary Muscle level
Adult Echocardiography Procedure Report | Adult Echo | Med E` Vel | Vel | meas | m_Med_PeakEVel | r_Vel | unilateral | 125200\G-037A\G-0391\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-037A, Left Ventricular Peak Early Diastolic Tissue Velocity |  |  | UCUM, mm/s, mm/s | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat E` Vel | Vel | meas | m_Lat_PeakEVel | r_Vel | unilateral | 125200\G-037A\G-0392\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-037A, Left Ventricular Peak Early Diastolic Tissue Velocity |  |  | UCUM, mm/s, mm/s | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Med A` Vel | Vel | meas | m_Med_PeakAVel | r_Vel | unilateral | 125200\G-037C\G-0391\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-037C, LV Peak Diastolic Tissue Velocity During Atrial Systole |  |  | UCUM, mm/s, mm/s | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat A` Vel | Vel | meas | m_Lat_PeakAVel | r_Vel | unilateral | 125200\G-037C\G-0392\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-037C, LV Peak Diastolic Tissue Velocity During Atrial Systole |  |  | UCUM, mm/s, mm/s | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Med S Vel | Vel | meas | m_Med_PeakSVel | r_Vel | unilateral | 125200\G-037D\G-0391\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-037D, Left Ventricular Peak Systolic Tissue Velocity |  |  | UCUM, mm/s, mm/s | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat S Vel | Vel | meas | m_Lat_PeakSVel | r_Vel | unilateral | 125200\G-037D\G-0392\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-037D, Left Ventricular Peak Systolic Tissue Velocity |  |  | UCUM, mm/s, mm/s | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | IVCT | Time | meas | m_IVCT | r_DopTime | unilateral | 125200\G-037E\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-037E, Left Ventricular Isovolumic Contraction Time |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | Med IVCT | Time | meas | m_TDI_Med_IVCT | r_DopTime | unilateral | 125200\G-037E\G-0391\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-037E, Left Ventricular Isovolumic Contraction Time |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat IVCT | Time | meas | m_TDI_Lat_IVCT | r_DopTime | unilateral | 125200\G-037E\G-0392\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-037E, Left Ventricular Isovolumic Contraction Time |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | LVOT Diam | Dist | meas | m_LVOT_Diam | r_Distance | unilateral | 125200\G-038F\T-32600\T-32650 | SNM3, T-32600, Left Ventricle | SRT, G-038F, Cardiovascular Orifice Diameter |  |  | UCUM, mm, mm | SNM3, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | MV TMAD AP2-a2DQ | MidPt Disp | meas | m_2DQ_TMAD_Ap2 | r_MidPoint_Disp | unilateral | 125200\C12203-62\C12228-09\C99-12228-06\G-03A2\G-A19B\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12203-62, Displacement Midpoint of L/R Mitral Annular Hingepoints to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | MV TMAD AP2-aCMQ | MidPt Disp | meas | m_CMQ_TMAD_Ap2 | r_MidPoint_Disp | unilateral | 125200\C12203-62\C12228-10\C99-12228-06\G-03A2\G-A19B\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12203-62, Displacement Midpoint of L/R Mitral Annular Hingepoints to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | MV TMAD AP4-a2DQ | MidPt Disp | meas | m_2DQ_TMAD_Ap4 | r_MidPoint_Disp | unilateral | 125200\C12203-62\C12228-09\C99-12228-06\G-03A2\G-A19C\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12203-62, Displacement Midpoint of L/R Mitral Annular Hingepoints to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | MV TMAD AP4-aCMQ | MidPt Disp | meas | m_CMQ_TMAD_Ap4 | r_MidPoint_Disp | unilateral | 125200\C12203-62\C12228-10\C99-12228-06\G-03A2\G-A19C\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12203-62, Displacement Midpoint of L/R Mitral Annular Hingepoints to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | MV TMAD AP2-a2DQ | MidPt Norm | meas | m_2DQ_TMAD_Ap2 | r_MidPoint_Norm | unilateral | 125200\C12203-63\C12228-09\C99-12228-06\G-03A2\G-A19B\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12203-63, Normalized Midpoint of L/R MV Annular Hingepoints to Apex |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | MV TMAD AP2-aCMQ | MidPt Norm | meas | m_CMQ_TMAD_Ap2 | r_MidPoint_Norm | unilateral | 125200\C12203-63\C12228-10\C99-12228-06\G-03A2\G-A19B\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12203-63, Normalized Midpoint of L/R MV Annular Hingepoints to Apex |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | MV TMAD AP4-a2DQ | MidPt Norm | meas | m_2DQ_TMAD_Ap4 | r_MidPoint_Norm | unilateral | 125200\C12203-63\C12228-09\C99-12228-06\G-03A2\G-A19C\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12203-63, Normalized Midpoint of L/R MV Annular Hingepoints to Apex |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | MV TMAD AP4-aCMQ | MidPt Norm | meas | m_CMQ_TMAD_Ap4 | r_MidPoint_Norm | unilateral | 125200\C12203-63\C12228-10\C99-12228-06\G-03A2\G-A19C\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12203-63, Normalized Midpoint of L/R MV Annular Hingepoints to Apex |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | MV TMAD AP2-a2DQ | Pt2 | meas | m_2DQ_TMAD_Ap2 | r_Point2 | unilateral | 125200\C12203-64\C12228-09\C99-12228-06\G-03A2\G-A19B\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12203-64, Displacement of the Right Mitral Annular Hingepoint to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | MV TMAD AP2-aCMQ | Pt2 | meas | m_CMQ_TMAD_Ap2 | r_Point2 | unilateral | 125200\C12203-64\C12228-10\C99-12228-06\G-03A2\G-A19B\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12203-64, Displacement of the Right Mitral Annular Hingepoint to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | MV TMAD AP4-a2DQ | Pt2 | meas | m_2DQ_TMAD_Ap4 | r_Point2 | unilateral | 125200\C12203-64\C12228-09\C99-12228-06\G-03A2\G-A19C\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12203-64, Displacement of the Right Mitral Annular Hingepoint to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | MV TMAD AP4-aCMQ | Pt2 | meas | m_CMQ_TMAD_Ap4 | r_Point2 | unilateral | 125200\C12203-64\C12228-10\C99-12228-06\G-03A2\G-A19C\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12203-64, Displacement of the Right Mitral Annular Hingepoint to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | MV TMAD AP2-a2DQ | Pt1 | meas | m_2DQ_TMAD_Ap2 | r_Point1 | unilateral | 125200\C12203-65\C12228-09\C99-12228-06\G-03A2\G-A19B\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12203-65, Displacement of the Left Mitral Annular Hingepoint to the Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | MV TMAD AP2-aCMQ | Pt1 | meas | m_CMQ_TMAD_Ap2 | r_Point1 | unilateral | 125200\C12203-65\C12228-10\C99-12228-06\G-03A2\G-A19B\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12203-65, Displacement of the Left Mitral Annular Hingepoint to the Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | MV TMAD AP4-a2DQ | Pt1 | meas | m_2DQ_TMAD_Ap4 | r_Point1 | unilateral | 125200\C12203-65\C12228-09\C99-12228-06\G-03A2\G-A19C\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12203-65, Displacement of the Left Mitral Annular Hingepoint to the Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | MV TMAD AP4-aCMQ | Pt1 | meas | m_CMQ_TMAD_Ap4 | r_Point1 | unilateral | 125200\C12203-65\C12228-10\C99-12228-06\G-03A2\G-A19C\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12203-65, Displacement of the Left Mitral Annular Hingepoint to the Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | MV D-E Exc | Dist | meas | m_MV_DE_Exc | r_Distance | unilateral | 125200\C12207-01\G-0394\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-01, Mitral Valve D-E Excursion |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | MV D-E Slope | Slope | meas | m_MV_DE_Slope | r_MMSlope | unilateral | 125200\C12207-02\G-0394\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-02, Mitral Valve D-E Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | MV E-E Sep | Dist | meas | m_MV_EE_Separ | r_Distance | unilateral | 125200\C12207-03\G-0394\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-03, Mitral Valve E-E Separation |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | MV A-C Interval | Time | meas | m_MV_AC_Int | r_MMTime | unilateral | 125200\C12207-04\G-0394\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-04, Mitral Valve A-C Interval |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | MV AP Diam-MVN | Dist | meas | m_mvq_APDiam | r_Distance | unilateral | 125200\C12207-07\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-07, Mitral Valve Anterior Posterior Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV AL-PM Diam-MVN | Dist | meas | m_mvq_ALPMDiam | r_Distance | unilateral | 125200\C12207-08\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-08, Mitral Valve Anterolateral Posteromedial Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Ann3D Circ-MVN | Dist | meas | m_mvq_Ann3DCirc | r_Distance | unilateral | 125200\C12207-09\C12228-08\T12224-01\T-35300\T-35313 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-09, Mitral Valve Annulus 3D Circumference |  |  | UCUM, mm, mm | SRT, T-35313, Mitral Annulus | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Ant Leaflet-MVN | Area | meas | m_mvq_AntLeaf | r_Area | unilateral | 125200\C12207-10\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-10, Mitral Valve Anterior Leaflet Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Post Leaflet-MVN | Area | meas | m_mvq_PostLeaf | r_Area | unilateral | 125200\C12207-11\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-11, Mitral Valve Posterior Leaflet Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV A Leaf 3D Area-MVN | Area | meas | m_mvq_AntLeaf3DA | r_Area | unilateral | 125200\C12207-12\C12228-08\T12224-01\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-12, Mitral Valve Anterior Leaflets 3D Area |  |  | UCUM, mm2, mm2 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV P Leaf 3D Area-MVN | Area | meas | m_mvq_PostLeaf3DA | r_Area | unilateral | 125200\C12207-13\C12228-08\T12224-01\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-13, Mitral Valve Posterior Leaflets 3D Area |  |  | UCUM, mm2, mm2 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Tenting Vol-MVN | Volume | meas | m_mvq_TentVol | r_Volume | unilateral | 125200\C12207-14\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-14, Mitral Valve Tenting Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Prolapse Vol-MVN | Volume | meas | m_mvq_ProVol | r_Volume | unilateral | 125200\C12207-15\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-15, Mitral Valve Prolapse Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV AL Chord-MVN | Dist | meas | m_mvq_ALChrd | r_Distance | unilateral | 125200\C12207-16\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-16, Mitral Valve Anterolateral Chord Length |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV PM Chord-MVN | Dist | meas | m_mvq_PMChord | r_Distance | unilateral | 125200\C12207-17\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-17, Mitral Valve Posteromedial Chord Length |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Ant Leaf Angle-MVN | Angle | meas | m_mvq_ALeafAngle | r_Angle | unilateral | 125200\C12207-18\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-18, Mitral Valve Anterior Leaflet Angle |  |  | UCUM, deg, deg |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Post Leaf Angle-MVN | Angle | meas | m_mvq_PLeafAngle | r_Angle | unilateral | 125200\C12207-19\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-19, Mitral Valve Posterior Leaflet Angle |  |  | UCUM, deg, deg |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Nonplanar Angle-MVN | Angle | meas | m_mvq_NonplnAngle | r_Angle | unilateral | 125200\C12207-20\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-20, Mitral Valve Nonplanar Angle |  |  | UCUM, deg, deg |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Ao OR-Mitr Ang-MVN | Angle | meas | m_mvq_AoOrMitrAngle | r_Angle | unilateral | 125200\C12207-21\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-21, Mitral Valve Aortic Orifice to Mitral Plane Angle |  |  | UCUM, deg, deg |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Coapt 2D-MVN | Dist | meas | m_mvq_Coapt2D | r_Distance | unilateral | 125200\C12207-22\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-22, Mitral Valve Coaptation 2D Length |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Tenting | Dist | meas | m_MV_Tenting | r_Distance | unilateral | 125200\C12207-23\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-23, Mitral Valve Tenting Height |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | MV Tenting Height-MVN | Dist | meas | m_mvq_TentHeight | r_Distance | unilateral | 125200\C12207-23\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-23, Mitral Valve Tenting Height |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Prolapse Height-MVN | Dist | meas | m_mvq_ProlHeight | r_Distance | unilateral | 125200\C12207-24\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-24, Mitral Valve Prolapse Height |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Ann2D Circ-MVN | Dist | meas | m_mvq_Ann2DCirc | r_Distance | unilateral | 125200\C12207-25\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-25, Mitral Valve Annulus 2D Circumference |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Ann3D Min Area-MVN | Area | meas | m_mvq_Ann3DMinArea | r_Area | unilateral | 125200\C12207-26\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-26, Mitral Valve 3D Annular Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Ann H/CC Diam-MVN | Ratio | meas | m_mvq_AnnHPDiam | r_MVQ_Ratio | unilateral | 125200\C12207-27\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-27, Mitral Valve Annulus Height/AP Diameter Ratio |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Ann Ellipsicity-MVN | Ellipsicity | meas | m_mvq_AnnEllip | r_Ellipsicity | unilateral | 125200\C12207-28\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-28, Mitral Valve Annulus Ellipsicity |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV MR Orifice Area-MVN | Area | meas | m_mvq_OriArea | r_Area | unilateral | 125200\C12207-29\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-29, Mitral Valve Regurgitant Orifice Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV A1 Tent Vol-MVN | Volume | meas | m_mvq_A1TentVol | r_Volume | unilateral | 125200\C12207-30\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-30, Mitral Valve A1 Tenting Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV A1 Prolapse Vol-MVN | Volume | meas | m_mvq_A1ProVol | r_Volume | unilateral | 125200\C12207-31\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-31, Mitral Valve A1 Prolapse Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV A2 Tent Vol-MVN | Volume | meas | m_mvq_A2TentVol | r_Volume | unilateral | 125200\C12207-32\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-32, Mitral Valve A2 Tenting Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV A2 Prolapse Vol-MVN | Volume | meas | m_mvq_A2ProVol | r_Volume | unilateral | 125200\C12207-33\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-33, Mitral Valve A2 Prolapse Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV A3 Tent Vol-MVN | Volume | meas | m_mvq_A3TentVol | r_Volume | unilateral | 125200\C12207-34\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-34, Mitral Valve A3 Tenting Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV A3 Prolapse Vol-MVN | Volume | meas | m_mvq_A3ProVol | r_Volume | unilateral | 125200\C12207-35\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-35, Mitral Valve A3 Prolapse Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV P1 Tent Vol-MVN | Volume | meas | m_mvq_P1TentVol | r_Volume | unilateral | 125200\C12207-36\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-36, Mitral Valve P1 Tenting Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV P1 Prolapse Vol-MVN | Volume | meas | m_mvq_P1ProVol | r_Volume | unilateral | 125200\C12207-37\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-37, Mitral Valve P1 Prolapse Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV P2 Tent Vol-MVN | Volume | meas | m_mvq_P2TentVol | r_Volume | unilateral | 125200\C12207-38\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-38, Mitral Valve P2 Tenting Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV P2 Prolapse Vol-MVN | Volume | meas | m_mvq_P2ProVol | r_Volume | unilateral | 125200\C12207-39\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-39, Mitral Valve P2 Prolapse Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV P3 Tent Vol-MVN | Volume | meas | m_mvq_P3TentVol | r_Volume | unilateral | 125200\C12207-40\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-40, Mitral Valve P3 Tenting Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | R to MV Open | Time | meas | m_RtoMVo_DopTime | r_DopTime | unilateral | 125200\C12207-41\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-41, R Wave to Mitral Valve Opening Time |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | R to MV Closure | Time | meas | m_RtoMVc_DopTime | r_DopTime | unilateral | 125200\C12207-42\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-42, R Wave to Mitral Valve Closure Time |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | MV P3 Prolapse Vol-MVN | Volume | meas | m_mvq_P3ProVol | r_Volume | unilateral | 125200\C12207-43\C12228-08\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-43, Mitral Valve P3 Prolapse Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Leaf 3D Area-MVN | Area | meas | m_mvq_Leaf3DArea | r_Area | unilateral | 125200\C12207-44\C12228-08\T12224-01\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-44, Mitral Valve Leaflets 3D Area |  |  | UCUM, mm2, mm2 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV Ant Leaflet | Dist | meas | m_MV_Ant_Leaflet | r_Distance | unilateral | 125200\C12207-47\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-47, Mitral Valve Anterior Leaflet Width |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | MV Post Leaflet | Dist | meas | m_MV_Post_Leaflet | r_Distance | unilateral | 125200\C12207-48\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-48, Mitral Valve Posterior Leaflet Width |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | MAPSE | Dist | meas | m_MAPSE | r_Distance | unilateral | 125200\C12207-50\G-0394\T-35300\T-35313 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-50, Mitral Annular Plane Systolic Excursion |  |  | UCUM, mm, mm | SRT, T-35313, Mitral Annulus | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | MV P½t | Vmax | meas | m_MV_PHalfT | r_Vmax | unilateral | 125200\C12222-03\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12222-03, Pressure Half-Time Peak velocity |  |  | UCUM, mm/s, mm/s
Adult Echocardiography Procedure Report | Adult Echo | MV Prop V | Slope | meas | m_MV_Prop_V | r_MMSlope | unilateral | 125200\C12222-24\G-0394\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12222-24, Flow Propagation Velocity |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | MV Prop V | Time | meas | m_MV_Prop_V | r_MMTime | unilateral | 125200\C12222-25\G-0394\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12222-25, Flow Propagation Time |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | MV Ann Height-MVN | Dist | meas | m_mvq_AnnHt | r_Distance | unilateral | 125200\C12228-08\C7470-01\G-03A2\T-35300\T-35313 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C7470-01, Height |  |  | UCUM, mm, mm | SRT, T-35313, Mitral Annulus | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MV R-R Interval | Time | meas | m_MV_RR_Interval | r_DopTime | unilateral | 125200\122182\T-35300 | SNM3, T-35300, Mitral Valve | DCM, 122182, R-R interval |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | MV Vmax | Vmax | meas | m_MV_Vmax | r_Vmax | unilateral | 125200\11726-7\R-42047\T-35300 | SNM3, T-35300, Mitral Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MR Vmax | Vmax | meas | m_MR_Vmax | r_Vmax | unilateral | 125200\11726-7\R-42E61\T-35300 | SNM3, T-35300, Mitral Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Peak A Vel | Vel | meas | m_MV_PeakAVel | r_Vel | unilateral | 125200\17978-8\R-42047\T-35300 | SNM3, T-35300, Mitral Valve | LN, 17978-8, Mitral Valve A-Wave Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | LV dP/dt | dP/dt | meas | m_LV_dPdt | r_DopDpDt | unilateral | 125200\18035-6\T-35300 | SNM3, T-35300, Mitral Valve | LN, 18035-6, Mitral Regurgitation dP/dt derived from Mitral Reg velocity |  |  | UCUM, mmHg, mmHg
Adult Echocardiography Procedure Report | Adult Echo | MV EPSS | Dist | meas | m_MV_EPSS | r_Distance | unilateral | 125200\18036-4\G-0394\T-35300 | SNM3, T-35300, Mitral Valve | LN, 18036-4, Mitral Valve EPSS, E wave |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | MV Peak E Vel | Vel | meas | m_MV_PeakEVel | r_Vel | unilateral | 125200\18037-2\R-42047\T-35300 | SNM3, T-35300, Mitral Valve | LN, 18037-2, Mitral Valve E-Wave Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV E-F Slope | Slope | meas | m_MV_EF_Slope | r_MMSlope | unilateral | 125200\18040-6\G-0394\T-35300 | SNM3, T-35300, Mitral Valve | LN, 18040-6, Mitral Valve E-F Slope by M-Mode |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | MV Accel Time | Slope | meas | m_MV_AccelTime | r_DopSlope | unilateral | 125200\20167-3\R-42047\T-35300 | SNM3, T-35300, Mitral Valve | LN, 20167-3, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Accel Time | Time | meas | m_MV_AccelTime | r_DopTime | unilateral | 125200\20168-1\R-42047\T-35300 | SNM3, T-35300, Mitral Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Dec Slope | Slope | meas | m_MV_DecSlope | r_DopSlope | unilateral | 125200\20216-8\R-42047\T-35300 | SNM3, T-35300, Mitral Valve | LN, 20216-8, Deceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Decel Time | Time | meas | m_MV_DecelTime | r_DopTime | unilateral | 125200\20217-6\R-42047\T-35300 | SNM3, T-35300, Mitral Valve | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Vmax | Max PG | meas | m_MV_Vmax | r_MaxPG | unilateral | 125200\20247-3\R-42047\T-35300 | SNM3, T-35300, Mitral Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MR Vmax | Max PG | meas | m_MR_Vmax | r_MaxPG | unilateral | 125200\20247-3\R-42E61\T-35300 | SNM3, T-35300, Mitral Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MV VTI | Mean PG | meas | m_MV_VTI | r_CardPGmean | unilateral | 125200\20256-4\R-42047\T-35300 | SNM3, T-35300, Mitral Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MR VTI | Mean PG | meas | m_MR_VTI | r_CardPGmean | unilateral | 125200\20256-4\R-42E61\T-35300 | SNM3, T-35300, Mitral Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MV P½t | P½t | meas | m_MV_PHalfT | r_PHalftime | unilateral | 125200\20280-4\T-35300 | SNM3, T-35300, Mitral Valve | LN, 20280-4, Pressure Half-Time |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | MV VTI | Vmean | meas | m_MV_VTI | r_CardTAPV | unilateral | 125200\20352-1\R-42047\T-35300 | SNM3, T-35300, Mitral Valve | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MR VTI | Vmean | meas | m_MR_VTI | r_CardTAPV | unilateral | 125200\20352-1\R-42E61\T-35300 | SNM3, T-35300, Mitral Valve | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MV VTI | VTI | meas | m_MV_VTI | r_VTI | unilateral | 125200\20354-7\R-42047\T-35300 | SNM3, T-35300, Mitral Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MR VTI | VTI | meas | m_MR_VTI | r_VTI | unilateral | 125200\20354-7\R-42E61\T-35300 | SNM3, T-35300, Mitral Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Radius | Dist | meas | m_MV_Radius | r_Distance | unilateral | 125200\59102-4\R-42047\T-35300 | SNM3, T-35300, Mitral Valve | LN, 59102-4, Flow Radius |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MR Radius | Dist | meas | m_MR_Radius | r_Distance | unilateral | 125200\59102-4\R-42E61\T-35300 | SNM3, T-35300, Mitral Valve | LN, 59102-4, Flow Radius |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MR Jet Area | Area | meas | m_MR_Jet_Area | r_Area | unilateral | 125200\59123-0\G-03A2\R-42E61\T-35300 | SNM3, T-35300, Mitral Valve | LN, 59123-0, Jet Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Alias Vel | Alias Vel | meas | m_MV_AliasVel | r_2DVel | unilateral | 125200\59130-5\R-42047\T-35300 | SNM3, T-35300, Mitral Valve | LN, 59130-5, Alias Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MR Alias Vel | Alias Vel | meas | m_MR_AliasVel | r_2DVel | unilateral | 125200\59130-5\R-42E61\T-35300 | SNM3, T-35300, Mitral Valve | LN, 59130-5, Alias Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | HR - MV | HR | meas | m_HR_MV | r_HeartRate | unilateral | 125200\8867-4\T-35300 | SNM3, T-35300, Mitral Valve | LN, 8867-4, Heart rate |  |  | UCUM, {H.B.}/min, bpm
Adult Echocardiography Procedure Report | Adult Echo | MV Ann2D Area-MVN | Area | meas | m_mvq_Ann2DArea | r_Area | unilateral | 125200\C12228-08\G-03A2\G-A166\T-35300\T-35313 | SNM3, T-35300, Mitral Valve | SNM3, G-A166, Area |  |  | UCUM, mm2, mm2 | SRT, T-35313, Mitral Annulus | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | MR VC | Dist | meas | m_MR_VC | r_Distance | unilateral | 125200\G-03A2\M-02550\R-421AA\R-42E61\T-35300 | SNM3, T-35300, Mitral Valve | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, R-421AA, Vena Contracta | SRT, G-03A2, 2D mode |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MV A Dur | Time | meas | m_MVA_Dur | r_DopTime | unilateral | 125200\G-0385\T-35300 | SNM3, T-35300, Mitral Valve | SRT, G-0385, Mitral Valve A-Wave Duration |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | MCOT | Time | meas | m_MCOT | r_DopTime | unilateral | 125200\G-0387\T-35300 | SNM3, T-35300, Mitral Valve | SRT, G-0387, Mitral Valve Closure to Opening Time |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | MV Area (Planim) | Area | meas | m_MV_Area | r_Area | unilateral | 125200\125220\G-038E\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125220, Planimetry
Adult Echocardiography Procedure Report | Adult Echo | MV Diam | Dist | meas | m_MV_Diam | r_Distance | unilateral | 125200\G-038F\G-03A2\T-35300 | SNM3, T-35300, Mitral Valve | SRT, G-038F, Cardiovascular Orifice Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | MV Ann Major | Dist | meas | m_MV_Ann_Major | r_Distance | unilateral | 125200\G-03A2\G-A193\T-35300\T-35313 | SNM3, T-35300, Mitral Valve | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm | SRT, T-35313, Mitral Annulus | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | MV Ann Minor | Dist | meas | m_MV_Ann_Minor | r_Distance | unilateral | 125200\G-03A2\G-A194\T-35300\T-35313 | SNM3, T-35300, Mitral Valve | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm | SRT, T-35313, Mitral Annulus | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Ao Arch Diam | Dist | meas | m_AoArch_Diam | r_Distance | unilateral | 125200\18011-7\G-03A2\T-42000 | SNM3, T-42000, Aorta | LN, 18011-7, Aortic Arch Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Asc Ao Diam | Dist | meas | m_AscAorta_Diam | r_Distance | unilateral | 125200\18012-5\G-03A2\T-42000 | SNM3, T-42000, Aorta | LN, 18012-5, Ascending Aortic Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Desc Ao Diam | Dist | meas | m_DescAorta_Diam | r_Distance | unilateral | 125200\18013-3\G-03A2\T-42000 | SNM3, T-42000, Aorta | LN, 18013-3, Descending Aortic Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Ao Isthmus Diam | Dist | meas | m_AoIsth_Diam | r_Distance | unilateral | 125200\18014-1\G-03A2\T-42000 | SNM3, T-42000, Aorta | LN, 18014-1, Aortic Isthmus Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | AoR Diam (MM) | Dist | meas | m_AoR_Diam_MM | r_Distance | unilateral | 125200\18015-8\G-0394\T-42000 | SNM3, T-42000, Aorta | LN, 18015-8, Aortic Root Diameter |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | AoR Diam (2D) | Dist | meas | m_AoR_Diam_2D | r_Distance | unilateral | 125200\18015-8\G-03A2\T-42000 | SNM3, T-42000, Aorta | LN, 18015-8, Aortic Root Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | PE to Sept Wall | Dist | meas | m_PE_to_Sept_Wall | r_Distance | unilateral | 125200\C12277-01\D3-90008\G-03A2 | SRT, D3-90008, Pericardial Effusion | 99PMSBLUS, C12277-01, Pericardial Effusion Septum Distance |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | PE to Lat Wall | Dist | meas | m_PE_to_Lat_Wall | r_Distance | unilateral | 125200\C12277-02\D3-90008\G-03A2 | SRT, D3-90008, Pericardial Effusion | 99PMSBLUS, C12277-02, Pericardial Effusion To Lateral Wall Distance |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | PE to Ant Wall | Dist | meas | m_PE_to_Ant_Wall | r_Distance | unilateral | 125200\C12277-03\D3-90008\G-03A2 | SRT, D3-90008, Pericardial Effusion | 99PMSBLUS, C12277-03, Pericardial Effusion To Anterior Distance |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | PE to Inf Wall | Dist | meas | m_PE_to_Inf_Wall | r_Distance | unilateral | 125200\C12277-04\D3-90008\G-03A2 | SRT, D3-90008, Pericardial Effusion | 99PMSBLUS, C12277-04, Pericardial Effusion To Inferior Distance |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | PE to Post Wall | Dist | meas | m_PE_to_Post_Wall | r_Distance | unilateral | 125200\C12277-05\D3-90008\G-03A2 | SRT, D3-90008, Pericardial Effusion | 99PMSBLUS, C12277-05, Pericardial Effusion To Posterior Distance |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | PE to RV | Dist | meas | m_PE_to_RV | r_Distance | unilateral | 125200\C12277-06\D3-90008\G-03A2 | SRT, D3-90008, Pericardial Effusion | 99PMSBLUS, C12277-06, Pericardial Effusion To Right Ventricle Wall Distance |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | PE to RA | Dist | meas | m_PE_to_RA | r_Distance | unilateral | 125200\C12277-07\D3-90008\G-03A2 | SRT, D3-90008, Pericardial Effusion | 99PMSBLUS, C12277-07, Pericardial Effusion To Right Atrium Wall Distance |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | VSD Vmax | Vmax | meas | m_VSD_Vmax | r_Vmax | unilateral | 125200\11726-7\D4-31150 | SRT, D4-31150, Ventricular Septal Defect | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s
Adult Echocardiography Procedure Report | Adult Echo | VSD Vmax | Max PG | meas | m_VSD_Vmax | r_MaxPG | unilateral | 125200\20247-3\D4-31150 | SRT, D4-31150, Ventricular Septal Defect | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Atrial Volume | meas | m_A2Cs_RA | r_ASimpVolume | unilateral | 125200\109070\125208\C12206-01\G-03A2\G-A19B\T-32200 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-01, Right Atrium Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Atrial Volume | meas | m_A4Cs_RA | r_ASimpVolume | unilateral | 125200\109070\125208\C12206-01\G-03A2\G-A19C\T-32200 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-01, Right Atrium Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2C-A/L | Length | meas | m_A2Cal_RA | r_Length | unilateral | 125200\125226\C12206-02\G-03A2\G-A19B\T-32200 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-02, Right Atrium Systolic Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Atrial Length | meas | m_A2Cs_RA | r_ASimpDistance | unilateral | 125200\109070\125208\C12206-02\G-03A2\G-A19B\T-32200 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-02, Right Atrium Systolic Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4C-A/L | Length | meas | m_A4Cal_RA | r_Length | unilateral | 125200\125226\C12206-02\G-03A2\G-A19C\T-32200 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-02, Right Atrium Systolic Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Atrial Length | meas | m_A4Cs_RA | r_ASimpDistance | unilateral | 125200\109070\125208\C12206-02\G-03A2\G-A19C\T-32200 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-02, Right Atrium Systolic Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 1 | meas | m_A2Cs_RA | r_Seg1 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 10 | meas | m_A2Cs_RA | r_Seg10 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 11 | meas | m_A2Cs_RA | r_Seg11 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 12 | meas | m_A2Cs_RA | r_Seg12 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 13 | meas | m_A2Cs_RA | r_Seg13 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 14 | meas | m_A2Cs_RA | r_Seg14 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 15 | meas | m_A2Cs_RA | r_Seg15 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 16 | meas | m_A2Cs_RA | r_Seg16 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 17 | meas | m_A2Cs_RA | r_Seg17 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 18 | meas | m_A2Cs_RA | r_Seg18 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 19 | meas | m_A2Cs_RA | r_Seg19 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 2 | meas | m_A2Cs_RA | r_Seg2 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 20 | meas | m_A2Cs_RA | r_Seg20 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 3 | meas | m_A2Cs_RA | r_Seg3 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 4 | meas | m_A2Cs_RA | r_Seg4 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 5 | meas | m_A2Cs_RA | r_Seg5 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 6 | meas | m_A2Cs_RA | r_Seg6 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 7 | meas | m_A2Cs_RA | r_Seg7 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 8 | meas | m_A2Cs_RA | r_Seg8 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Seg 9 | meas | m_A2Cs_RA | r_Seg9 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 1 | meas | m_A4Cs_RA | r_Seg1 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 10 | meas | m_A4Cs_RA | r_Seg10 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 11 | meas | m_A4Cs_RA | r_Seg11 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 12 | meas | m_A4Cs_RA | r_Seg12 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 13 | meas | m_A4Cs_RA | r_Seg13 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 14 | meas | m_A4Cs_RA | r_Seg14 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 15 | meas | m_A4Cs_RA | r_Seg15 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 16 | meas | m_A4Cs_RA | r_Seg16 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 17 | meas | m_A4Cs_RA | r_Seg17 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 18 | meas | m_A4Cs_RA | r_Seg18 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 19 | meas | m_A4Cs_RA | r_Seg19 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 2 | meas | m_A4Cs_RA | r_Seg2 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 20 | meas | m_A4Cs_RA | r_Seg20 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 3 | meas | m_A4Cs_RA | r_Seg3 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 4 | meas | m_A4Cs_RA | r_Seg4 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 5 | meas | m_A4Cs_RA | r_Seg5 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 6 | meas | m_A4Cs_RA | r_Seg6 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 7 | meas | m_A4Cs_RA | r_Seg7 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 8 | meas | m_A4Cs_RA | r_Seg8 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Seg 9 | meas | m_A4Cs_RA | r_Seg9 | unilateral | 125200\109070\125208\C12206-04\G-03A2\G-A19C\T-32200\T5203-01 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-04, Right Atrium MOD Diam |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A2C-A/L | Area | meas | m_A2Cal_RA | r_Area | unilateral | 125200\125226\17988-7\G-03A2\G-A19B\T-32200 | SRT, T-32200, Right Atrium | LN, 17988-7, Right Atrium Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | RA A2Cs | Atrial Area | meas | m_A2Cs_RA | r_ASimpArea | unilateral | 125200\109070\125208\17988-7\G-03A2\G-A19B\T-32200 | SRT, T-32200, Right Atrium | LN, 17988-7, Right Atrium Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA A4C-A/L | Area | meas | m_A4Cal_RA | r_Area | unilateral | 125200\125226\17988-7\G-03A2\G-A19C\T-32200 | SRT, T-32200, Right Atrium | LN, 17988-7, Right Atrium Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | RA ESA | Area | meas | m_RA_ESA | r_Area | unilateral | 125200\109070\17988-7\G-03A2\G-A19C\T-32200 | SRT, T-32200, Right Atrium | LN, 17988-7, Right Atrium Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | RA A4Cs | Atrial Area | meas | m_A4Cs_RA | r_ASimpArea | unilateral | 125200\109070\125208\17988-7\G-03A2\G-A19C\T-32200 | SRT, T-32200, Right Atrium | LN, 17988-7, Right Atrium Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA Pressure | Pressure | meas | m_RA_Press | r_UserPressure | unilateral | 125200\18070-3\T-32200 | SRT, T-32200, Right Atrium | LN, 18070-3, Right Atrium Systolic Pressure |  |  | UCUM, mmHg, mmHg
Adult Echocardiography Procedure Report | Adult Echo | RA Length | Dist | meas | m_RA_Length | r_Distance | unilateral | 125200\G-03A2\G-A193\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | RA Width | Dist | meas | m_RA_Width | r_Distance | unilateral | 125200\G-03A2\G-A194\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Atrial Length | meas | m_A2Cs_LA | r_ASimpDistance | unilateral | 125200\125208\C12205-01\G-03A2\G-A19B\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-01, Left Atrium systolic major axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LA A2C-A/L | Length | meas | m_A2Cal_LA | r_Length | unilateral | 125200\125226\C12205-01\G-03A2\G-A19B\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-01, Left Atrium systolic major axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Atrial Length | meas | m_A4Cs_LA | r_ASimpDistance | unilateral | 125200\125208\C12205-01\G-03A2\G-A19C\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-01, Left Atrium systolic major axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LA A4C-A/L | Length | meas | m_A4Cal_LA | r_Length | unilateral | 125200\125226\C12205-01\G-03A2\G-A19C\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-01, Left Atrium systolic major axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 1 | meas | m_A2Cs_LA | r_Seg1 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 10 | meas | m_A2Cs_LA | r_Seg10 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 11 | meas | m_A2Cs_LA | r_Seg11 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 12 | meas | m_A2Cs_LA | r_Seg12 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 13 | meas | m_A2Cs_LA | r_Seg13 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 14 | meas | m_A2Cs_LA | r_Seg14 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 15 | meas | m_A2Cs_LA | r_Seg15 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 16 | meas | m_A2Cs_LA | r_Seg16 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 17 | meas | m_A2Cs_LA | r_Seg17 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 18 | meas | m_A2Cs_LA | r_Seg18 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 19 | meas | m_A2Cs_LA | r_Seg19 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 2 | meas | m_A2Cs_LA | r_Seg2 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 20 | meas | m_A2Cs_LA | r_Seg20 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 3 | meas | m_A2Cs_LA | r_Seg3 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 4 | meas | m_A2Cs_LA | r_Seg4 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 5 | meas | m_A2Cs_LA | r_Seg5 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 6 | meas | m_A2Cs_LA | r_Seg6 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 7 | meas | m_A2Cs_LA | r_Seg7 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 8 | meas | m_A2Cs_LA | r_Seg8 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Seg 9 | meas | m_A2Cs_LA | r_Seg9 | unilateral | 125200\109070\C12205-02\G-A19B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 1 | meas | m_A4Cs_LA | r_Seg1 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 10 | meas | m_A4Cs_LA | r_Seg10 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 11 | meas | m_A4Cs_LA | r_Seg11 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 12 | meas | m_A4Cs_LA | r_Seg12 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 13 | meas | m_A4Cs_LA | r_Seg13 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 14 | meas | m_A4Cs_LA | r_Seg14 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 15 | meas | m_A4Cs_LA | r_Seg15 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 16 | meas | m_A4Cs_LA | r_Seg16 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 17 | meas | m_A4Cs_LA | r_Seg17 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 18 | meas | m_A4Cs_LA | r_Seg18 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 19 | meas | m_A4Cs_LA | r_Seg19 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 2 | meas | m_A4Cs_LA | r_Seg2 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 20 | meas | m_A4Cs_LA | r_Seg20 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 3 | meas | m_A4Cs_LA | r_Seg3 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 4 | meas | m_A4Cs_LA | r_Seg4 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 5 | meas | m_A4Cs_LA | r_Seg5 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 6 | meas | m_A4Cs_LA | r_Seg6 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 7 | meas | m_A4Cs_LA | r_Seg7 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 8 | meas | m_A4Cs_LA | r_Seg8 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Seg 9 | meas | m_A4Cs_LA | r_Seg9 | unilateral | 125200\109070\C12205-02\G-A19C\T-32300\T5203-01 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LA Volume (HM) | LA Vmax (HM) | meas | m_HM_LA_Vol | r_HM_ESV_LA | unilateral | 125200\C12205-05\C12228-11\T12224-01\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-05, Left Atrium Maximum Volume |  |  | UCUM, mm3, mm3 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | LA EF BP-a2DQ | Max Vol | meas | m_2DQ_EF_BP_LA | r_MaxVol | unilateral | 125200\C12205-05\C12228-09\C99-12228-05\G-03A2\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-05, Left Atrium Maximum Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LA EF AP2-a2DQ | Max Vol | meas | m_2DQ_EF_Ap2_LA | r_MaxVol | unilateral | 125200\C12205-05\C12228-09\C99-12228-05\G-03A2\G-A19B\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-05, Left Atrium Maximum Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LA EF AP4-a2DQ | Max Vol | meas | m_2DQ_EF_Ap4_LA | r_MaxVol | unilateral | 125200\C12205-05\C12228-09\C99-12228-05\G-03A2\G-A19C\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-05, Left Atrium Maximum Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LA Volume (HM) | LA Vmin (HM) | meas | m_HM_LA_Vol | r_HM_EDV_LA | unilateral | 125200\C12205-06\C12228-11\T12224-01\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-06, Left Atrium Minimum Volume |  |  | UCUM, mm3, mm3 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | LA EF BP-a2DQ | Min Vol | meas | m_2DQ_EF_BP_LA | r_MinVol | unilateral | 125200\C12205-06\C12228-09\C99-12228-05\G-03A2\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-06, Left Atrium Minimum Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LA EF AP2-a2DQ | Min Vol | meas | m_2DQ_EF_Ap2_LA | r_MinVol | unilateral | 125200\C12205-06\C12228-09\C99-12228-05\G-03A2\G-A19B\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-06, Left Atrium Minimum Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LA EF AP4-a2DQ | Min Vol | meas | m_2DQ_EF_Ap4_LA | r_MinVol | unilateral | 125200\C12205-06\C12228-09\C99-12228-05\G-03A2\G-A19C\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-06, Left Atrium Minimum Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LA Volume (HM) | LA EF (HM) | meas | m_HM_LA_Vol | r_HM_LA_EF | unilateral | 125200\C12205-07\C12228-11\T12224-01\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-07, Left Atrium Ejection Fraction |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | LA EF BP-a2DQ | EF | meas | m_2DQ_EF_BP_LA | r_EF | unilateral | 125200\C12205-07\C12228-09\C99-12228-05\G-03A2\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-07, Left Atrium Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LA EF AP2-a2DQ | EF | meas | m_2DQ_EF_Ap2_LA | r_EF | unilateral | 125200\C12205-07\C12228-09\C99-12228-05\G-03A2\G-A19B\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-07, Left Atrium Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LA EF AP4-a2DQ | EF | meas | m_2DQ_EF_Ap4_LA | r_EF | unilateral | 125200\C12205-07\C12228-09\C99-12228-05\G-03A2\G-A19C\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-07, Left Atrium Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LA FAC AP2-a2DQ | Max Area | meas | m_2DQ_FAC_Ap2_LA | r_MaxA | unilateral | 125200\C12205-08\C12228-09\C99-12228-05\G-03A2\G-A19B\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-08, Left Atrium Maximum Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LA FAC AP4-a2DQ | Max Area | meas | m_2DQ_FAC_Ap4_LA | r_MaxA | unilateral | 125200\C12205-08\C12228-09\C99-12228-05\G-03A2\G-A19C\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-08, Left Atrium Maximum Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LA FAC AP2-a2DQ | Min Area | meas | m_2DQ_FAC_Ap2_LA | r_MinA | unilateral | 125200\C12205-09\C12228-09\C99-12228-05\G-03A2\G-A19B\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-09, Left Atrium Minimum Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LA FAC AP4-a2DQ | Min Area | meas | m_2DQ_FAC_Ap4_LA | r_MinA | unilateral | 125200\C12205-09\C12228-09\C99-12228-05\G-03A2\G-A19C\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-09, Left Atrium Minimum Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LA FAC AP2-a2DQ | FAC | meas | m_2DQ_FAC_Ap2_LA | r_FAC | unilateral | 125200\C12205-10\C12228-09\C99-12228-05\G-03A2\G-A19B\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-10, Left Atrium Fractional Area Change |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LA FAC AP4-a2DQ | FAC | meas | m_2DQ_FAC_Ap4_LA | r_FAC | unilateral | 125200\C12205-10\C12228-09\C99-12228-05\G-03A2\G-A19C\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-10, Left Atrium Fractional Area Change |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | LA Volume (HM) | LA VI (HM) | meas | m_HM_LA_Vol | r_HM_Vol_Index_LA | unilateral | 125200\C12205-11\C12228-11\T12224-01\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-11, Left Atrium Maximum Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | LA Volume (HM) | avg LA Vmax (HM) | meas | m_HM_LA_Vol | r_HM_ESV_LA_AVG | unilateral | 125200\C12205-12\C12228-11\T12224-01\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-12, Left Atrium MultiBeat Maximum Volume |  |  | UCUM, mm3, mm3 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | LA Volume (HM) | avg LA Vmin (HM) | meas | m_HM_LA_Vol | r_HM_EDV_LA_AVG | unilateral | 125200\C12205-13\C12228-11\T12224-01\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-13, Left Atrium MultiBeat Minimum Volume |  |  | UCUM, mm3, mm3 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | LA Volume (HM) | avg LA VI (HM) | meas | m_HM_LA_Vol | r_HM_LA_VI_AVG | unilateral | 125200\C12205-14\C12228-11\T12224-01\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-14, Left Atrium MultiBeat Maximum Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | LA Volume (HM) | avg LA EF (HM) | meas | m_HM_LA_Vol | r_HM_LA_EF_AVG | unilateral | 125200\C12205-15\C12228-11\T12224-01\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-15, Left Atrium MultiBeat Ejection Fraction |  |  | UCUM, %, Percent |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, C12228-11, QLab HM
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Atrial Area | meas | m_A2Cs_LA | r_ASimpArea | unilateral | 125200\125208\17977-0\G-03A2\G-A19B\T-32300 | SRT, T-32300, Left Atrium | LN, 17977-0, Left Atrium Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LA A2C-A/L | Area | meas | m_A2Cal_LA | r_Area | unilateral | 125200\125226\17977-0\G-03A2\G-A19B\T-32300 | SRT, T-32300, Left Atrium | LN, 17977-0, Left Atrium Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Atrial Area | meas | m_A4Cs_LA | r_ASimpArea | unilateral | 125200\125208\17977-0\G-03A2\G-A19C\T-32300 | SRT, T-32300, Left Atrium | LN, 17977-0, Left Atrium Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LA A4C-A/L | Area | meas | m_A4Cal_LA | r_Area | unilateral | 125200\125226\17977-0\G-03A2\G-A19C\T-32300 | SRT, T-32300, Left Atrium | LN, 17977-0, Left Atrium Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | LA Dimen (MM) | Dist | meas | m_LA_Dimen_MM | r_Distance | unilateral | 125200\29469-4\G-0394\T-32300 | SRT, T-32300, Left Atrium | LN, 29469-4, Left Atrium Antero-posterior Systolic Dimension |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LA Dimen (2D) | Dist | meas | m_LA_Dimen_2D | r_Distance | unilateral | 125200\29469-4\G-03A2\T-32300 | SRT, T-32300, Left Atrium | LN, 29469-4, Left Atrium Antero-posterior Systolic Dimension |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LA Area | Area | meas | m_LA_Area | r_Area | unilateral | 125200\G-03A2\G-A166\T-32300 | SRT, T-32300, Left Atrium | SNM3, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LA A2Cs | Atrial Volume | meas | m_A2Cs_LA | r_ASimpVolume | unilateral | 125200\125208\G-0383\G-03A2\G-A19B\T-32300 | SRT, T-32300, Left Atrium | SRT, G-0383, Left Atrium Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LA A4Cs | Atrial Volume | meas | m_A4Cs_LA | r_ASimpVolume | unilateral | 125200\125208\G-0383\G-03A2\G-A19C\T-32300 | SRT, T-32300, Left Atrium | SRT, G-0383, Left Atrium Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LA Length | Dist | meas | m_LA_Length | r_Distance | unilateral | 125200\G-03A2\G-A193\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LA Width | Dist | meas | m_LA_Width | r_Distance | unilateral | 125200\G-03A2\G-A194\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | IV Septum Wall | Dist | meas | m_IV_Septum_Wall | r_Distance | unilateral | 125200\C12201-09\G-03A2\T-32410 | SRT, T-32410, Interventricular septum | 99PMSBLUS, C12201-09, Interventricular Septum Wall Thickness |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | RV EDA | Area | meas | m_RV_EDA | r_Area | unilateral | 125200\C12204-01\F-32011\G-03A2\G-A19C\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-01, Right Ventricular Diastolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | RV ESA | Area | meas | m_RV_ESA | r_Area | unilateral | 125200\109070\C12204-02\G-03A2\G-A19C\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-02, Right Ventricular Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Adult Echocardiography Procedure Report | Adult Echo | RV Wall | Dist | meas | m_RV_Wall | r_Distance | unilateral | 125200\C12204-08\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-08, Right Ventricular Wall Thickness |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | RV S Vel | Vel | meas | m_RV_S_Vel | r_Vel | unilateral | 125200\C12204-10\P5-B0128\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-10, Right Ventricle S Velocity |  |  | UCUM, mm/s, mm/s |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | RV E€™ Vel | Vel | meas | m_RV_E_Vel | r_Vel | unilateral | 125200\C12204-11\P5-B0128\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-11, Right Ventricle E Velocity |  |  | UCUM, mm/s, mm/s |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | RV A€™ Vel | Vel | meas | m_RV_A_Vel | r_Vel | unilateral | 125200\C12204-12\P5-B0128\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-12, Right Ventricle A Velocity |  |  | UCUM, mm/s, mm/s |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Time to RV S | Time | meas | m_Time_to_RV_S | r_DopTime | unilateral | 125200\C12204-13\P5-B0128\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-13, Right Ventricle Time To S |  |  | UCUM, ms, ms |  | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | RV Base | Dist | meas | m_RV_Base | r_Distance | unilateral | 125200\C12204-14\G-03A2\G-A19C\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-14, Right Ventricle Basal Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber
Adult Echocardiography Procedure Report | Adult Echo | RV Mid | Dist | meas | m_RV_Mid | r_Distance | unilateral | 125200\C12204-15\G-03A2\G-A19C\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-15, Right Ventricle Mid Cavity Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber
Adult Echocardiography Procedure Report | Adult Echo | RVOT Distal | Dist | meas | m_RVOT_Distal | r_Distance | unilateral | 125200\C12204-16\G-03A2\T-32500\T-32550 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-16, Right Ventricle Outflow Tract Distal Diameter |  |  | UCUM, mm, mm | SNM3, T-32550, Right Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | RVOT Prox | Dist | meas | m_RVOT_Prox | r_Distance | unilateral | 125200\C12204-17\G-03A2\T-32500\T-32550 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-17, Right Ventricle Outflow Tract Proximal Diameter |  |  | UCUM, mm, mm | SNM3, T-32550, Right Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | RVET (DOP) | Time | meas | m_RVET_DOP | r_DopTime | unilateral | 125200\122213\T-32500 | SRT, T-32500, Right Ventricle | DCM, 122213, Right Ventricular ejection time |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | RVOT Vmax | Vmax | meas | m_RVOT_Vmax | r_Vmax | unilateral | 125200\11726-7\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s | SNM3, T-32550, Right Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | RV dP/dt | dP/dt | meas | m_RV_DpDt | r_DopDpDt | unilateral | 125200\18034-9\T-32500 | SRT, T-32500, Right Ventricle | LN, 18034-9, Tricuspid Regurgitation dP/dt |  |  | UCUM, mmHg, mmHg
Adult Echocardiography Procedure Report | Adult Echo | RVAWd (MM) | Dist | meas | m_RVAWd_MM | r_Distance | unilateral | 125200\18153-7\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 18153-7, Right Ventricular Anterior Wall Diastolic Thickness |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | RVAWd (2D) | Dist | meas | m_RVAWd_2D | r_Distance | unilateral | 125200\18153-7\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | LN, 18153-7, Right Ventricular Anterior Wall Diastolic Thickness |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | RV ET | Time | meas | m_RV_Eject_Tm | r_MMTime | unilateral | 125200\20222-6\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 20222-6, Ejection Time |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | RVOT Vmax | Max PG | meas | m_RVOT_Vmax | r_MaxPG | unilateral | 125200\20247-3\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SNM3, T-32550, Right Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | RVOT VTI | Mean PG | meas | m_RVOT_VTI | r_CardPGmean | unilateral | 125200\20256-4\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg | SNM3, T-32550, Right Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | RVIDd (MM) | Dist | meas | m_RVIDd_MM | r_Distance | unilateral | 125200\20304-2\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 20304-2, Right Ventricular Internal Diastolic Dimension |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | RVIDd (2D) | Dist | meas | m_RVIDd_2D | r_Distance | unilateral | 125200\20304-2\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | LN, 20304-2, Right Ventricular Internal Diastolic Dimension |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | RVOT VTI | Vmean | meas | m_RVOT_VTI | r_CardTAPV | unilateral | 125200\20352-1\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s | SNM3, T-32550, Right Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | RVOT VTI | VTI | meas | m_RVOT_VTI | r_VTI | unilateral | 125200\20354-7\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm | SNM3, T-32550, Right Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | RV PEP | Time | meas | m_RV_PreEject_Tm | r_MMTime | unilateral | 125200\59085-1\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 59085-1, Pre-Ejection Period |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | RVOT Diam | Dist | meas | m_RVOT_Diam | r_Distance | unilateral | 125200\G-038F\G-03A2\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, G-038F, Cardiovascular Orifice Diameter |  |  | UCUM, mm, mm | SNM3, T-32550, Right Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | RV Length | Dist | meas | m_RV_Length | r_Distance | unilateral | 125200\G-03A2\G-A193\G-A19C\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber
Adult Echocardiography Procedure Report | Adult Echo | TV D-E Exc | Dist | meas | m_TV_DE_Exc | r_Distance | unilateral | 125200\C12208-01\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-01, Tricuspid Valve D-E Excursion |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | TV TMAD AP2-a2DQ | Pt1 | meas | m_2DQ_TMAD_Ap2_TV | r_Point1 | unilateral | 125200\C12208-010\C12228-09\C99-12228-06\G-03A2\G-A19B\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-010, Displacement Left Tricuspid Annular Hingepoint to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TV TMAD AP2-aCMQ | Pt1 | meas | m_CMQ_TMAD_Ap2_TV | r_Point1 | unilateral | 125200\C12208-010\C12228-10\C99-12228-06\G-03A2\G-A19B\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-010, Displacement Left Tricuspid Annular Hingepoint to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TV TMAD AP4-a2DQ | Pt1 | meas | m_2DQ_TMAD_Ap4_TV | r_Point1 | unilateral | 125200\C12208-010\C12228-09\C99-12228-06\G-03A2\G-A19C\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-010, Displacement Left Tricuspid Annular Hingepoint to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TV TMAD AP4-aCMQ | Pt1 | meas | m_CMQ_TMAD_Ap4_TV | r_Point1 | unilateral | 125200\C12208-010\C12228-10\C99-12228-06\G-03A2\G-A19C\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-010, Displacement Left Tricuspid Annular Hingepoint to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TV D-E Slope | Slope | meas | m_TV_DE_Slope | r_MMSlope | unilateral | 125200\C12208-02\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-02, Tricuspid Valve D-E Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | TV E-F Slope | Slope | meas | m_TV_EF_Slope | r_MMSlope | unilateral | 125200\C12208-03\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-03, Tricuspid Valve E-F Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | TV A-C Interval | Time | meas | m_TV_AC_Int | r_MMTime | unilateral | 125200\C12208-04\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-04, Tricuspid Valve A-C Interval |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | TV TMAD AP2-a2DQ | Pt2 | meas | m_2DQ_TMAD_Ap2_TV | r_Point2 | unilateral | 125200\C12208-07\C12228-09\C99-12228-06\G-03A2\G-A19B\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-07, Displacement Right Tricuspid Annular Hingepoint to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TV TMAD AP2-aCMQ | Pt2 | meas | m_CMQ_TMAD_Ap2_TV | r_Point2 | unilateral | 125200\C12208-07\C12228-10\C99-12228-06\G-03A2\G-A19B\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-07, Displacement Right Tricuspid Annular Hingepoint to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TV TMAD AP4-a2DQ | Pt2 | meas | m_2DQ_TMAD_Ap4_TV | r_Point2 | unilateral | 125200\C12208-07\C12228-09\C99-12228-06\G-03A2\G-A19C\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-07, Displacement Right Tricuspid Annular Hingepoint to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TV TMAD AP4-aCMQ | Pt2 | meas | m_CMQ_TMAD_Ap4_TV | r_Point2 | unilateral | 125200\C12208-07\C12228-10\C99-12228-06\G-03A2\G-A19C\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-07, Displacement Right Tricuspid Annular Hingepoint to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TV TMAD AP2-a2DQ | MidPt Disp | meas | m_2DQ_TMAD_Ap2_TV | r_MidPoint_Disp | unilateral | 125200\C12208-08\C12228-09\C99-12228-06\G-03A2\G-A19B\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-08, Displacement Midpoint of L/R TV Annular Hingepoints to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TV TMAD AP2-aCMQ | MidPt Disp | meas | m_CMQ_TMAD_Ap2_TV | r_MidPoint_Disp | unilateral | 125200\C12208-08\C12228-10\C99-12228-06\G-03A2\G-A19B\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-08, Displacement Midpoint of L/R TV Annular Hingepoints to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TV TMAD AP4-a2DQ | MidPt Disp | meas | m_2DQ_TMAD_Ap4_TV | r_MidPoint_Disp | unilateral | 125200\C12208-08\C12228-09\C99-12228-06\G-03A2\G-A19C\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-08, Displacement Midpoint of L/R TV Annular Hingepoints to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TV TMAD AP4-aCMQ | MidPt Disp | meas | m_CMQ_TMAD_Ap4_TV | r_MidPoint_Disp | unilateral | 125200\C12208-08\C12228-10\C99-12228-06\G-03A2\G-A19C\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-08, Displacement Midpoint of L/R TV Annular Hingepoints to Apex |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TV TMAD AP2-a2DQ | MidPt Norm | meas | m_2DQ_TMAD_Ap2_TV | r_MidPoint_Norm | unilateral | 125200\C12208-09\C12228-09\C99-12228-06\G-03A2\G-A19B\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-09, Normalized Midpoint of L/R TV Annular Hingepoints to Apex |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TV TMAD AP2-aCMQ | MidPt Norm | meas | m_CMQ_TMAD_Ap2_TV | r_MidPoint_Norm | unilateral | 125200\C12208-09\C12228-10\C99-12228-06\G-03A2\G-A19B\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-09, Normalized Midpoint of L/R TV Annular Hingepoints to Apex |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TV TMAD AP4-a2DQ | MidPt Norm | meas | m_2DQ_TMAD_Ap4_TV | r_MidPoint_Norm | unilateral | 125200\C12208-09\C12228-09\C99-12228-06\G-03A2\G-A19C\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-09, Normalized Midpoint of L/R TV Annular Hingepoints to Apex |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TV TMAD AP4-aCMQ | MidPt Norm | meas | m_CMQ_TMAD_Ap4_TV | r_MidPoint_Norm | unilateral | 125200\C12208-09\C12228-10\C99-12228-06\G-03A2\G-A19C\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-09, Normalized Midpoint of L/R TV Annular Hingepoints to Apex |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TAPSE | Dist | meas | m_TAPSE | r_Distance | unilateral | 125200\C12208-11\G-0394\T-35100\T-35111 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-11, Tricuspid Annular Plane Systolic Excursion |  |  | UCUM, mm, mm | SRT, T-35111, Tricuspid Annulus | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | TV Vmax | Vmax | meas | m_TV_Vmax | r_Vmax | unilateral | 125200\11726-7\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TR Vmax | Vmax | meas | m_TR_Vmax | r_Vmax | unilateral | 125200\11726-7\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TV Peak A Vel | Vel | meas | m_TV_PeakAVel | r_Vel | unilateral | 125200\18030-7\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 18030-7, Tricuspid Valve A Wave Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TV Peak E Vel | Vel | meas | m_TV_PeakEVel | r_Vel | unilateral | 125200\18031-5\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 18031-5, Tricuspid Valve E Wave Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TV Accel Time | Slope | meas | m_TV_AccelTime | r_DopSlope | unilateral | 125200\20167-3\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20167-3, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TV Accel Time | Time | meas | m_TV_AccelTime | r_DopTime | unilateral | 125200\20168-1\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TV Vmax | Max PG | meas | m_TV_Vmax | r_MaxPG | unilateral | 125200\20247-3\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TR Vmax | Max PG | meas | m_TR_Vmax | r_MaxPG | unilateral | 125200\20247-3\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TV VTI | Mean PG | meas | m_TV_VTI | r_CardPGmean | unilateral | 125200\20256-4\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TR VTI | Mean PG | meas | m_TR_VTI | r_CardPGmean | unilateral | 125200\20256-4\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TV P1/2t | P½t | meas | m_TV_PHalfT | r_PHalftime | unilateral | 125200\20280-4\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20280-4, Pressure Half-Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TV VTI | Vmean | meas | m_TV_VTI | r_CardTAPV | unilateral | 125200\20352-1\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TR VTI | Vmean | meas | m_TR_VTI | r_CardTAPV | unilateral | 125200\20352-1\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TV VTI | VTI | meas | m_TV_VTI | r_VTI | unilateral | 125200\20354-7\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TR VTI | VTI | meas | m_TR_VTI | r_VTI | unilateral | 125200\20354-7\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TV Radius | Dist | meas | m_TV_Radius | r_Distance | unilateral | 125200\59102-4\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59102-4, Flow Radius |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TR Radius | Dist | meas | m_TR_Radius | r_Distance | unilateral | 125200\59102-4\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59102-4, Flow Radius |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TR Jet Area | Area | meas | m_TR_Jet_Area | r_Area | unilateral | 125200\59123-0\G-03A2\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59123-0, Jet Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TV Alias Vel | Alias Vel | meas | m_TV_AliasVel | r_2DVel | unilateral | 125200\59130-5\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59130-5, Alias Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TR Alias Vel | Alias Vel | meas | m_TR_AliasVel | r_2DVel | unilateral | 125200\59130-5\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59130-5, Alias Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | HR - TV | HR | meas | m_HR_TV | r_HeartRate | unilateral | 125200\8867-4\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 8867-4, Heart rate |  |  | UCUM, {H.B.}/min, bpm
Adult Echocardiography Procedure Report | Adult Echo | TR VC | Dist | meas | m_TR_VC | r_Distance | unilateral | 125200\G-03A2\M-02550\R-421AA\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, R-421AA, Vena Contracta | SRT, G-03A2, 2D mode |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TV Annulus | Dist | meas | m_TV_Annulus | r_Distance | unilateral | 125200\G-03A2\G-A19C\M-02550\T-35100\T-35111 | SRT, T-35100, Tricuspid Valve | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-35111, Tricuspid Annulus | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber
Adult Echocardiography Procedure Report | Adult Echo | TCOT | Time | meas | m_TCOT | r_DopTime | unilateral | 125200\G-0389\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-0389, Tricuspid Valve Closure to Opening Time |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | TV Diam | Dist | meas | m_TV_Diam | r_Distance | unilateral | 125200\G-038F\G-03A2\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038F, Cardiovascular Orifice Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Late Dias Slope | Slope | meas | m_LateDiasSlope | r_MMSlope | unilateral | 125200\C12209-01\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | 99PMSBLUS, C12209-01, Late Diastolic Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | PI End Dias Vel | Vel | meas | m_PI_EndDiVel | r_Vel | unilateral | 125200\11653-3\R-42E61\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | PV Vmax | Vmax | meas | m_PV_Vmax | r_Vmax | unilateral | 125200\11726-7\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | PI Vmax | Vmax | meas | m_PI_Vmax | r_Vmax | unilateral | 125200\11726-7\R-42E61\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | PV Accel Time | Slope | meas | m_PV_AccelTime | r_DopSlope | unilateral | 125200\20167-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20167-3, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | PV Accel Time | Time | meas | m_PV_AccelTime | r_DopTime | unilateral | 125200\20168-1\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | PV Vmax | Max PG | meas | m_PV_Vmax | r_MaxPG | unilateral | 125200\20247-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | PI Vmax | Max PG | meas | m_PI_Vmax | r_MaxPG | unilateral | 125200\20247-3\R-42E61\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | PI End Dias Vel | PG | meas | m_PI_EndDiVel | r_PG | unilateral | 125200\20247-3\F-32011\R-42E61\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  | SRT, F-32011, End Diastole | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | PV VTI | Mean PG | meas | m_PV_VTI | r_CardPGmean | unilateral | 125200\20256-4\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | PR P1/2t | P½t | meas | m_PR_PHalfT | r_PHalftime | unilateral | 125200\20280-4\R-42E61\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20280-4, Pressure Half-Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | PV VTI | Vmean | meas | m_PV_VTI | r_CardTAPV | unilateral | 125200\20352-1\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | PV VTI | VTI | meas | m_PV_VTI | r_VTI | unilateral | 125200\20354-7\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | A Wave Amp | Dist | meas | m_AWaveAmp | r_Distance | unilateral | 125200\59100-8\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 59100-8, A-Wave Amplitude |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | B-C Slope | Slope | meas | m_BC_Slope | r_MMSlope | unilateral | 125200\59126-3\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 59126-3, B-C Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | HR - PV | HR | meas | m_HR_PV | r_HeartRate | unilateral | 125200\8867-4\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 8867-4, Heart rate |  |  | UCUM, {H.B.}/min, bpm
Adult Echocardiography Procedure Report | Adult Echo | PVA Ann Area | Area | meas | m_PVA_Ann_Area | r_Area | unilateral | 125200\C12209-04\G-03A2\G-A166\T-35200 | SRT, T-35200, Pulmonic Valve | SNM3, G-A166, Area |  |  | UCUM, mm2, mm2 | 99PMSBLUS, C12209-04, Pulmonic Annulus | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | PR VC | Dist | meas | m_PR_VC | r_Distance | unilateral | 125200\G-03A2\M-02550\R-421AA\R-42E61\T-35200 | SRT, T-35200, Pulmonic Valve | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, R-421AA, Vena Contracta | SRT, G-03A2, 2D mode |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | R to AV Open | Time | meas | m_RtoAVo_DopTime | r_DopTime | unilateral | 125200\C12211-06\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12211-06, R Wave to Aortic Valve Opening Time |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | R to AV Closure | Time | meas | m_RtoAVc_DopTime | r_DopTime | unilateral | 125200\C12211-07\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12211-07, R Wave to Aortic Valve Closure Time |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | Ao Ann-LCA Dist | Dist | meas | m_Ao_Ann_LCA_Dist | r_Distance | unilateral | 125200\C12211-09\G-03A2\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12211-09, Aortic Annulus to Left coronary artery orifice distance |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Ao Ann-RCA Dist | Dist | meas | m_Ao_Ann_RCA_Dist | r_Distance | unilateral | 125200\C12211-10\G-03A2\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12211-10, Aortic Annulus to Right coronary artery orifice distance |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Ao Ann-STJ Dist | Dist | meas | m_Ao_Ann_STJ_Dist | r_Distance | unilateral | 125200\C12211-19\G-03A2\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12211-19, Aortic Annulus to Sino-tubular Junction Distance |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | AV R-R Interval | Time | meas | m_AV_RR_Interval | r_DopTime | unilateral | 125200\122182\T-35400 | SRT, T-35400, Aortic Valve | DCM, 122182, R-R interval |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | AI End Dias Vel | Vel | meas | m_AI_EndDiVel | r_Vel | unilateral | 125200\11653-3\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AV Vmax | Vmax | meas | m_AV_Vmax | r_Vmax | unilateral | 125200\11726-7\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Vmax | Vmax | meas | m_AI_Vmax | r_Vmax | unilateral | 125200\11726-7\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AV Cusp Sep | Dist | meas | m_AV_Cusp_Sep | r_Distance | unilateral | 125200\17996-0\G-0394\T-35400 | SRT, T-35400, Aortic Valve | LN, 17996-0, Aortic Valve Cusp Separation |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | AV Accel Time | Slope | meas | m_AV_AccelTime | r_DopSlope | unilateral | 125200\20167-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20167-3, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Accel Time | Slope | meas | m_AI_AccelTime | r_DopSlope | unilateral | 125200\20167-3\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20167-3, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AV Accel Time | Time | meas | m_AV_AccelTime | r_DopTime | unilateral | 125200\20168-1\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Accel Time | Time | meas | m_AI_AccelTime | r_DopTime | unilateral | 125200\20168-1\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Dec Slope | Slope | meas | m_AI_DecSlope | r_DopSlope | unilateral | 125200\20216-8\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20216-8, Deceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AV Decel Time | Time | meas | m_AV_DecelTime | r_DopTime | unilateral | 125200\20217-6\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Dec Slope | Time | meas | m_AI_DecSlope | r_DopTime | unilateral | 125200\20217-6\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AV Vmax | Max PG | meas | m_AV_Vmax | r_MaxPG | unilateral | 125200\20247-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Vmax | Max PG | meas | m_AI_Vmax | r_MaxPG | unilateral | 125200\20247-3\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AV VTI | Mean PG | meas | m_AV_VTI | r_CardPGmean | unilateral | 125200\20256-4\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AI VTI | Mean PG | meas | m_AI_VTI | r_CardPGmean | unilateral | 125200\20256-4\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI P½t | P½t | meas | m_AI_PHalfT | r_PHalftime | unilateral | 125200\20280-4\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20280-4, Pressure Half-Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AV VTI | Vmean | meas | m_AV_VTI | r_CardTAPV | unilateral | 125200\20352-1\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AI VTI | Vmean | meas | m_AI_VTI | r_CardTAPV | unilateral | 125200\20352-1\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AV VTI | VTI | meas | m_AV_VTI | r_VTI | unilateral | 125200\20354-7\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AI VTI | VTI | meas | m_AI_VTI | r_VTI | unilateral | 125200\20354-7\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Radius | Dist | meas | m_AI_Radius | r_Distance | unilateral | 125200\59102-4\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 59102-4, Flow Radius |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Alias Vel | Alias Vel | meas | m_AI_AliasVel | r_2DVel | unilateral | 125200\59130-5\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 59130-5, Alias Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | HR - AV | HR | meas | m_HR_AV | r_HeartRate | unilateral | 125200\8867-4\T-35400 | SRT, T-35400, Aortic Valve | LN, 8867-4, Heart rate |  |  | UCUM, {H.B.}/min, bpm
Adult Echocardiography Procedure Report | Adult Echo | Ao Ann Area | Area | meas | m_Ao_Ann_Area | r_Area | unilateral | 125200\G-03A2\G-A166\T-35400\T-35410 | SRT, T-35400, Aortic Valve | SNM3, G-A166, Area |  |  | UCUM, mm2, mm2 | SRT, T-35410, Aortic Valve Ring | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | AR Diam(MM) | Dist | meas | m_AR_Diam_MM | r_Distance | unilateral | 125200\G-0394\M-02550\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AR VC | Dist | meas | m_AR_VC | r_Distance | unilateral | 125200\G-03A2\M-02550\R-421AA\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, R-421AA, Vena Contracta | SRT, G-03A2, 2D mode |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | Ao Ann Diam | Dist | meas | m_Ao_Ann_Diam | r_Distance | unilateral | 125200\G-03A2\M-02550\T-35400\T-35410 | SRT, T-35400, Aortic Valve | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-35410, Aortic Valve Ring | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | AV Area | Area | meas | m_AV_Area | r_Area | unilateral | 125200\125220\G-038E\G-03A2\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125220, Planimetry
Adult Echocardiography Procedure Report | Adult Echo | Ao Ann Major | Dist | meas | m_Ao_Ann_Major | r_Distance | unilateral | 125200\G-03A2\G-A193\T-35400\T-35410 | SRT, T-35400, Aortic Valve | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm | SRT, T-35410, Aortic Valve Ring | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Ao Ann Minor | Dist | meas | m_Ao_Ann_Minor | r_Distance | unilateral | 125200\G-03A2\G-A194\T-35400\T-35410 | SRT, T-35400, Aortic Valve | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm | SRT, T-35410, Aortic Valve Ring | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Desc Ao Vmax | Vmax | meas | m_Desc_Ao_Vmax | r_Vmax | unilateral | 125200\11726-7\R-42047\T-42070 | SRT, T-42070, Thoracic aorta | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | Desc Ao Vmax | Max PG | meas | m_Desc_Ao_Vmax | r_MaxPG | unilateral | 125200\20247-3\R-42047\T-42070 | SRT, T-42070, Thoracic aorta | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | Prox ascAo Diam | Dist | meas | m_Prox_ascAo_Diam | r_Distance | unilateral | 125200\G-03A2\G-A118\M-02550\T-42100 | SRT, T-42100, Ascending aorta | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  |  |  |  | SRT, G-A118, Proximal
Adult Echocardiography Procedure Report | Adult Echo | Ao STJ Area | Area | meas | m_Ao_STJ_Area | r_Area | unilateral | 125200\G-03A2\G-A166\T-42102 | SRT, T-42102, Aortic Sinotubular Junction | SNM3, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Ao STJ Diam | Dist | meas | m_Ao_STJ_Diam | r_Distance | unilateral | 125200\G-03A2\M-02550\T-42102 | SRT, T-42102, Aortic Sinotubular Junction | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Ao STJ Major | Dist | meas | m_Ao_STJ_Major | r_Distance | unilateral | 125200\G-03A2\G-A193\T-42102 | SRT, T-42102, Aortic Sinotubular Junction | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Ao STJ Minor | Dist | meas | m_Ao_STJ_Minor | r_Distance | unilateral | 125200\G-03A2\G-A194\T-42102 | SRT, T-42102, Aortic Sinotubular Junction | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Ao Sinus Area | Area | meas | m_Ao_Sinus_Area | r_Area | unilateral | 125200\G-03A2\G-A166\T-42200 | SRT, T-42200, Structure Sinus of Valsalva | SNM3, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Ao Sinus Diam | Dist | meas | m_Ao_Sinus_Diam | r_Distance | unilateral | 125200\G-03A2\M-02550\T-42200 | SRT, T-42200, Structure Sinus of Valsalva | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Ao Sinus Major | Dist | meas | m_Ao_Sinus_Major | r_Distance | unilateral | 125200\G-03A2\G-A193\T-42200 | SRT, T-42200, Structure Sinus of Valsalva | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Ao Sinus Minor | Dist | meas | m_Ao_Sinus_Minor | r_Distance | unilateral | 125200\G-03A2\G-A194\T-42200 | SRT, T-42200, Structure Sinus of Valsalva | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | RPA Vmax | Vmax | meas | m_RPA_Vmax | r_Vmax | unilateral | 125200\11726-7\T-44000\T-44200 | SRT, T-44000, Pulmonary artery | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, T-44200, Right Pulmonary Artery
Adult Echocardiography Procedure Report | Adult Echo | LPA Vmax | Vmax | meas | m_LPA_Vmax | r_Vmax | unilateral | 125200\11726-7\T-44000\T-44400 | SRT, T-44000, Pulmonary artery | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, T-44400, Left Pulmonary Artery
Adult Echocardiography Procedure Report | Adult Echo | LPA Diam | Dist | meas | m_LPA_Diam | r_Distance | unilateral | 125200\18019-0\G-03A2\T-44000 | SRT, T-44000, Pulmonary artery | LN, 18019-0, Left Pulmonary Artery Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | MPA Diam | Dist | meas | m_MPA_Diam | r_Distance | unilateral | 125200\18020-8\G-03A2\T-44000 | SRT, T-44000, Pulmonary artery | LN, 18020-8, Main Pulmonary Artery Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | RPA Diam | Dist | meas | m_RPA_Diam | r_Distance | unilateral | 125200\18021-6\G-03A2\T-44000 | SRT, T-44000, Pulmonary artery | LN, 18021-6, Right Pulmonary Artery Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | RPA Vmax | Max PG | meas | m_RPA_Vmax | r_MaxPG | unilateral | 125200\20247-3\T-44000\T-44200 | SRT, T-44000, Pulmonary artery | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-44200, Right Pulmonary Artery
Adult Echocardiography Procedure Report | Adult Echo | LPA Vmax | Max PG | meas | m_LPA_Vmax | r_MaxPG | unilateral | 125200\20247-3\T-44000\T-44400 | SRT, T-44000, Pulmonary artery | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-44400, Left Pulmonary Artery
Adult Echocardiography Procedure Report | Adult Echo | Pulm Sys Vel | Vel | meas | m_PulmSysVel | r_Vel | unilateral | 125200\29450-4\T-48581 | SRT, T-48581, Pulmonary Venous Structure | LN, 29450-4, Pulmonary Vein Systolic Peak Velocity |  |  | UCUM, mm/s, mm/s
Adult Echocardiography Procedure Report | Adult Echo | Pulm Dias Vel | Vel | meas | m_PulmDiasVel | r_Vel | unilateral | 125200\29451-2\T-48581 | SRT, T-48581, Pulmonary Venous Structure | LN, 29451-2, Pulmonary Vein Diastolic Peak Velocity |  |  | UCUM, mm/s, mm/s
Adult Echocardiography Procedure Report | Adult Echo | Pulm A Revs Vel | Vel | meas | m_PulmARevsVel | r_Vel | unilateral | 125200\29453-8\T-48581 | SRT, T-48581, Pulmonary Venous Structure | LN, 29453-8, Pulmonary Vein Atrial Contraction Reversal Peak Velocity |  |  | UCUM, mm/s, mm/s
Adult Echocardiography Procedure Report | Adult Echo | Pulm A Revs Dur | Time | meas | m_PulmARevsDur | r_DopTime | unilateral | 125200\G-038B\T-48581 | SRT, T-48581, Pulmonary Venous Structure | SRT, G-038B, Pulmonary Vein A-Wave Duration |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | IVC Diam Ins(MM) | Dist | meas | m_IVC_ins_MM | r_Distance | unilateral | 125200\18006-7\F-20010\G-0394\T-48710 | SRT, T-48710, Inferior vena cava | LN, 18006-7, Inferior Vena Cava Diameter |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode | [9 empty] | SRT, F-20010, During Inspiration
Adult Echocardiography Procedure Report | Adult Echo | IVC Diam Exp(MM) | Dist | meas | m_IVC_exp_MM | r_Distance | unilateral | 125200\18006-7\F-20020\G-0394\T-48710 | SRT, T-48710, Inferior vena cava | LN, 18006-7, Inferior Vena Cava Diameter |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode | [9 empty] | SRT, F-20020, During Expiration
Adult Echocardiography Procedure Report | Adult Echo | IVC Diam | Dist | meas | m_IVC_Diam | r_Distance | unilateral | 125200\18006-7\G-03A2\T-48710 | SRT, T-48710, Inferior vena cava | LN, 18006-7, Inferior Vena Cava Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | IVC Diam Ins | Dist | meas | m_IVC_Diam_Ins | r_Distance | unilateral | 125200\18006-7\F-20010\G-03A2\T-48710 | SRT, T-48710, Inferior vena cava | LN, 18006-7, Inferior Vena Cava Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | [9 empty] | SRT, F-20010, During Inspiration
Adult Echocardiography Procedure Report | Adult Echo | IVC Diam Exp | Dist | meas | m_IVC_Diam_Exp | r_Distance | unilateral | 125200\18006-7\F-20020\G-03A2\T-48710 | SRT, T-48710, Inferior vena cava | LN, 18006-7, Inferior Vena Cava Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | [9 empty] | SRT, F-20020, During Expiration
Adult Echocardiography Procedure Report | Adult Echo | Hep. A Revs Dur | Time | meas | m_HepARevsDur | r_DopTime | unilateral | 125200\C12216-01\T-48720 | SRT, T-48720, Hepatic Vein | 99PMSBLUS, C12216-01, Hepatic Vein A-Wave Duration |  |  | UCUM, ms, ms
Adult Echocardiography Procedure Report | Adult Echo | Hepatic Sys Vel | Vel | meas | m_HepSysVel | r_Vel | unilateral | 125200\29471-0\T-48720 | SRT, T-48720, Hepatic Vein | LN, 29471-0, Hepatic Vein Systolic Peak Velocity |  |  | UCUM, mm/s, mm/s
Adult Echocardiography Procedure Report | Adult Echo | Hepatic Dias Vel | Vel | meas | m_HepDiasVel | r_Vel | unilateral | 125200\29472-8\T-48720 | SRT, T-48720, Hepatic Vein | LN, 29472-8, Hepatic Vein Diastolic Peak Velocity |  |  | UCUM, mm/s, mm/s
Adult Echocardiography Procedure Report | Adult Echo | Hep. A Revs Vel | Vel | meas | m_HepARevsVel | r_Vel | unilateral | 125200\29474-4\T-48720 | SRT, T-48720, Hepatic Vein | LN, 29474-4, Hepatic Vein Atrial Contraction Reversal Peak Velocity |  |  | UCUM, mm/s, mm/s
Adult Echocardiography Procedure Report | Adult Echo | LV Global Endo L End S-aCMQ | Strain | meas | m_CMQ_LVGEnStr_ESAP | r_Strain | unilateral | 125200\C12203-100\C12226-02\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-100, Global Endo End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-02, Combination of AP2-AP3 and AP4 views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Endo C End S-aCMQ | Strain | meas | m_CMQ_LVGEnStr_ESSAX | r_Strain | unilateral | 125200\C12203-100\C12226-08\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-100, Global Endo End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-08, Combination of SAX A, SAX M and SAX B views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Mid L Peak-aCMQ | Strain | meas | m_CMQ_LVGMiStr_PAP | r_Strain | unilateral | 125200\C12203-101\C12226-02\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-101, Global Mid Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-02, Combination of AP2-AP3 and AP4 views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Mid C Peak-aCMQ | Strain | meas | m_CMQ_LVGMiStr_PSAX | r_Strain | unilateral | 125200\C12203-101\C12226-08\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-101, Global Mid Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-08, Combination of SAX A, SAX M and SAX B views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Mid L Peak S-aCMQ | Strain | meas | m_CMQ_LVGMiStr_PSAP | r_Strain | unilateral | 125200\C12203-102\C12226-02\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-102, Global Mid Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-02, Combination of AP2-AP3 and AP4 views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Mid C Peak S-aCMQ | Strain | meas | m_CMQ_LVGMiStr_PSSAX | r_Strain | unilateral | 125200\C12203-102\C12226-08\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-102, Global Mid Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-08, Combination of SAX A, SAX M and SAX B views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Mid L End S-aCMQ | Strain | meas | m_CMQ_LVGMiStr_ESAP | r_Strain | unilateral | 125200\C12203-103\C12226-02\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-103, Global Mid End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-02, Combination of AP2-AP3 and AP4 views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Mid C End S-aCMQ | Strain | meas | m_CMQ_LVGMiStr_ESSAX | r_Strain | unilateral | 125200\C12203-103\C12226-08\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-103, Global Mid End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-08, Combination of SAX A, SAX M and SAX B views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Epi L Peak-aCMQ | Strain | meas | m_CMQ_LVGEpStr_PAP | r_Strain | unilateral | 125200\C12203-104\C12226-02\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-104, Global Epi Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-02, Combination of AP2-AP3 and AP4 views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Epi C Peak-aCMQ | Strain | meas | m_CMQ_LVGEpStr_PSAX | r_Strain | unilateral | 125200\C12203-104\C12226-08\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-104, Global Epi Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-08, Combination of SAX A, SAX M and SAX B views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Epi L Peak S-aCMQ | Strain | meas | m_CMQ_LVGEpStr_PSAP | r_Strain | unilateral | 125200\C12203-105\C12226-02\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-105, Global Epi Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-02, Combination of AP2-AP3 and AP4 views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Epi C Peak S-aCMQ | Strain | meas | m_CMQ_LVGEpStr_PSSAX | r_Strain | unilateral | 125200\C12203-105\C12226-08\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-105, Global Epi Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-08, Combination of SAX A, SAX M and SAX B views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Epi L End S-aCMQ | Strain | meas | m_CMQ_LVGEpStr_ESAP | r_Strain | unilateral | 125200\C12203-106\C12226-02\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-106, Global Epi End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-02, Combination of AP2-AP3 and AP4 views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Epi C End S-aCMQ | Strain | meas | m_CMQ_LVGEpStr_ESSAX | r_Strain | unilateral | 125200\C12203-106\C12226-08\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-106, Global Epi End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-08, Combination of SAX A, SAX M and SAX B views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP3 Endo L Peak-aCMQ | Strain | meas | m_CMQ_Strain_Ap3 | r_Strain | unilateral | 125200\C12203-89\C12226-04\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-89, Endo Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-04, Apical three chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXA Endo C Peak-aCMQ | Strain | meas | m_CMQ_Strain_SaxA | r_Strain | unilateral | 125200\C12203-89\C12226-05\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-89, Endo Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-05, Short Axis Apical |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXM Endo C Peak-aCMQ | Strain | meas | m_CMQ_Strain_SaxM | r_Strain | unilateral | 125200\C12203-89\C12226-06\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-89, Endo Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXB Endo C Peak-aCMQ | Strain | meas | m_CMQ_Strain_SaxB | r_Strain | unilateral | 125200\C12203-89\C12226-07\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-89, Endo Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-07, Short Axis Basal |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP2 Endo L Peak-aCMQ | Strain | meas | m_CMQ_Strain_Ap2 | r_Strain | unilateral | 125200\C12203-89\C12228-10\C99-12228-07\G-03A2\G-A19B\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-89, Endo Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP4 Endo L Peak-aCMQ | Strain | meas | m_CMQ_Strain_Ap4 | r_Strain | unilateral | 125200\C12203-89\C12228-10\C99-12228-07\G-03A2\G-A19C\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-89, Endo Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP3 Endo L Peak S-aCMQ | Strain | meas | m_CMQ_LVEnStr_PSAP3 | r_Strain | unilateral | 125200\C12203-90\C12226-04\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-90, Endo Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-04, Apical three chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXA Endo C Peak S-aCMQ | Strain | meas | m_CMQ_LVEnStr_PSSAXA | r_Strain | unilateral | 125200\C12203-90\C12226-05\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-90, Endo Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-05, Short Axis Apical |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXM Endo C Peak S-aCMQ | Strain | meas | m_CMQ_LVEnStr_PSSAXM | r_Strain | unilateral | 125200\C12203-90\C12226-06\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-90, Endo Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXB Endo C Peak S-aCMQ | Strain | meas | m_CMQ_LVEnStr_PSSAXB | r_Strain | unilateral | 125200\C12203-90\C12226-07\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-90, Endo Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-07, Short Axis Basal |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP2 Endo L Peak S-aCMQ | Strain | meas | m_CMQ_LVEnStr_PSAP2 | r_Strain | unilateral | 125200\C12203-90\C12228-10\C99-12228-07\G-03A2\G-A19B\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-90, Endo Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP4 Endo L Peak S-aCMQ | Strain | meas | m_CMQ_LVEnStr_PSAP4 | r_Strain | unilateral | 125200\C12203-90\C12228-10\C99-12228-07\G-03A2\G-A19C\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-90, Endo Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP3 Endo L End S-aCMQ | Strain | meas | m_CMQ_LVEnStr_ESAP3 | r_Strain | unilateral | 125200\C12203-91\C12226-04\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-91, Endo End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-04, Apical three chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXA Endo C End S-aCMQ | Strain | meas | m_CMQ_LVEnStr_ESSAXA | r_Strain | unilateral | 125200\C12203-91\C12226-05\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-91, Endo End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-05, Short Axis Apical |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXM Endo C End S-aCMQ | Strain | meas | m_CMQ_LVEnStr_ESSAXM | r_Strain | unilateral | 125200\C12203-91\C12226-06\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-91, Endo End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXB Endo C End S-aCMQ | Strain | meas | m_CMQ_LVEnStr_ESSAXB | r_Strain | unilateral | 125200\C12203-91\C12226-07\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-91, Endo End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-07, Short Axis Basal |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP2 Endo L End S-aCMQ | Strain | meas | m_CMQ_LVEnStr_ESAP2 | r_Strain | unilateral | 125200\C12203-91\C12228-10\C99-12228-07\G-03A2\G-A19B\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-91, Endo End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP4 Endo L End S-aCMQ | Strain | meas | m_CMQ_LVEnStr_ESAP4 | r_Strain | unilateral | 125200\C12203-91\C12228-10\C99-12228-07\G-03A2\G-A19C\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-91, Endo End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP3 Mid L Peak-aCMQ | Strain | meas | m_CMQ_LVMiStr_PAP3 | r_Strain | unilateral | 125200\C12203-92\C12226-04\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-92, Mid Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-04, Apical three chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXA Mid C Peak-aCMQ | Strain | meas | m_CMQ_LVMiStr_PSAXA | r_Strain | unilateral | 125200\C12203-92\C12226-05\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-92, Mid Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-05, Short Axis Apical |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXM Mid C Peak-aCMQ | Strain | meas | m_CMQ_LVMiStr_PSAXM | r_Strain | unilateral | 125200\C12203-92\C12226-06\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-92, Mid Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXB Mid C Peak-aCMQ | Strain | meas | m_CMQ_LVMiStr_PSAXB | r_Strain | unilateral | 125200\C12203-92\C12226-07\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-92, Mid Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-07, Short Axis Basal |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP2 Mid L Peak-aCMQ | Strain | meas | m_CMQ_LVMiStr_PAP2 | r_Strain | unilateral | 125200\C12203-92\C12228-10\C99-12228-07\G-03A2\G-A19B\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-92, Mid Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP4 Mid L Peak-aCMQ | Strain | meas | m_CMQ_LVMiStr_PAP4 | r_Strain | unilateral | 125200\C12203-92\C12228-10\C99-12228-07\G-03A2\G-A19C\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-92, Mid Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP3 Mid L Peak S-aCMQ | Strain | meas | m_CMQ_LVMiStr_PSAP3 | r_Strain | unilateral | 125200\C12203-93\C12226-04\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-93, Mid Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-04, Apical three chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXA Mid C Peak S-aCMQ | Strain | meas | m_CMQ_LVMiStr_PSSAXA | r_Strain | unilateral | 125200\C12203-93\C12226-05\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-93, Mid Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-05, Short Axis Apical |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXM Mid C Peak S-aCMQ | Strain | meas | m_CMQ_LVMiStr_PSSAXM | r_Strain | unilateral | 125200\C12203-93\C12226-06\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-93, Mid Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXB Mid C Peak S-aCMQ | Strain | meas | m_CMQ_LVMiStr_PSSAXB | r_Strain | unilateral | 125200\C12203-93\C12226-07\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-93, Mid Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-07, Short Axis Basal |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP2 Mid L Peak S-aCMQ | Strain | meas | m_CMQ_LVMiStr_PSAP2 | r_Strain | unilateral | 125200\C12203-93\C12228-10\C99-12228-07\G-03A2\G-A19B\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-93, Mid Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP4 Mid L Peak S-aCMQ | Strain | meas | m_CMQ_LVMiStr_PSAP4 | r_Strain | unilateral | 125200\C12203-93\C12228-10\C99-12228-07\G-03A2\G-A19C\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-93, Mid Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP3 Mid L End S-aCMQ | Strain | meas | m_CMQ_LVMiStr_ESAP3 | r_Strain | unilateral | 125200\C12203-94\C12226-04\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-94, Mid End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-04, Apical three chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXA Mid C End S-aCMQ | Strain | meas | m_CMQ_LVMiStr_ESSAXA | r_Strain | unilateral | 125200\C12203-94\C12226-05\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-94, Mid End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-05, Short Axis Apical |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXM Mid C End S-aCMQ | Strain | meas | m_CMQ_LVMiStr_ESSAXM | r_Strain | unilateral | 125200\C12203-94\C12226-06\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-94, Mid End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXB Mid C End S-aCMQ | Strain | meas | m_CMQ_LVMiStr_ESSAXB | r_Strain | unilateral | 125200\C12203-94\C12226-07\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-94, Mid End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-07, Short Axis Basal |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP2 Mid L End S-aCMQ | Strain | meas | m_CMQ_LVMiStr_ESAP2 | r_Strain | unilateral | 125200\C12203-94\C12228-10\C99-12228-07\G-03A2\G-A19B\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-94, Mid End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP4 Mid L End S-aCMQ | Strain | meas | m_CMQ_LVMiStr_ESAP4 | r_Strain | unilateral | 125200\C12203-94\C12228-10\C99-12228-07\G-03A2\G-A19C\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-94, Mid End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP3 Epi L Peak-aCMQ | Strain | meas | m_CMQ_LVEpStr_PAP3 | r_Strain | unilateral | 125200\C12203-95\C12226-04\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-95, Epi Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-04, Apical three chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXA Epi C Peak-aCMQ | Strain | meas | m_CMQ_LVEpStr_PSAXA | r_Strain | unilateral | 125200\C12203-95\C12226-05\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-95, Epi Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-05, Short Axis Apical |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXM Epi C Peak-aCMQ | Strain | meas | m_CMQ_LVEpStr_PSAXM | r_Strain | unilateral | 125200\C12203-95\C12226-06\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-95, Epi Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXB Epi C Peak-aCMQ | Strain | meas | m_CMQ_LVEpStr_PSAXB | r_Strain | unilateral | 125200\C12203-95\C12226-07\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-95, Epi Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-07, Short Axis Basal |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP2 Epi L Peak-aCMQ | Strain | meas | m_CMQ_LVEpStr_PAP2 | r_Strain | unilateral | 125200\C12203-95\C12228-10\C99-12228-07\G-03A2\G-A19B\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-95, Epi Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP4 Epi L Peak-aCMQ | Strain | meas | m_CMQ_LVEpStr_PAP4 | r_Strain | unilateral | 125200\C12203-95\C12228-10\C99-12228-07\G-03A2\G-A19C\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-95, Epi Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP3 Epi L Peak S-aCMQ | Strain | meas | m_CMQ_LVEpStr_PSAP3 | r_Strain | unilateral | 125200\C12203-96\C12226-04\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-96, Epi Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-04, Apical three chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXA Epi C Peak S-aCMQ | Strain | meas | m_CMQ_LVEpStr_PSSAXA | r_Strain | unilateral | 125200\C12203-96\C12226-05\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-96, Epi Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-05, Short Axis Apical |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXM Epi C Peak S-aCMQ | Strain | meas | m_CMQ_LVEpStr_PSSAXM | r_Strain | unilateral | 125200\C12203-96\C12226-06\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-96, Epi Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXB Epi C Peak S-aCMQ | Strain | meas | m_CMQ_LVEpStr_PSSAXB | r_Strain | unilateral | 125200\C12203-96\C12226-07\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-96, Epi Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-07, Short Axis Basal |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP2 Epi L Peak S-aCMQ | Strain | meas | m_CMQ_LVEpStr_PSAP2 | r_Strain | unilateral | 125200\C12203-96\C12228-10\C99-12228-07\G-03A2\G-A19B\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-96, Epi Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP4 Epi L Peak S-aCMQ | Strain | meas | m_CMQ_LVEpStr_PSAP4 | r_Strain | unilateral | 125200\C12203-96\C12228-10\C99-12228-07\G-03A2\G-A19C\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-96, Epi Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP3 Epi L End S-aCMQ | Strain | meas | m_CMQ_LVEpStr_ESAP3 | r_Strain | unilateral | 125200\C12203-97\C12226-04\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-97, Epi End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-04, Apical three chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXA Epi C End S-aCMQ | Strain | meas | m_CMQ_LVEpStr_ESSAXA | r_Strain | unilateral | 125200\C12203-97\C12226-05\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-97, Epi End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-05, Short Axis Apical |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXM Epi C End S-aCMQ | Strain | meas | m_CMQ_LVEpStr_ESSAXM | r_Strain | unilateral | 125200\C12203-97\C12226-06\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-97, Epi End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV SAXB Epi C End S-aCMQ | Strain | meas | m_CMQ_LVEpStr_ESSAXB | r_Strain | unilateral | 125200\C12203-97\C12226-07\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-97, Epi End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-07, Short Axis Basal |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP2 Epi L End S-aCMQ | Strain | meas | m_CMQ_LVEpStr_ESAP2 | r_Strain | unilateral | 125200\C12203-97\C12228-10\C99-12228-07\G-03A2\G-A19B\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-97, Epi End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV AP4 Epi L End S-aCMQ | Strain | meas | m_CMQ_LVEpStr_ESAP4 | r_Strain | unilateral | 125200\C12203-97\C12228-10\C99-12228-07\G-03A2\G-A19C\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-97, Epi End Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Endo L Peak-aCMQ | Strain | meas | m_CMQ_Strain_GL | r_Strain | unilateral | 125200\C12203-98\C12226-02\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-98, Global Endo Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-02, Combination of AP2-AP3 and AP4 views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Endo C Peak-aCMQ | Strain | meas | m_CMQ_Strain_GC | r_Strain | unilateral | 125200\C12203-98\C12226-08\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-98, Global Endo Peak Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-08, Combination of SAX A, SAX M and SAX B views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Endo L Peak S-aCMQ | Strain | meas | m_CMQ_LVGEnStr_PSAP | r_Strain | unilateral | 125200\C12203-99\C12226-02\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-99, Global Endo Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-02, Combination of AP2-AP3 and AP4 views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | LV Global Endo C Peak S-aCMQ | Strain | meas | m_CMQ_LVGEnStr_PSSAX | r_Strain | unilateral | 125200\C12203-99\C12226-08\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-99, Global Endo Peak Systolic Strain of Left Ventricle |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-08, Combination of SAX A, SAX M and SAX B views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [7 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | MV DVI-pr | none | calc | c_MV_DVI_pr | none | unilateral | 125200\C12207-46\C12220-10 | 99PMSBLUS, C12207-46, Mitral Valve (Prosthetic) | 99PMSBLUS, C12220-10, Doppler Velocity Index |  |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | MV EOA-pr | none | calc | c_MV_EOA_pr | none | unilateral | 125200\C12207-46\C12222-26 | 99PMSBLUS, C12207-46, Mitral Valve (Prosthetic) | 99PMSBLUS, C12222-26, Effective Orifice Area |  |  | UCUM, mm2, mm2
Adult Echocardiography Procedure Report | Adult Echo | AV EOAI-pr | none | calc | c_AV_EOAI_pr | none | unilateral | 125200\C12211-14\C12211-15\G-03A2 | 99PMSBLUS, C12211-15, Aortic Valve (Prosthetic) | 99PMSBLUS, C12211-14, Prosthetic Aortic Valve Effective Orifice Area Indexed To BSA |  |  | UCUM, cm2/m2, cm2/m2 |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | AV DVI-pr | none | calc | c_AV_DVI_pr | none | unilateral | 125200\C12211-15\C12220-10 | 99PMSBLUS, C12211-15, Aortic Valve (Prosthetic) | 99PMSBLUS, C12220-10, Doppler Velocity Index |  |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | AV EOA-pr | none | calc | c_AV_EOA_pr | none | unilateral | 125200\C12211-15\C12222-26\G-03A2 | 99PMSBLUS, C12211-15, Aortic Valve (Prosthetic) | 99PMSBLUS, C12222-26, Effective Orifice Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LV Mass Index (A/L) | none | calc | c_LVMass_A_LBSA | none | unilateral | 125200\C12203-01\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-01, Left Ventricle Mass Index |  |  | UCUM, g/m2, g/m2
Adult Echocardiography Procedure Report | Adult Echo | LV Mass Index(Cubed) | none | calc | c_LV_Mass_BSA | none | unilateral | 125200\125206\C12203-01\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-01, Left Ventricle Mass Index |  |  | UCUM, g/m2, g/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | LV PEP/ET | none | calc | c_LV_PEP_ET | none | unilateral | 125200\C12203-04\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C12203-04, PEP/ET |  |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | EF (Dumesnil) | none | calc | c_EF_Dumesnil | none | unilateral | 125200\C3467-04\T-32600 | SNM3, T-32600, Left Ventricle | 99PMSBLUS, C3467-04, Left Ventricular Ejection Fraction by Dumesnil 1995 |  |  | UCUM, %, Percent
Adult Echocardiography Procedure Report | Adult Echo | EDV (MM-Cubed) | none | calc | c_EDV_Cubed_MM | none | unilateral | 125200\125206\18026-5\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | EDV (MM-Teich) | none | calc | c_EDV_Teich_MM | none | unilateral | 125200\125209\18026-5\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | EDV (2D-Cubed) | none | calc | c_EDV_Cubed | none | unilateral | 125200\125206\18026-5\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | EDV (BP) | none | calc | c_EDV_BP | none | unilateral | 125200\125207\18026-5\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | EDV (2D-Teich) | none | calc | c_EDV_Teich | none | unilateral | 125200\125209\18026-5\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | EDV (A/L) | none | calc | c_EDV_SPE | none | unilateral | 125200\125226\18026-5\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | EF (MM-Cubed) | none | calc | c_EF_Cubed_MM | none | unilateral | 125200\125206\18043-0\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | EF (MM-Teich) | none | calc | c_EF_Teich_MM | none | unilateral | 125200\125209\18043-0\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | EF (2D-Cubed) | none | calc | c_EF_Cubed | none | unilateral | 125200\125206\18043-0\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | EF (BP) | none | calc | c_EF_BP | none | unilateral | 125200\125207\18043-0\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | EF (2D-Teich) | none | calc | c_EF_Teich | none | unilateral | 125200\125209\18043-0\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | EF (A/L) | none | calc | c_EF_SPE | none | unilateral | 125200\125226\18043-0\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | EF (A2C) | none | calc | c_EF_AP2 | none | unilateral | 125200\125208\18043-0\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | EF (A4C) | none | calc | c_EF_AP4 | none | unilateral | 125200\125208\18043-0\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | FS (MM-Cubed) | none | calc | c_FS_Cubed_MM | none | unilateral | 125200\125206\18051-3\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18051-3, Left Ventricular Fractional Shortening |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | FS (MM-Teich) | none | calc | c_FS_Teich_MM | none | unilateral | 125200\125209\18051-3\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18051-3, Left Ventricular Fractional Shortening |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | FS (2D-Cubed) | none | calc | c_FS_Cubed | none | unilateral | 125200\125206\18051-3\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18051-3, Left Ventricular Fractional Shortening |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | FS (2D-Teich) | none | calc | c_FS_Teich | none | unilateral | 125200\125209\18051-3\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18051-3, Left Ventricular Fractional Shortening |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | LVPW % (MM) | none | calc | c_LVPW_MM | none | unilateral | 125200\18053-9\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18053-9, Left Ventricle Posterior Wall % Thickening |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LVPW % (2D) | none | calc | c_LVPW | none | unilateral | 125200\18053-9\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18053-9, Left Ventricle Posterior Wall % Thickening |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | IVS % (MM) | none | calc | c_IVS_MM | none | unilateral | 125200\18054-7\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18054-7, Interventricular Septum % Thickening |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | IVS % (2D) | none | calc | c_IVS | none | unilateral | 125200\18054-7\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18054-7, Interventricular Septum % Thickening |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LV Mass (Cubed) | none | calc | c_LV_Mass | none | unilateral | 125200\125206\18087-7\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18087-7, Left Ventricle Mass |  |  | UCUM, g, grams |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | LV Mass (A/L) | none | calc | c_LVMass_A_L | none | unilateral | 125200\18087-7\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18087-7, Left Ventricle Mass |  |  | UCUM, g, grams |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | ESV (MM-Cubed) | none | calc | c_ESV_Cubed_MM | none | unilateral | 125200\125206\18148-7\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | ESV (MM-Teich) | none | calc | c_ESV_Teich_MM | none | unilateral | 125200\125209\18148-7\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | ESV (2D-Cubed) | none | calc | c_ESV_Cubed | none | unilateral | 125200\125206\18148-7\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | ESV (BP) | none | calc | c_ESV_BP | none | unilateral | 125200\125207\18148-7\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | ESV (2D-Teich) | none | calc | c_ESV_Teich | none | unilateral | 125200\125209\18148-7\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | ESV (A/L) | none | calc | c_ESV_SPE | none | unilateral | 125200\125226\18148-7\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | IVS/LVPW (MM) | none | calc | c_IVS_LVPW_MM | none | unilateral | 125200\18155-2\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18155-2, Interventricular Septum to Posterior Wall Thickness Ratio |  |  | UCUM, 1, no units |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | IVS/LVPW (2D) | none | calc | c_IVS_LVPW | none | unilateral | 125200\18155-2\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | LN, 18155-2, Interventricular Septum to Posterior Wall Thickness Ratio |  |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | E`/A` Medial | none | calc | c_Med_E_A | none | unilateral | 125200\59129-7\G-0391\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | LN, 59129-7, Left Ventricle E to A Tissue Velocity Ratio |  |  | UCUM, 1, no units | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | E`/A` Lateral | none | calc | c_Lat_E_A | none | unilateral | 125200\59129-7\G-0392\P5-B0128\T-32600 | SNM3, T-32600, Left Ventricle | LN, 59129-7, Left Ventricle E to A Tissue Velocity Ratio |  |  | UCUM, 1, no units | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | SI (MM-Cubed) | none | calc | c_SI_Cubed_MM | none | unilateral | 125200\125206\F-00078\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | SI (MM-Teich) | none | calc | c_SI_Teich_MM | none | unilateral | 125200\125209\F-00078\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | SI (2D-Cubed) | none | calc | c_SI_Cubed | none | unilateral | 125200\125206\F-00078\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | SI (BP) | none | calc | c_SI_BP | none | unilateral | 125200\125207\F-00078\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | SI (2D-Teich) | none | calc | c_SI_Teich | none | unilateral | 125200\125209\F-00078\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | SI (A/L) | none | calc | c_SI_SPE | none | unilateral | 125200\125226\F-00078\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | SI (A2C) | none | calc | c_SI_AP2 | none | unilateral | 125200\125208\F-00078\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | SI (A4C) | none | calc | c_SI_AP4 | none | unilateral | 125200\125208\F-00078\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | CO (MM-Cubed) | none | calc | c_CO_Cubed_MM | none | unilateral | 125200\125206\F-32100\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | CO (MM-Teich) | none | calc | c_CO_Teich_MM | none | unilateral | 125200\125209\F-32100\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | CO (2D-Cubed) | none | calc | c_CO_Cubed | none | unilateral | 125200\125206\F-32100\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | CO (BP) | none | calc | c_CO_BP | none | unilateral | 125200\125207\F-32100\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | CO (2D-Teich) | none | calc | c_CO_Teich | none | unilateral | 125200\125209\F-32100\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | CO (A/L) | none | calc | c_CO_SPE | none | unilateral | 125200\125226\F-32100\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | CO (A2C) | none | calc | c_CO_AP2 | none | unilateral | 125200\125208\F-32100\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | CO (A4C) | none | calc | c_CO_AP4 | none | unilateral | 125200\125208\F-32100\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | CO (LVOT) | none | calc | c_CO_LVOT | none | unilateral | 125200\F-32100\T-32600\T-32650 | SNM3, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s | SNM3, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | CI (MM-Cubed) | none | calc | c_CI_Cubed_MM | none | unilateral | 125200\125206\F-32110\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | CI (MM-Teich) | none | calc | c_CI_Teich_MM | none | unilateral | 125200\125209\F-32110\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | CI (2D-Cubed) | none | calc | c_CI_Cubed | none | unilateral | 125200\125206\F-32110\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | CI (BP) | none | calc | c_CI_BP | none | unilateral | 125200\125207\F-32110\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | CI (2D-Teich) | none | calc | c_CI_Teich | none | unilateral | 125200\125209\F-32110\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | CI (A/L) | none | calc | c_CI_SPE | none | unilateral | 125200\125226\F-32110\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | CI (A2C) | none | calc | c_CI_AP2 | none | unilateral | 125200\125208\F-32110\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | CI (A4C) | none | calc | c_CI_AP4 | none | unilateral | 125200\125208\F-32110\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | SV (MM-Cubed) | none | calc | c_SV_Cubed_MM | none | unilateral | 125200\125206\F-32120\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | SV (MM-Teich) | none | calc | c_SV_Teich_MM | none | unilateral | 125200\125209\F-32120\G-0394\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | SV (2D-Cubed) | none | calc | c_SV_Cubed | none | unilateral | 125200\125206\F-32120\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | SV (BP) | none | calc | c_SV_BP | none | unilateral | 125200\125207\F-32120\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | SV (2D-Teich) | none | calc | c_SV_Teich | none | unilateral | 125200\125209\F-32120\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | SV (A/L) | none | calc | c_SV_SPE | none | unilateral | 125200\125226\F-32120\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | SV (A2C) | none | calc | c_SV_AP2 | none | unilateral | 125200\125208\F-32120\G-03A2\G-A19B\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | SV (A4C) | none | calc | c_SV_AP4 | none | unilateral | 125200\125208\F-32120\G-03A2\G-A19C\T-32600 | SNM3, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | SV (LVOT) | none | calc | c_SV_LVOT | none | unilateral | 125200\F-32120\T-32600\T-32650 | SNM3, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 | SNM3, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | LV FAC | none | calc | c_LV_FAC | none | unilateral | 125200\G-0376\G-039B\G-03A2\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-0376, Left Ventricular Fractional Area Change |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level
Adult Echocardiography Procedure Report | Adult Echo | E/Med E` | none | calc | c_E_MedE | none | unilateral | 125200\G-037B\G-0391\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-037B, Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave |  |  | UCUM, 1, no units | SRT, G-0391, Medial Mitral Annulus
Adult Echocardiography Procedure Report | Adult Echo | E/Lat E` | none | calc | c_E_LatE | none | unilateral | 125200\G-037B\G-0392\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-037B, Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave |  |  | UCUM, 1, no units | SRT, G-0392, Lateral Mitral Annulus
Adult Echocardiography Procedure Report | Adult Echo | LV MPI | none | calc | c_LV_MPI | none | unilateral | 125200\G-037F\T-32600 | SNM3, T-32600, Left Ventricle | SRT, G-037F, Left Ventricular Index of Myocardial Performance |  |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | LVOT Area | none | calc | c_LVOT_Area | none | unilateral | 125200\G-038E\G-03A2\T-32600\T-32650 | SNM3, T-32600, Left Ventricle | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | SNM3, T-32650, Left Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | PISA (MR) | none | calc | c_PISA_MR | none | unilateral | 125200\125216\C12207-06\R-42E61\T-35300 | SNM3, T-35300, Mitral Valve | 99PMSBLUS, C12207-06, Mitral Valve Flow Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | MV E/A | none | calc | c_MV_EA | none | unilateral | 125200\18038-0\T-35300 | SNM3, T-35300, Mitral Valve | LN, 18038-0, Mitral Valve E to A Ratio |  |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | MR Volume | none | calc | c_MR_Vol | none | unilateral | 125200\125216\33878-0\R-42E61\T-35300 | SNM3, T-35300, Mitral Valve | LN, 33878-0, Volume Flow |  |  | UCUM, mm3, mm3 |  |  |  |  | SRT, R-42E61, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | MR Flow Rate | none | calc | c_MR_FlowRate | none | unilateral | 125200\34141-2\R-42E61\T-35300 | SNM3, T-35300, Mitral Valve | LN, 34141-2, Peak Instantaneous Flow Rate |  |  | UCUM, mm3/s, mm3/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | Tei Index | none | calc | c_TeiIndex | none | unilateral | 125200\59099-2\T-35300 | SNM3, T-35300, Mitral Valve | LN, 59099-2, Myocardial Performance Index (Tei) |  |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | CO (MV) | none | calc | c_CO_MV | none | unilateral | 125200\F-32100\T-35300 | SNM3, T-35300, Mitral Valve | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s
Adult Echocardiography Procedure Report | Adult Echo | SV (MV) | none | calc | c_SV_MV | none | unilateral | 125200\F-32120\T-35300 | SNM3, T-35300, Mitral Valve | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3
Adult Echocardiography Procedure Report | Adult Echo | MVA (P½t ) | none | calc | c_MVA_PHalfT | none | unilateral | 125200\125210\G-038E\T-35300 | SNM3, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125210, Area by Pressure Half-Time
Adult Echocardiography Procedure Report | Adult Echo | MVA (VTI) | none | calc | c_MVA_VTI | none | unilateral | 125200\125215\G-038E\T-35300 | SNM3, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125215, Continuity Equation by Velocity Time Integral
Adult Echocardiography Procedure Report | Adult Echo | MVA (PISA) | none | calc | c_MVA_PISA | none | unilateral | 125200\125216\G-038E\T-35300 | SNM3, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | MR ERO | none | calc | c_MR_ERO | none | unilateral | 125200\125216\G-038E\R-42E61\T-35300 | SNM3, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | MR Fraction | none | calc | c_MR_Frac | none | unilateral | 125200\G-0390\R-42E61\T-35300 | SNM3, T-35300, Mitral Valve | SRT, G-0390, Regurgitant Fraction |  |  | UCUM, %, Percent |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | Qp/Qs | none | calc | c_QpQs | none | unilateral | 125200\29462-9\P5-30031 | SRT, P5-30031, Cardiac Shunt Study | LN, 29462-9, Pulmonary-to-Systemic Shunt Flow Ratio |  |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | RA ESV-A/L | none | calc | c_RA_ESV_AL | none | unilateral | 125200\109070\125204\C12206-01\G-03A2\T-32200 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-01, Right Atrium Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | DCM, 109070, End Systole |  | DCM, 125204, Area-Length Biplane
Adult Echocardiography Procedure Report | Adult Echo | RA ESV (BP) | none | calc | c_ESV_RA_BP | none | unilateral | 125200\109070\125207\C12206-01\G-03A2\T-32200 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-01, Right Atrium Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | DCM, 109070, End Systole |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | RA ESV Index (BP) | none | calc | c_BP_RA_ESV_BSA | none | unilateral | 125200\109070\125207\C12206-03\G-03A2\T-32200 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-03, Right Atrium Systolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  | DCM, 109070, End Systole |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | RA ESV Index (A2C) | none | calc | c_A2C_RA_ESV_BSA | none | unilateral | 125200\109070\125208\C12206-03\G-03A2\G-A19B\T-32200 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-03, Right Atrium Systolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | RA ESV Index (A4C) | none | calc | c_A4C_RA_ESV_BSA | none | unilateral | 125200\109070\125208\C12206-03\G-03A2\G-A19C\T-32200 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-03, Right Atrium Systolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LA ESV Index (BP) | none | calc | c_BP_LA_ESV_BSA | none | unilateral | 125200\125207\C12205-03\G-03A2\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-03, Left Atrium Systolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | LA ESV Index (A2C) | none | calc | c_A2C_LA_ESV_BSA | none | unilateral | 125200\125208\C12205-03\G-03A2\G-A19B\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-03, Left Atrium Systolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LA ESV Index (A4C) | none | calc | c_A4C_LA_ESV_BSA | none | unilateral | 125200\125208\C12205-03\G-03A2\G-A19C\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-03, Left Atrium Systolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LA/Ao (MM) | none | calc | c_LA_Ao_MM | none | unilateral | 125200\17985-3\G-0394\T-32300 | SRT, T-32300, Left Atrium | LN, 17985-3, Left Atrium to Aortic Root Ratio |  |  | UCUM, 1, no units |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LA/Ao (2D) | none | calc | c_LA_Ao_2D | none | unilateral | 125200\17985-3\G-03A2\T-32300 | SRT, T-32300, Left Atrium | LN, 17985-3, Left Atrium to Aortic Root Ratio |  |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LA ESV (BP) | none | calc | c_ESV_LA_BP | none | unilateral | 125200\125207\G-0383\G-03A2\T-32300 | SRT, T-32300, Left Atrium | SRT, G-0383, Left Atrium Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | LA ESV-A/L | none | calc | c_LA_ESV_AL | none | unilateral | 125200\109070\125204\G-0383\G-03A2\T-32300 | SRT, T-32300, Left Atrium | SRT, G-0383, Left Atrium Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | DCM, 109070, End Systole |  | DCM, 125204, Area-Length Biplane
Adult Echocardiography Procedure Report | Adult Echo | RV PEP/ET | none | calc | c_RV_PEP_ET | none | unilateral | 125200\C12203-04\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12203-04, PEP/ET |  |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | RV FAC | none | calc | c_RV_FAC | none | unilateral | 125200\C12204-03\G-03A2\G-A19C\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-03, Right Ventricular Fractional Area Change |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber
Adult Echocardiography Procedure Report | Adult Echo | RVIDd/LVIDd | none | calc | c_RVID_LVID_ratio | none | unilateral | 125200\C12204-04\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-04, Right and Left Ventricular End Diastolic Diameter Ratio |  |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | CO (RVOT) | none | calc | c_CO_RVOT | none | unilateral | 125200\F-32100\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s | SNM3, T-32550, Right Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | SV (RVOT) | none | calc | c_SV_RVOT | none | unilateral | 125200\F-32120\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 | SNM3, T-32550, Right Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | RVSP | none | calc | c_RVSP | none | unilateral | 125200\G-0380\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-0380, Right Ventricular Peak Systolic Pressure |  |  | UCUM, mmHg, mmHg
Adult Echocardiography Procedure Report | Adult Echo | RV MPI | none | calc | c_RV_MPI | none | unilateral | 125200\G-0381\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-0381, Right Ventricular Index of Myocardial Performance |  |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | RVOT Area | none | calc | c_RVOT_Area | none | unilateral | 125200\G-038E\G-03A2\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | SNM3, T-32550, Right Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | PISA (TR) | none | calc | c_PISA_TR | none | unilateral | 125200\125216\C12208-05\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-05, Tricuspid Valve Flow Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | TV E/A | none | calc | c_TV_EA | none | unilateral | 125200\18039-8\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 18039-8, Tricuspid Valve E to A Ratio |  |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | TR Volume | none | calc | c_TR_Vol | none | unilateral | 125200\33878-0\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 33878-0, Volume Flow |  |  | UCUM, mm3, mm3 |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TR Flow Rate | none | calc | c_TR_FlowRate | none | unilateral | 125200\34141-2\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 34141-2, Peak Instantaneous Flow Rate |  |  | UCUM, mm3/s, mm3/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | CO (TV) | none | calc | c_CO_TV | none | unilateral | 125200\F-32100\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s
Adult Echocardiography Procedure Report | Adult Echo | SV (TV) | none | calc | c_SV_TV | none | unilateral | 125200\F-32120\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3
Adult Echocardiography Procedure Report | Adult Echo | TVA (PISA) | none | calc | c_TVA_PISA | none | unilateral | 125200\125216\G-038E\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | TR ERO | none | calc | c_TR_ERO | none | unilateral | 125200\125216\G-038E\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | TV Area | none | calc | c_TV_Area | none | unilateral | 125200\G-038E\G-03A2\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | TR Fraction | none | calc | c_TR_Frac | none | unilateral | 125200\G-0390\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-0390, Regurgitant Fraction |  |  | UCUM, %, Percent |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | PVA (Vmax) | none | calc | c_PVA_Vmax | none | unilateral | 125200\125214\G-038E\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125214, Continuity Equation by Peak Velocity
Adult Echocardiography Procedure Report | Adult Echo | PVA (VTI) | none | calc | c_PVA_VTI | none | unilateral | 125200\125215\G-038E\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125215, Continuity Equation by Velocity Time Integral
Adult Echocardiography Procedure Report | Adult Echo | PISA (AI) | none | calc | c_PISA_AI | none | unilateral | 125200\125216\C12211-01\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12211-01, Aortic Valve Flow Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | Ao Ann Diam Ratio | none | calc | c_Ao_Ann_Diam_Ratio | none | unilateral | 125200\C12211-11\G-03A2\T-35400\T-35410 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12211-11, Aortic Annulus Diameter Ratio |  |  | UCUM, 1, no units | SRT, T-35410, Aortic Valve Ring | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | AV VR | none | calc | c_AV_VR | none | unilateral | 125200\C12211-17\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12211-17, Aortic Valve Velocity Ratio |  |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | AVA(VTI)/BSA | none | calc | c_AVA_VTI_BSA | none | unilateral | 125200\C12211-18\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12211-18, Aortic Valve Area Indexed To BSA |  |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | AI Volume | none | calc | c_AI_Vol | none | unilateral | 125200\33878-0\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 33878-0, Volume Flow |  |  | UCUM, mm3, mm3 |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Flow Rate | none | calc | c_AI_FlowRate | none | unilateral | 125200\34141-2\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 34141-2, Peak Instantaneous Flow Rate |  |  | UCUM, mm3/s, mm3/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AVA (Vmax) | none | calc | c_AVA_Vmax | none | unilateral | 125200\125214\G-038E\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125214, Continuity Equation by Peak Velocity
Adult Echocardiography Procedure Report | Adult Echo | AVA (VTI) | none | calc | c_AVA_VTI | none | unilateral | 125200\125215\G-038E\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125215, Continuity Equation by Velocity Time Integral
Adult Echocardiography Procedure Report | Adult Echo | AI ERO | none | calc | c_AI_ERO | none | unilateral | 125200\125216\G-038E\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | AVA(3DQ-SV | none | calc | c_3DQ_AVA | none | unilateral | 125200\12229-01\G-038E\T12224-01\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, 12229-01, Aortic Valve Area via 3DQ Stroke Volume and AV VTI
Adult Echocardiography Procedure Report | Adult Echo | AVA(3DQA-SV | none | calc | c_3DQA_AVA | none | unilateral | 125200\12229-02\G-038E\T12224-01\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  | 99PMSBLUS, T12224-01, 3D mode |  |  |  | 99PMSBLUS, 12229-02, Aortic Valve Area via 3DQA Stroke Volume and AV VTI
Adult Echocardiography Procedure Report | Adult Echo | AI Fraction | none | calc | c_AI_Frac | none | unilateral | 125200\G-0390\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-0390, Regurgitant Fraction |  |  | UCUM, %, Percent |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | Ao STJ Diam Ratio | none | calc | c_Ao_STJ_Diam_Ratio | none | unilateral | 125200\C12211-12\G-03A2\T-42102 | SRT, T-42102, Aortic Sinotubular Junction | 99PMSBLUS, C12211-12, Aortic Sinotubular Junction Diameter Ratio |  |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Ao Sinus Diam Ratio | none | calc | c_Ao_Sin_Diam_Ratio | none | unilateral | 125200\C12211-13\G-03A2\T-42200 | SRT, T-42200, Structure Sinus of Valsalva | 99PMSBLUS, C12211-13, Aortic Sinus Diameter Ratio |  |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Pulm S/D | none | calc | c_PulmSD | none | unilateral | 125200\29452-0\T-48581 | SRT, T-48581, Pulmonary Venous Structure | LN, 29452-0, Pulmonary Vein Systolic to Diastolic Ratio |  |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | Hepatic S/D | none | calc | c_HepSD | none | unilateral | 125200\29473-6\T-48720 | SRT, T-48720, Hepatic Vein | LN, 29473-6, Hepatic Vein Systolic to Diastolic Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Heart Area | Area | meas | m_fet_hrtArea | r_Area | unilateral | 125196\99999\C12005-09\G-03A2 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, C12005-09, Heart Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | 2 Beat Pk-to-Pk(DOP) | Time | meas | m_DOP_2BeatPKToPK | r_DopTime | unilateral | 125196\99999\C12019-01 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, C12019-01, Peak-to-Peak time interval over two beats |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | 2 Beat Pk-to-Pk(MM) | Time | meas | m_MM_2BeatPKToPK | r_MMTime | unilateral | 125196\99999\C12019-01\G-0394 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, C12019-01, Peak-to-Peak time interval over two beats |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Ant ICA | ? | meas | m_fet_antIca | r_DopAngle | unilateral | 125196\C12104-03\C12122-01 | 99PMSBLUS, C12104-03, Anterior Internal Carotid Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | Ant ICA | EDV | meas | m_fet_antIca | r_EDV | unilateral | 125196\11653-3\C12104-03 | 99PMSBLUS, C12104-03, Anterior Internal Carotid Artery | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Ant ICA | MDV | meas | m_fet_antIca | r_MDV | unilateral | 125196\11665-7\C12104-03 | 99PMSBLUS, C12104-03, Anterior Internal Carotid Artery | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Ant ICA | TAPV | meas | m_fet_antIca | r_TAPV | unilateral | 125196\11692-1\C12104-03 | 99PMSBLUS, C12104-03, Anterior Internal Carotid Artery | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Ant ICA | PSV | meas | m_fet_antIca | r_PSV | unilateral | 125196\11726-7\C12104-03 | 99PMSBLUS, C12104-03, Anterior Internal Carotid Artery | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Ant ICA | PI | meas | m_fet_antIca | r_PI | unilateral | 125196\12008-9\C12104-03 | 99PMSBLUS, C12104-03, Anterior Internal Carotid Artery | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Ant ICA | RI | meas | m_fet_antIca | r_RI | unilateral | 125196\12023-8\C12104-03 | 99PMSBLUS, C12104-03, Anterior Internal Carotid Artery | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Ant ICA | S/D | meas | m_fet_antIca | r_SD | unilateral | 125196\12144-2\C12104-03 | 99PMSBLUS, C12104-03, Anterior Internal Carotid Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Ant ICA | AI | meas | m_fet_antIca | r_AccelIndex | unilateral | 125196\20167-3\C12104-03 | 99PMSBLUS, C12104-03, Anterior Internal Carotid Artery | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | Ant ICA | AT | meas | m_fet_antIca | r_AccelTime | unilateral | 125196\20168-1\C12104-03 | 99PMSBLUS, C12104-03, Anterior Internal Carotid Artery | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Ant ICA | DT | meas | m_fet_antIca | r_DecelTime | unilateral | 125196\20217-6\C12104-03 | 99PMSBLUS, C12104-03, Anterior Internal Carotid Artery | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Ant ICA | TAMV | meas | m_fet_antIca | r_TAMV | unilateral | 125196\20352-1\C12104-03 | 99PMSBLUS, C12104-03, Anterior Internal Carotid Artery | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Ant ICA | VTI | meas | m_fet_antIca | r_VTI | unilateral | 125196\20354-7\C12104-03 | 99PMSBLUS, C12104-03, Anterior Internal Carotid Artery | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Ant ICA | HR | meas | m_fet_antIca | r_HeartRate | unilateral | 125196\8867-4\C12104-03 | 99PMSBLUS, C12104-03, Anterior Internal Carotid Artery | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | Ant ICA | Dist | meas | m_fet_antIca | r_Distance | unilateral | 125196\C12104-03\G-0364 | 99PMSBLUS, C12104-03, Anterior Internal Carotid Artery | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Post ICA | ? | meas | m_fet_postIca | r_DopAngle | unilateral | 125196\C12104-04\C12122-01 | 99PMSBLUS, C12104-04, Posterior Internal Carotid Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | Post ICA | EDV | meas | m_fet_postIca | r_EDV | unilateral | 125196\11653-3\C12104-04 | 99PMSBLUS, C12104-04, Posterior Internal Carotid Artery | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Post ICA | MDV | meas | m_fet_postIca | r_MDV | unilateral | 125196\11665-7\C12104-04 | 99PMSBLUS, C12104-04, Posterior Internal Carotid Artery | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Post ICA | TAPV | meas | m_fet_postIca | r_TAPV | unilateral | 125196\11692-1\C12104-04 | 99PMSBLUS, C12104-04, Posterior Internal Carotid Artery | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Post ICA | PSV | meas | m_fet_postIca | r_PSV | unilateral | 125196\11726-7\C12104-04 | 99PMSBLUS, C12104-04, Posterior Internal Carotid Artery | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Post ICA | PI | meas | m_fet_postIca | r_PI | unilateral | 125196\12008-9\C12104-04 | 99PMSBLUS, C12104-04, Posterior Internal Carotid Artery | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Post ICA | RI | meas | m_fet_postIca | r_RI | unilateral | 125196\12023-8\C12104-04 | 99PMSBLUS, C12104-04, Posterior Internal Carotid Artery | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Post ICA | S/D | meas | m_fet_postIca | r_SD | unilateral | 125196\12144-2\C12104-04 | 99PMSBLUS, C12104-04, Posterior Internal Carotid Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Post ICA | AI | meas | m_fet_postIca | r_AccelIndex | unilateral | 125196\20167-3\C12104-04 | 99PMSBLUS, C12104-04, Posterior Internal Carotid Artery | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | Post ICA | AT | meas | m_fet_postIca | r_AccelTime | unilateral | 125196\20168-1\C12104-04 | 99PMSBLUS, C12104-04, Posterior Internal Carotid Artery | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Post ICA | DT | meas | m_fet_postIca | r_DecelTime | unilateral | 125196\20217-6\C12104-04 | 99PMSBLUS, C12104-04, Posterior Internal Carotid Artery | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Post ICA | TAMV | meas | m_fet_postIca | r_TAMV | unilateral | 125196\20352-1\C12104-04 | 99PMSBLUS, C12104-04, Posterior Internal Carotid Artery | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Post ICA | VTI | meas | m_fet_postIca | r_VTI | unilateral | 125196\20354-7\C12104-04 | 99PMSBLUS, C12104-04, Posterior Internal Carotid Artery | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Post ICA | HR | meas | m_fet_postIca | r_HeartRate | unilateral | 125196\8867-4\C12104-04 | 99PMSBLUS, C12104-04, Posterior Internal Carotid Artery | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | Post ICA | Dist | meas | m_fet_postIca | r_Distance | unilateral | 125196\C12104-04\G-0364 | 99PMSBLUS, C12104-04, Posterior Internal Carotid Artery | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | R ICA | ? | meas | m_fet_rtIca | r_DopAngle | unilateral | 125196\C12104-05\C12122-01 | 99PMSBLUS, C12104-05, Right Internal Carotid Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | R ICA | EDV | meas | m_fet_rtIca | r_EDV | unilateral | 125196\11653-3\C12104-05 | 99PMSBLUS, C12104-05, Right Internal Carotid Artery | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R ICA | MDV | meas | m_fet_rtIca | r_MDV | unilateral | 125196\11665-7\C12104-05 | 99PMSBLUS, C12104-05, Right Internal Carotid Artery | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R ICA | TAPV | meas | m_fet_rtIca | r_TAPV | unilateral | 125196\11692-1\C12104-05 | 99PMSBLUS, C12104-05, Right Internal Carotid Artery | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R ICA | PSV | meas | m_fet_rtIca | r_PSV | unilateral | 125196\11726-7\C12104-05 | 99PMSBLUS, C12104-05, Right Internal Carotid Artery | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R ICA | PI | meas | m_fet_rtIca | r_PI | unilateral | 125196\12008-9\C12104-05 | 99PMSBLUS, C12104-05, Right Internal Carotid Artery | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | R ICA | RI | meas | m_fet_rtIca | r_RI | unilateral | 125196\12023-8\C12104-05 | 99PMSBLUS, C12104-05, Right Internal Carotid Artery | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | R ICA | S/D | meas | m_fet_rtIca | r_SD | unilateral | 125196\12144-2\C12104-05 | 99PMSBLUS, C12104-05, Right Internal Carotid Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | R ICA | AI | meas | m_fet_rtIca | r_AccelIndex | unilateral | 125196\20167-3\C12104-05 | 99PMSBLUS, C12104-05, Right Internal Carotid Artery | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | R ICA | AT | meas | m_fet_rtIca | r_AccelTime | unilateral | 125196\20168-1\C12104-05 | 99PMSBLUS, C12104-05, Right Internal Carotid Artery | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | R ICA | DT | meas | m_fet_rtIca | r_DecelTime | unilateral | 125196\20217-6\C12104-05 | 99PMSBLUS, C12104-05, Right Internal Carotid Artery | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | R ICA | TAMV | meas | m_fet_rtIca | r_TAMV | unilateral | 125196\20352-1\C12104-05 | 99PMSBLUS, C12104-05, Right Internal Carotid Artery | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R ICA | VTI | meas | m_fet_rtIca | r_VTI | unilateral | 125196\20354-7\C12104-05 | 99PMSBLUS, C12104-05, Right Internal Carotid Artery | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | R ICA | HR | meas | m_fet_rtIca | r_HeartRate | unilateral | 125196\8867-4\C12104-05 | 99PMSBLUS, C12104-05, Right Internal Carotid Artery | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | R ICA | Dist | meas | m_fet_rtIca | r_Distance | unilateral | 125196\C12104-05\G-0364 | 99PMSBLUS, C12104-05, Right Internal Carotid Artery | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | L ICA | ? | meas | m_fet_ltIca | r_DopAngle | unilateral | 125196\C12104-06\C12122-01 | 99PMSBLUS, C12104-06, Left Internal Carotid Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | L ICA | EDV | meas | m_fet_ltIca | r_EDV | unilateral | 125196\11653-3\C12104-06 | 99PMSBLUS, C12104-06, Left Internal Carotid Artery | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L ICA | MDV | meas | m_fet_ltIca | r_MDV | unilateral | 125196\11665-7\C12104-06 | 99PMSBLUS, C12104-06, Left Internal Carotid Artery | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L ICA | TAPV | meas | m_fet_ltIca | r_TAPV | unilateral | 125196\11692-1\C12104-06 | 99PMSBLUS, C12104-06, Left Internal Carotid Artery | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L ICA | PSV | meas | m_fet_ltIca | r_PSV | unilateral | 125196\11726-7\C12104-06 | 99PMSBLUS, C12104-06, Left Internal Carotid Artery | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L ICA | PI | meas | m_fet_ltIca | r_PI | unilateral | 125196\12008-9\C12104-06 | 99PMSBLUS, C12104-06, Left Internal Carotid Artery | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | L ICA | RI | meas | m_fet_ltIca | r_RI | unilateral | 125196\12023-8\C12104-06 | 99PMSBLUS, C12104-06, Left Internal Carotid Artery | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | L ICA | S/D | meas | m_fet_ltIca | r_SD | unilateral | 125196\12144-2\C12104-06 | 99PMSBLUS, C12104-06, Left Internal Carotid Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | L ICA | AI | meas | m_fet_ltIca | r_AccelIndex | unilateral | 125196\20167-3\C12104-06 | 99PMSBLUS, C12104-06, Left Internal Carotid Artery | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | L ICA | AT | meas | m_fet_ltIca | r_AccelTime | unilateral | 125196\20168-1\C12104-06 | 99PMSBLUS, C12104-06, Left Internal Carotid Artery | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | L ICA | DT | meas | m_fet_ltIca | r_DecelTime | unilateral | 125196\20217-6\C12104-06 | 99PMSBLUS, C12104-06, Left Internal Carotid Artery | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | L ICA | TAMV | meas | m_fet_ltIca | r_TAMV | unilateral | 125196\20352-1\C12104-06 | 99PMSBLUS, C12104-06, Left Internal Carotid Artery | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L ICA | VTI | meas | m_fet_ltIca | r_VTI | unilateral | 125196\20354-7\C12104-06 | 99PMSBLUS, C12104-06, Left Internal Carotid Artery | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | L ICA | HR | meas | m_fet_ltIca | r_HeartRate | unilateral | 125196\8867-4\C12104-06 | 99PMSBLUS, C12104-06, Left Internal Carotid Artery | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | L ICA | Dist | meas | m_fet_ltIca | r_Distance | unilateral | 125196\C12104-06\G-0364 | 99PMSBLUS, C12104-06, Left Internal Carotid Artery | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | R MCA | ? | meas | m_fet_rtMca | r_DopAngle | unilateral | 125196\C12105-06\C12122-01 | 99PMSBLUS, C12105-06, Right Middle Cerebral Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | R MCA | EDV | meas | m_fet_rtMca | r_EDV | unilateral | 125196\11653-3\C12105-06 | 99PMSBLUS, C12105-06, Right Middle Cerebral Artery | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R MCA | MDV | meas | m_fet_rtMca | r_MDV | unilateral | 125196\11665-7\C12105-06 | 99PMSBLUS, C12105-06, Right Middle Cerebral Artery | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R MCA | TAPV | meas | m_fet_rtMca | r_TAPV | unilateral | 125196\11692-1\C12105-06 | 99PMSBLUS, C12105-06, Right Middle Cerebral Artery | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R MCA | PSV | meas | m_fet_rtMca | r_PSV | unilateral | 125196\11726-7\C12105-06 | 99PMSBLUS, C12105-06, Right Middle Cerebral Artery | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R MCA | PI | meas | m_fet_rtMca | r_PI | unilateral | 125196\12008-9\C12105-06 | 99PMSBLUS, C12105-06, Right Middle Cerebral Artery | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | R MCA | RI | meas | m_fet_rtMca | r_RI | unilateral | 125196\12023-8\C12105-06 | 99PMSBLUS, C12105-06, Right Middle Cerebral Artery | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | R MCA | S/D | meas | m_fet_rtMca | r_SD | unilateral | 125196\12144-2\C12105-06 | 99PMSBLUS, C12105-06, Right Middle Cerebral Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | R MCA | AI | meas | m_fet_rtMca | r_AccelIndex | unilateral | 125196\20167-3\C12105-06 | 99PMSBLUS, C12105-06, Right Middle Cerebral Artery | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | R MCA | AT | meas | m_fet_rtMca | r_AccelTime | unilateral | 125196\20168-1\C12105-06 | 99PMSBLUS, C12105-06, Right Middle Cerebral Artery | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | R MCA | DT | meas | m_fet_rtMca | r_DecelTime | unilateral | 125196\20217-6\C12105-06 | 99PMSBLUS, C12105-06, Right Middle Cerebral Artery | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | R MCA | TAMV | meas | m_fet_rtMca | r_TAMV | unilateral | 125196\20352-1\C12105-06 | 99PMSBLUS, C12105-06, Right Middle Cerebral Artery | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R MCA | VTI | meas | m_fet_rtMca | r_VTI | unilateral | 125196\20354-7\C12105-06 | 99PMSBLUS, C12105-06, Right Middle Cerebral Artery | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | R MCA | HR | meas | m_fet_rtMca | r_HeartRate | unilateral | 125196\8867-4\C12105-06 | 99PMSBLUS, C12105-06, Right Middle Cerebral Artery | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | R MCA | Dist | meas | m_fet_rtMca | r_Distance | unilateral | 125196\C12105-06\G-0364 | 99PMSBLUS, C12105-06, Right Middle Cerebral Artery | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | L MCA | ? | meas | m_fet_ltMca | r_DopAngle | unilateral | 125196\C12105-07\C12122-01 | 99PMSBLUS, C12105-07, Left Middle Cerebral Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | L MCA | EDV | meas | m_fet_ltMca | r_EDV | unilateral | 125196\11653-3\C12105-07 | 99PMSBLUS, C12105-07, Left Middle Cerebral Artery | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L MCA | MDV | meas | m_fet_ltMca | r_MDV | unilateral | 125196\11665-7\C12105-07 | 99PMSBLUS, C12105-07, Left Middle Cerebral Artery | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L MCA | TAPV | meas | m_fet_ltMca | r_TAPV | unilateral | 125196\11692-1\C12105-07 | 99PMSBLUS, C12105-07, Left Middle Cerebral Artery | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L MCA | PSV | meas | m_fet_ltMca | r_PSV | unilateral | 125196\11726-7\C12105-07 | 99PMSBLUS, C12105-07, Left Middle Cerebral Artery | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L MCA | PI | meas | m_fet_ltMca | r_PI | unilateral | 125196\12008-9\C12105-07 | 99PMSBLUS, C12105-07, Left Middle Cerebral Artery | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | L MCA | RI | meas | m_fet_ltMca | r_RI | unilateral | 125196\12023-8\C12105-07 | 99PMSBLUS, C12105-07, Left Middle Cerebral Artery | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | L MCA | S/D | meas | m_fet_ltMca | r_SD | unilateral | 125196\12144-2\C12105-07 | 99PMSBLUS, C12105-07, Left Middle Cerebral Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | L MCA | AI | meas | m_fet_ltMca | r_AccelIndex | unilateral | 125196\20167-3\C12105-07 | 99PMSBLUS, C12105-07, Left Middle Cerebral Artery | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | L MCA | AT | meas | m_fet_ltMca | r_AccelTime | unilateral | 125196\20168-1\C12105-07 | 99PMSBLUS, C12105-07, Left Middle Cerebral Artery | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | L MCA | DT | meas | m_fet_ltMca | r_DecelTime | unilateral | 125196\20217-6\C12105-07 | 99PMSBLUS, C12105-07, Left Middle Cerebral Artery | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | L MCA | TAMV | meas | m_fet_ltMca | r_TAMV | unilateral | 125196\20352-1\C12105-07 | 99PMSBLUS, C12105-07, Left Middle Cerebral Artery | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L MCA | VTI | meas | m_fet_ltMca | r_VTI | unilateral | 125196\20354-7\C12105-07 | 99PMSBLUS, C12105-07, Left Middle Cerebral Artery | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | L MCA | HR | meas | m_fet_ltMca | r_HeartRate | unilateral | 125196\8867-4\C12105-07 | 99PMSBLUS, C12105-07, Left Middle Cerebral Artery | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | L MCA | Dist | meas | m_fet_ltMca | r_Distance | unilateral | 125196\C12105-07\G-0364 | 99PMSBLUS, C12105-07, Left Middle Cerebral Artery | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | R Renal A | ? | meas | m_fet_rtRenArt | r_DopAngle | unilateral | 125196\C12115-01\C12122-01 | 99PMSBLUS, C12115-01, Right Renal Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | R Renal A | EDV | meas | m_fet_rtRenArt | r_EDV | unilateral | 125196\11653-3\C12115-01 | 99PMSBLUS, C12115-01, Right Renal Artery | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R Renal A | MDV | meas | m_fet_rtRenArt | r_MDV | unilateral | 125196\11665-7\C12115-01 | 99PMSBLUS, C12115-01, Right Renal Artery | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R Renal A | TAPV | meas | m_fet_rtRenArt | r_TAPV | unilateral | 125196\11692-1\C12115-01 | 99PMSBLUS, C12115-01, Right Renal Artery | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R Renal A | PSV | meas | m_fet_rtRenArt | r_PSV | unilateral | 125196\11726-7\C12115-01 | 99PMSBLUS, C12115-01, Right Renal Artery | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R Renal A | PI | meas | m_fet_rtRenArt | r_PI | unilateral | 125196\12008-9\C12115-01 | 99PMSBLUS, C12115-01, Right Renal Artery | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | R Renal A | RI | meas | m_fet_rtRenArt | r_RI | unilateral | 125196\12023-8\C12115-01 | 99PMSBLUS, C12115-01, Right Renal Artery | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | R Renal A | S/D | meas | m_fet_rtRenArt | r_SD | unilateral | 125196\12144-2\C12115-01 | 99PMSBLUS, C12115-01, Right Renal Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | R Renal A | AI | meas | m_fet_rtRenArt | r_AccelIndex | unilateral | 125196\20167-3\C12115-01 | 99PMSBLUS, C12115-01, Right Renal Artery | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | R Renal A | AT | meas | m_fet_rtRenArt | r_AccelTime | unilateral | 125196\20168-1\C12115-01 | 99PMSBLUS, C12115-01, Right Renal Artery | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | R Renal A | DT | meas | m_fet_rtRenArt | r_DecelTime | unilateral | 125196\20217-6\C12115-01 | 99PMSBLUS, C12115-01, Right Renal Artery | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | R Renal A | TAMV | meas | m_fet_rtRenArt | r_TAMV | unilateral | 125196\20352-1\C12115-01 | 99PMSBLUS, C12115-01, Right Renal Artery | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R Renal A | VTI | meas | m_fet_rtRenArt | r_VTI | unilateral | 125196\20354-7\C12115-01 | 99PMSBLUS, C12115-01, Right Renal Artery | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | R Renal A | HR | meas | m_fet_rtRenArt | r_HeartRate | unilateral | 125196\8867-4\C12115-01 | 99PMSBLUS, C12115-01, Right Renal Artery | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | R Renal A | Dist | meas | m_fet_rtRenArt | r_Distance | unilateral | 125196\C12115-01\G-0364 | 99PMSBLUS, C12115-01, Right Renal Artery | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | L Renal A | ? | meas | m_fet_ltRenArt | r_DopAngle | unilateral | 125196\C12115-02\C12122-01 | 99PMSBLUS, C12115-02, Left Renal Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | L Renal A | EDV | meas | m_fet_ltRenArt | r_EDV | unilateral | 125196\11653-3\C12115-02 | 99PMSBLUS, C12115-02, Left Renal Artery | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L Renal A | MDV | meas | m_fet_ltRenArt | r_MDV | unilateral | 125196\11665-7\C12115-02 | 99PMSBLUS, C12115-02, Left Renal Artery | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L Renal A | TAPV | meas | m_fet_ltRenArt | r_TAPV | unilateral | 125196\11692-1\C12115-02 | 99PMSBLUS, C12115-02, Left Renal Artery | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L Renal A | PSV | meas | m_fet_ltRenArt | r_PSV | unilateral | 125196\11726-7\C12115-02 | 99PMSBLUS, C12115-02, Left Renal Artery | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L Renal A | PI | meas | m_fet_ltRenArt | r_PI | unilateral | 125196\12008-9\C12115-02 | 99PMSBLUS, C12115-02, Left Renal Artery | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | L Renal A | RI | meas | m_fet_ltRenArt | r_RI | unilateral | 125196\12023-8\C12115-02 | 99PMSBLUS, C12115-02, Left Renal Artery | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | L Renal A | S/D | meas | m_fet_ltRenArt | r_SD | unilateral | 125196\12144-2\C12115-02 | 99PMSBLUS, C12115-02, Left Renal Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | L Renal A | AI | meas | m_fet_ltRenArt | r_AccelIndex | unilateral | 125196\20167-3\C12115-02 | 99PMSBLUS, C12115-02, Left Renal Artery | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | L Renal A | AT | meas | m_fet_ltRenArt | r_AccelTime | unilateral | 125196\20168-1\C12115-02 | 99PMSBLUS, C12115-02, Left Renal Artery | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | L Renal A | DT | meas | m_fet_ltRenArt | r_DecelTime | unilateral | 125196\20217-6\C12115-02 | 99PMSBLUS, C12115-02, Left Renal Artery | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | L Renal A | TAMV | meas | m_fet_ltRenArt | r_TAMV | unilateral | 125196\20352-1\C12115-02 | 99PMSBLUS, C12115-02, Left Renal Artery | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L Renal A | VTI | meas | m_fet_ltRenArt | r_VTI | unilateral | 125196\20354-7\C12115-02 | 99PMSBLUS, C12115-02, Left Renal Artery | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | L Renal A | HR | meas | m_fet_ltRenArt | r_HeartRate | unilateral | 125196\8867-4\C12115-02 | 99PMSBLUS, C12115-02, Left Renal Artery | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | L Renal A | Dist | meas | m_fet_ltRenArt | r_Distance | unilateral | 125196\C12115-02\G-0364 | 99PMSBLUS, C12115-02, Left Renal Artery | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | R Ovarian A | ? | meas | m_fet_rtovary | r_DopAngle | unilateral | 125196\C12122-01\C12140-01 | 99PMSBLUS, C12140-01, Right Ovarian Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | R Ovarian A | EDV | meas | m_fet_rtovary | r_EDV | unilateral | 125196\11653-3\C12140-01 | 99PMSBLUS, C12140-01, Right Ovarian Artery | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R Ovarian A | MDV | meas | m_fet_rtovary | r_MDV | unilateral | 125196\11665-7\C12140-01 | 99PMSBLUS, C12140-01, Right Ovarian Artery | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R Ovarian A | TAPV | meas | m_fet_rtovary | r_TAPV | unilateral | 125196\11692-1\C12140-01 | 99PMSBLUS, C12140-01, Right Ovarian Artery | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R Ovarian A | PSV | meas | m_fet_rtovary | r_PSV | unilateral | 125196\11726-7\C12140-01 | 99PMSBLUS, C12140-01, Right Ovarian Artery | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R Ovarian A | PI | meas | m_fet_rtovary | r_PI | unilateral | 125196\12008-9\C12140-01 | 99PMSBLUS, C12140-01, Right Ovarian Artery | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | R Ovarian A | RI | meas | m_fet_rtovary | r_RI | unilateral | 125196\12023-8\C12140-01 | 99PMSBLUS, C12140-01, Right Ovarian Artery | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | R Ovarian A | S/D | meas | m_fet_rtovary | r_SD | unilateral | 125196\12144-2\C12140-01 | 99PMSBLUS, C12140-01, Right Ovarian Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | R Ovarian A | AI | meas | m_fet_rtovary | r_AccelIndex | unilateral | 125196\20167-3\C12140-01 | 99PMSBLUS, C12140-01, Right Ovarian Artery | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | R Ovarian A | AT | meas | m_fet_rtovary | r_AccelTime | unilateral | 125196\20168-1\C12140-01 | 99PMSBLUS, C12140-01, Right Ovarian Artery | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | R Ovarian A | DT | meas | m_fet_rtovary | r_DecelTime | unilateral | 125196\20217-6\C12140-01 | 99PMSBLUS, C12140-01, Right Ovarian Artery | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | R Ovarian A | TAMV | meas | m_fet_rtovary | r_TAMV | unilateral | 125196\20352-1\C12140-01 | 99PMSBLUS, C12140-01, Right Ovarian Artery | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R Ovarian A | VTI | meas | m_fet_rtovary | r_VTI | unilateral | 125196\20354-7\C12140-01 | 99PMSBLUS, C12140-01, Right Ovarian Artery | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | R Ovarian A | HR | meas | m_fet_rtovary | r_HeartRate | unilateral | 125196\8867-4\C12140-01 | 99PMSBLUS, C12140-01, Right Ovarian Artery | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | R Ovarian A | Dist | meas | m_fet_rtovary | r_Distance | unilateral | 125196\C12140-01\G-0364 | 99PMSBLUS, C12140-01, Right Ovarian Artery | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | L Ovarian A | ? | meas | m_fet_ltovary | r_DopAngle | unilateral | 125196\C12122-01\C12140-02 | 99PMSBLUS, C12140-02, Left Ovarian Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | L Ovarian A | EDV | meas | m_fet_ltovary | r_EDV | unilateral | 125196\11653-3\C12140-02 | 99PMSBLUS, C12140-02, Left Ovarian Artery | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L Ovarian A | MDV | meas | m_fet_ltovary | r_MDV | unilateral | 125196\11665-7\C12140-02 | 99PMSBLUS, C12140-02, Left Ovarian Artery | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L Ovarian A | TAPV | meas | m_fet_ltovary | r_TAPV | unilateral | 125196\11692-1\C12140-02 | 99PMSBLUS, C12140-02, Left Ovarian Artery | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L Ovarian A | PSV | meas | m_fet_ltovary | r_PSV | unilateral | 125196\11726-7\C12140-02 | 99PMSBLUS, C12140-02, Left Ovarian Artery | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L Ovarian A | PI | meas | m_fet_ltovary | r_PI | unilateral | 125196\12008-9\C12140-02 | 99PMSBLUS, C12140-02, Left Ovarian Artery | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | L Ovarian A | RI | meas | m_fet_ltovary | r_RI | unilateral | 125196\12023-8\C12140-02 | 99PMSBLUS, C12140-02, Left Ovarian Artery | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | L Ovarian A | S/D | meas | m_fet_ltovary | r_SD | unilateral | 125196\12144-2\C12140-02 | 99PMSBLUS, C12140-02, Left Ovarian Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | L Ovarian A | AI | meas | m_fet_ltovary | r_AccelIndex | unilateral | 125196\20167-3\C12140-02 | 99PMSBLUS, C12140-02, Left Ovarian Artery | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | L Ovarian A | AT | meas | m_fet_ltovary | r_AccelTime | unilateral | 125196\20168-1\C12140-02 | 99PMSBLUS, C12140-02, Left Ovarian Artery | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | L Ovarian A | DT | meas | m_fet_ltovary | r_DecelTime | unilateral | 125196\20217-6\C12140-02 | 99PMSBLUS, C12140-02, Left Ovarian Artery | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | L Ovarian A | TAMV | meas | m_fet_ltovary | r_TAMV | unilateral | 125196\20352-1\C12140-02 | 99PMSBLUS, C12140-02, Left Ovarian Artery | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L Ovarian A | VTI | meas | m_fet_ltovary | r_VTI | unilateral | 125196\20354-7\C12140-02 | 99PMSBLUS, C12140-02, Left Ovarian Artery | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | L Ovarian A | HR | meas | m_fet_ltovary | r_HeartRate | unilateral | 125196\8867-4\C12140-02 | 99PMSBLUS, C12140-02, Left Ovarian Artery | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | L Ovarian A | Dist | meas | m_fet_ltovary | r_Distance | unilateral | 125196\C12140-02\G-0364 | 99PMSBLUS, C12140-02, Left Ovarian Artery | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | R Uterine A | ? | meas | m_fet_rtUterArt | r_DopAngle | unilateral | 125196\C12122-01\C12140-03 | 99PMSBLUS, C12140-03, Right Uterine Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | R Uterine A | EDV | meas | m_fet_rtUterArt | r_EDV | unilateral | 125196\11653-3\C12140-03 | 99PMSBLUS, C12140-03, Right Uterine Artery | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R Uterine A | MDV | meas | m_fet_rtUterArt | r_MDV | unilateral | 125196\11665-7\C12140-03 | 99PMSBLUS, C12140-03, Right Uterine Artery | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R Uterine A | TAPV | meas | m_fet_rtUterArt | r_TAPV | unilateral | 125196\11692-1\C12140-03 | 99PMSBLUS, C12140-03, Right Uterine Artery | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R Uterine A | PSV | meas | m_fet_rtUterArt | r_PSV | unilateral | 125196\11726-7\C12140-03 | 99PMSBLUS, C12140-03, Right Uterine Artery | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R Uterine A | PI | meas | m_fet_rtUterArt | r_PI | unilateral | 125196\12008-9\C12140-03 | 99PMSBLUS, C12140-03, Right Uterine Artery | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | R Uterine A | RI | meas | m_fet_rtUterArt | r_RI | unilateral | 125196\12023-8\C12140-03 | 99PMSBLUS, C12140-03, Right Uterine Artery | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | R Uterine A | S/D | meas | m_fet_rtUterArt | r_SD | unilateral | 125196\12144-2\C12140-03 | 99PMSBLUS, C12140-03, Right Uterine Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | R Uterine A | AI | meas | m_fet_rtUterArt | r_AccelIndex | unilateral | 125196\20167-3\C12140-03 | 99PMSBLUS, C12140-03, Right Uterine Artery | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | R Uterine A | AT | meas | m_fet_rtUterArt | r_AccelTime | unilateral | 125196\20168-1\C12140-03 | 99PMSBLUS, C12140-03, Right Uterine Artery | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | R Uterine A | DT | meas | m_fet_rtUterArt | r_DecelTime | unilateral | 125196\20217-6\C12140-03 | 99PMSBLUS, C12140-03, Right Uterine Artery | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | R Uterine A | TAMV | meas | m_fet_rtUterArt | r_TAMV | unilateral | 125196\20352-1\C12140-03 | 99PMSBLUS, C12140-03, Right Uterine Artery | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | R Uterine A | VTI | meas | m_fet_rtUterArt | r_VTI | unilateral | 125196\20354-7\C12140-03 | 99PMSBLUS, C12140-03, Right Uterine Artery | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | R Uterine A | HR | meas | m_fet_rtUterArt | r_HeartRate | unilateral | 125196\8867-4\C12140-03 | 99PMSBLUS, C12140-03, Right Uterine Artery | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | R Uterine A | Dist | meas | m_fet_rtUterArt | r_Distance | unilateral | 125196\C12140-03\G-0364 | 99PMSBLUS, C12140-03, Right Uterine Artery | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | L Uterine A | ? | meas | m_fet_ltUterArt | r_DopAngle | unilateral | 125196\C12122-01\C12140-04 | 99PMSBLUS, C12140-04, Left Uterine Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | L Uterine A | EDV | meas | m_fet_ltUterArt | r_EDV | unilateral | 125196\11653-3\C12140-04 | 99PMSBLUS, C12140-04, Left Uterine Artery | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L Uterine A | MDV | meas | m_fet_ltUterArt | r_MDV | unilateral | 125196\11665-7\C12140-04 | 99PMSBLUS, C12140-04, Left Uterine Artery | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L Uterine A | TAPV | meas | m_fet_ltUterArt | r_TAPV | unilateral | 125196\11692-1\C12140-04 | 99PMSBLUS, C12140-04, Left Uterine Artery | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L Uterine A | PSV | meas | m_fet_ltUterArt | r_PSV | unilateral | 125196\11726-7\C12140-04 | 99PMSBLUS, C12140-04, Left Uterine Artery | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L Uterine A | PI | meas | m_fet_ltUterArt | r_PI | unilateral | 125196\12008-9\C12140-04 | 99PMSBLUS, C12140-04, Left Uterine Artery | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | L Uterine A | RI | meas | m_fet_ltUterArt | r_RI | unilateral | 125196\12023-8\C12140-04 | 99PMSBLUS, C12140-04, Left Uterine Artery | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | L Uterine A | S/D | meas | m_fet_ltUterArt | r_SD | unilateral | 125196\12144-2\C12140-04 | 99PMSBLUS, C12140-04, Left Uterine Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | L Uterine A | AI | meas | m_fet_ltUterArt | r_AccelIndex | unilateral | 125196\20167-3\C12140-04 | 99PMSBLUS, C12140-04, Left Uterine Artery | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | L Uterine A | AT | meas | m_fet_ltUterArt | r_AccelTime | unilateral | 125196\20168-1\C12140-04 | 99PMSBLUS, C12140-04, Left Uterine Artery | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | L Uterine A | DT | meas | m_fet_ltUterArt | r_DecelTime | unilateral | 125196\20217-6\C12140-04 | 99PMSBLUS, C12140-04, Left Uterine Artery | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | L Uterine A | TAMV | meas | m_fet_ltUterArt | r_TAMV | unilateral | 125196\20352-1\C12140-04 | 99PMSBLUS, C12140-04, Left Uterine Artery | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | L Uterine A | VTI | meas | m_fet_ltUterArt | r_VTI | unilateral | 125196\20354-7\C12140-04 | 99PMSBLUS, C12140-04, Left Uterine Artery | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | L Uterine A | HR | meas | m_fet_ltUterArt | r_HeartRate | unilateral | 125196\8867-4\C12140-04 | 99PMSBLUS, C12140-04, Left Uterine Artery | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | L Uterine A | Dist | meas | m_fet_ltUterArt | r_Distance | unilateral | 125196\C12140-04\G-0364 | 99PMSBLUS, C12140-04, Left Uterine Artery | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Foramen Ovale | Max PG | meas | m_fet_foreOvale | r_MaxPG | unilateral | 125196\20247-3\D4-31012 | SRT, D4-31012, Patent Foramen Ovale | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Fetal Cardiac Ultrasound Report | Fetal Echo | Foramen Ovale | Vmax | meas | m_fet_foreOvale | r_Vmax | unilateral | 125196\20351-3\D4-31012 | SRT, D4-31012, Patent Foramen Ovale | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | SVC | PG | meas | m_fet_svcDop | r_PG | unilateral | 125196\C12122-10\M-2460D | SRT, M-2460D, Right Superior vena cava | 99PMSBLUS, C12122-10, Pressure Gradient |  |  | UCUM, mmHg, mmHg
Fetal Cardiac Ultrasound Report | Fetal Echo | SVC | Vel | meas | m_fet_svcDop | r_Vel | unilateral | 125196\20351-3\M-2460D | SRT, M-2460D, Right Superior vena cava | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | SVC Diam (2D) | Dist | meas | m_fet_svc2D | r_Distance | unilateral | 125196\G-03A2\M-02550\M-2460D | SRT, M-2460D, Right Superior vena cava | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | RA Wall Diam (MM) | Dist | meas | m_fet_rtAtrWMM | r_Distance | unilateral | 125196\59090-1\G-0394\T-32200 | SRT, T-32200, Right Atrium | LN, 59090-1, ROI Internal Dimension by US |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | RA Width | Dist | meas | m_fet_raWidth | r_Distance | unilateral | 125196\G-03A2\G-A220\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A220, Width |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | RA Length | Dist | meas | m_fet_raLen | r_Distance | unilateral | 125196\G-03A2\G-A22A\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A22A, Length |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LA Dimen (MM) | Dist | meas | m_fet_laDimenMM | r_Distance | unilateral | 125196\29469-4\G-0394\T-32300 | SRT, T-32300, Left Atrium | LN, 29469-4, Left Atrium Antero-posterior Systolic Dimension |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LA Dimen (2D) | Dist | meas | m_fet_laDimen2D | r_Distance | unilateral | 125196\29469-4\G-03A2\T-32300 | SRT, T-32300, Left Atrium | LN, 29469-4, Left Atrium Antero-posterior Systolic Dimension |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LA Wall Diam (MM) | Dist | meas | m_fet_ltAtrWMM | r_Distance | unilateral | 125196\59090-1\G-0394\T-32300 | SRT, T-32300, Left Atrium | LN, 59090-1, ROI Internal Dimension by US |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LA Width | Dist | meas | m_fet_laWidth | r_Distance | unilateral | 125196\G-03A2\G-A220\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A220, Width |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LA Length | Dist | meas | m_fet_laLen | r_Distance | unilateral | 125196\G-03A2\G-A22A\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A22A, Length |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | IVSd (MM) | Dist | meas | m_fet_ivsdMM | r_Distance | unilateral | 125196\59089-3\F-32011\G-0394\T-32410 | SRT, T-32410, Interventricular septum | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Fetal Cardiac Ultrasound Report | Fetal Echo | IVSs (MM) | Dist | meas | m_fet_ivssMM | r_Distance | unilateral | 125196\59089-3\G-0394\R-FAB5B\T-32410 | SRT, T-32410, Interventricular septum | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | SRT, R-FAB5B, End Systole
Fetal Cardiac Ultrasound Report | Fetal Echo | IVSd (2D) | Dist | meas | m_fet_ivsd2D | r_Distance | unilateral | 125196\59089-3\F-32011\G-03A2\T-32410 | SRT, T-32410, Interventricular septum | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Fetal Cardiac Ultrasound Report | Fetal Echo | IVSs (2D) | Dist | meas | m_fet_ivss2D | r_Distance | unilateral | 125196\59089-3\G-03A2\R-FAB5B\T-32410 | SRT, T-32410, Interventricular septum | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole
Fetal Cardiac Ultrasound Report | Fetal Echo | IVS Width (2D) | Dist | meas | m_fet_ivsWall | r_Distance | unilateral | 125196\G-03A2\G-A220\T-32410 | SRT, T-32410, Interventricular septum | SRT, G-A220, Width |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | RVOT VTI | Vmean | meas | m_fet_rvotVTI | r_CardTAPV | unilateral | 125196\11692-1\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s | SRT, T-32550, Right Ventricle Outflow Tract
Fetal Cardiac Ultrasound Report | Fetal Echo | RVOT Acc Time | Time | meas | m_fet_rvotAccTime | r_DopTime | unilateral | 125196\20168-1\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms | SRT, T-32550, Right Ventricle Outflow Tract
Fetal Cardiac Ultrasound Report | Fetal Echo | RV ET | Time | meas | m_fet_RV_Eject_Tm | r_MMTime | unilateral | 125196\20222-6\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 20222-6, Ejection Time |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | RVOT Vmax | Max PG | meas | m_fet_rvotVmax | r_MaxPG | unilateral | 125196\20247-3\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-32550, Right Ventricle Outflow Tract
Fetal Cardiac Ultrasound Report | Fetal Echo | RVOT VTI | Mean PG | meas | m_fet_rvotVTI | r_CardPGmean | unilateral | 125196\20256-4\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-32550, Right Ventricle Outflow Tract
Fetal Cardiac Ultrasound Report | Fetal Echo | RVOT Vmax | Vmax | meas | m_fet_rvotVmax | r_Vmax | unilateral | 125196\20351-3\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, T-32550, Right Ventricle Outflow Tract
Fetal Cardiac Ultrasound Report | Fetal Echo | RVOT VTI | VTI | meas | m_fet_rvotVTI | r_VTI | unilateral | 125196\20354-7\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm | SRT, T-32550, Right Ventricle Outflow Tract
Fetal Cardiac Ultrasound Report | Fetal Echo | RV PEP | Time | meas | m_fet_RV_PreEject_Tm | r_MMTime | unilateral | 125196\59085-1\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 59085-1, Pre-Ejection Period |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | RVAWd (MM) | Dist | meas | m_fet_rvawdMM | r_Distance | unilateral | 125196\59089-3\F-32011\G-0394\R-4210B\T-32500 | SRT, T-32500, Right Ventricle | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm | SRT, R-4210B, Anterior Wall | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Fetal Cardiac Ultrasound Report | Fetal Echo | RV Wall Width (2D) | Dist | meas | m_fet_rvWall | r_Distance | unilateral | 125196\59089-3\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | RVAWd (2D) | Dist | meas | m_fet_rvawd2D | r_Distance | unilateral | 125196\59089-3\F-32011\G-03A2\R-4210B\T-32500 | SRT, T-32500, Right Ventricle | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm | SRT, R-4210B, Anterior Wall | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Fetal Cardiac Ultrasound Report | Fetal Echo | RVIDd (MM) | Dist | meas | m_fet_rviddMM | r_Distance | unilateral | 125196\59090-1\F-32011\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 59090-1, ROI Internal Dimension by US |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Fetal Cardiac Ultrasound Report | Fetal Echo | RVIDd (2D) | Dist | meas | m_fet_rvidd2D | r_Distance | unilateral | 125196\59090-1\F-32011\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | LN, 59090-1, ROI Internal Dimension by US |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Fetal Cardiac Ultrasound Report | Fetal Echo | RVOT Diam | Dist | meas | m_fet_rvotDiam | r_Distance | unilateral | 125196\G-038F\G-03A2\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, G-038F, Cardiovascular Orifice Diameter |  |  | UCUM, mm, mm | SRT, T-32550, Right Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | RV Width | Dist | meas | m_fet_rvWidth | r_Distance | unilateral | 125196\G-03A2\G-A220\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-A220, Width |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | RV Length | Dist | meas | m_fet_rvLen | r_Distance | unilateral | 125196\G-03A2\G-A22A\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-A22A, Length |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LVOT VTI | Vmean | meas | m_fet_lvotVTI | r_CardTAPV | unilateral | 125196\11692-1\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s | SRT, T-32650, Left Ventricle Outflow Tract
Fetal Cardiac Ultrasound Report | Fetal Echo | Med Acc Time | Time | meas | m_f_TDIMedAT | r_DopTime | unilateral | 125196\20168-1\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | Lat Acc Time | Time | meas | m_f_TDILatAT | r_DopTime | unilateral | 125196\20168-1\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | LVOT Acc Time | Time | meas | m_fet_lvotAccTime | r_DopTime | unilateral | 125196\20168-1\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms | SRT, T-32650, Left Ventricle Outflow Tract
Fetal Cardiac Ultrasound Report | Fetal Echo | Med Dec Time | Time | meas | m_f_TDIMedDT | r_DopTime | unilateral | 125196\20217-6\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | Lat Dec Time | Time | meas | m_f_TDILatDT | r_DopTime | unilateral | 125196\20217-6\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | LV ET | Time | meas | m_fet_LV_Eject_Tm | r_MMTime | unilateral | 125196\20222-6\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 20222-6, Ejection Time |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LVOT Vmax | Max PG | meas | m_fet_lvotVmax | r_MaxPG | unilateral | 125196\20247-3\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-32650, Left Ventricle Outflow Tract
Fetal Cardiac Ultrasound Report | Fetal Echo | LVOT VTI | Mean PG | meas | m_fet_lvotVTI | r_CardPGmean | unilateral | 125196\20256-4\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-32650, Left Ventricle Outflow Tract
Fetal Cardiac Ultrasound Report | Fetal Echo | LVOT Vmax | Vmax | meas | m_fet_lvotVmax | r_Vmax | unilateral | 125196\20351-3\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, T-32650, Left Ventricle Outflow Tract
Fetal Cardiac Ultrasound Report | Fetal Echo | LVOT VTI | VTI | meas | m_fet_lvotVTI | r_VTI | unilateral | 125196\20354-7\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm | SRT, T-32650, Left Ventricle Outflow Tract
Fetal Cardiac Ultrasound Report | Fetal Echo | Med E` Vel | Vel | meas | m_f_MedPeakEVel | r_Vel | unilateral | 125196\59080-2\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59080-2, E-Wave Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | Lat E` Vel | Vel | meas | m_f_LatPeakEVel | r_Vel | unilateral | 125196\59080-2\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59080-2, E-Wave Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | Med IVRT | Time | meas | m_f_TDI_MedIVRT | r_DopTime | unilateral | 125196\59083-6\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59083-6, Isovolumic Relaxation Time |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | Lat IVRT | Time | meas | m_f_TDI_LatIVRT | r_DopTime | unilateral | 125196\59083-6\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59083-6, Isovolumic Relaxation Time |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | Med IVCT | Time | meas | m_f_TDI_MedIVCT | r_DopTime | unilateral | 125196\59084-4\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59084-4, Isovolumic Contraction Time |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | Lat IVCT | Time | meas | m_f_TDI_LatIVCT | r_DopTime | unilateral | 125196\59084-4\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59084-4, Isovolumic Contraction Time |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | LV PEP | Time | meas | m_fet_LV_PreEject_Tm | r_MMTime | unilateral | 125196\59085-1\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59085-1, Pre-Ejection Period |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LVPWd (MM) | Dist | meas | m_fet_lvpwdMM | r_Distance | unilateral | 125196\59089-3\F-32011\G-0394\R-42175\T-32600 | SRT, T-32600, Left Ventricle | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm | SRT, R-42175, Posterior Wall | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Fetal Cardiac Ultrasound Report | Fetal Echo | LVPWs (MM) | Dist | meas | m_fet_lvpwsMM | r_Distance | unilateral | 125196\59089-3\G-0394\R-42175\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm | SRT, R-42175, Posterior Wall | SRT, G-0394, M mode |  | SRT, R-FAB5B, End Systole
Fetal Cardiac Ultrasound Report | Fetal Echo | LV Wall Width (2D) | Dist | meas | m_fet_lvWall | r_Distance | unilateral | 125196\59089-3\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LVPWd (2D) | Dist | meas | m_fet_lvpwd2D | r_Distance | unilateral | 125196\59089-3\F-32011\G-03A2\R-42175\T-32600 | SRT, T-32600, Left Ventricle | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm | SRT, R-42175, Posterior Wall | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Fetal Cardiac Ultrasound Report | Fetal Echo | LVPWs (2D) | Dist | meas | m_fet_lvpws2D | r_Distance | unilateral | 125196\59089-3\G-03A2\R-42175\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm | SRT, R-42175, Posterior Wall | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole
Fetal Cardiac Ultrasound Report | Fetal Echo | LVIDd (MM) | Dist | meas | m_fet_lviddMM | r_Distance | unilateral | 125196\59090-1\F-32011\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59090-1, ROI Internal Dimension by US |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Fetal Cardiac Ultrasound Report | Fetal Echo | LVIDs (MM) | Dist | meas | m_fet_lvidsMM | r_Distance | unilateral | 125196\59090-1\G-0394\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59090-1, ROI Internal Dimension by US |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | SRT, R-FAB5B, End Systole
Fetal Cardiac Ultrasound Report | Fetal Echo | LVIDd (2D) | Dist | meas | m_fet_lvidd2D | r_Distance | unilateral | 125196\59090-1\F-32011\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 59090-1, ROI Internal Dimension by US |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Fetal Cardiac Ultrasound Report | Fetal Echo | LVIDs (2D) | Dist | meas | m_fet_lvids2D | r_Distance | unilateral | 125196\59090-1\G-03A2\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59090-1, ROI Internal Dimension by US |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole
Fetal Cardiac Ultrasound Report | Fetal Echo | Time to Med S | Time | meas | m_f_TimeMedPeakS | r_DopTime | unilateral | 125196\59095-0\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59095-0, Time to Left Ventricle S Tissue Velocity |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | Time to Lat S | Time | meas | m_f_TimeLatPeakS | r_DopTime | unilateral | 125196\59095-0\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59095-0, Time to Left Ventricle S Tissue Velocity |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | Time to Med E` | Time | meas | m_f_TimeMedPeakE | r_DopTime | unilateral | 125196\59096-8\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59096-8, Time to Left Ventricle E Tissue Velocity |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | Time to Lat E` | Time | meas | m_f_TimeLatPeakE | r_DopTime | unilateral | 125196\59096-8\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59096-8, Time to Left Ventricle E Tissue Velocity |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | Med E` Area | VTI | meas | m_f_AUndMedPeakE | r_VTI | unilateral | 125196\59124-8\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59124-8, LV VTI E wave US |  |  | UCUM, mm, mm | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | Lat E` Area | VTI | meas | m_f_AUndLatPeakE | r_VTI | unilateral | 125196\59124-8\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59124-8, LV VTI E wave US |  |  | UCUM, mm, mm | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | Med A` Area | VTI | meas | m_f_AUndMedPeakA | r_VTI | unilateral | 125196\59125-5\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59125-5, LV VTI A wave US |  |  | UCUM, mm, mm | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | Lat A` Area | VTI | meas | m_f_AUndLatPeakA | r_VTI | unilateral | 125196\59125-5\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59125-5, LV VTI A wave US |  |  | UCUM, mm, mm | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | Med S Vel | Vel | meas | m_f_MedPeakSVel | r_Vel | unilateral | 125196\59133-9\G-0391\P5-B0128\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59133-9, Peak Tissue Velocity |  |  | UCUM, mm/s, mm/s | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging |  | SRT, R-FAB5B, End Systole
Fetal Cardiac Ultrasound Report | Fetal Echo | Lat S Vel | Vel | meas | m_f_LatPeakSVel | r_Vel | unilateral | 125196\59133-9\G-0392\P5-B0128\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59133-9, Peak Tissue Velocity |  |  | UCUM, mm/s, mm/s | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging |  | SRT, R-FAB5B, End Systole
Fetal Cardiac Ultrasound Report | Fetal Echo | HR-LV (2D) | HR | meas | m_fet_HRLv2D | r_HeartRate | unilateral | 125196\8867-4\T-32600 | SRT, T-32600, Left Ventricle | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | HR-LV (MM) | HR | meas | m_fet_HRLvMM | r_HeartRate | unilateral | 125196\8867-4\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Med A` Vel | Vel | meas | m_f_MedPeakAVel | r_Vel | unilateral | 125196\F-32011\G-037C\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037C, LV Peak Diastolic Tissue Velocity During Atrial Systole |  |  | UCUM, mm/s, mm/s | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging |  | SRT, F-32011, End Diastole
Fetal Cardiac Ultrasound Report | Fetal Echo | Lat A` Vel | Vel | meas | m_f_LatPeakAVel | r_Vel | unilateral | 125196\F-32011\G-037C\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037C, LV Peak Diastolic Tissue Velocity During Atrial Systole |  |  | UCUM, mm/s, mm/s | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging |  | SRT, F-32011, End Diastole
Fetal Cardiac Ultrasound Report | Fetal Echo | LVOT Diam | Dist | meas | m_fet_lvotDiam | r_Distance | unilateral | 125196\G-038F\G-03A2\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, G-038F, Cardiovascular Orifice Diameter |  |  | UCUM, mm, mm | SRT, T-32650, Left Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LV Width | Dist | meas | m_fet_lvWidth | r_Distance | unilateral | 125196\G-03A2\G-A220\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A220, Width |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LV Length | Dist | meas | m_fet_lvLen | r_Distance | unilateral | 125196\G-03A2\G-A22A\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A22A, Length |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | TR VTI | Vmean | meas | m_fet_trVTI | r_CardTAPV | unilateral | 125196\11692-1\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | TV VTI | Vmean | meas | m_fet_tvVTI | r_CardTAPV | unilateral | 125196\11692-1\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | TV Acc Time | Time | meas | m_fet_tvAccTime | r_DopTime | unilateral | 125196\20168-1\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | TV Dec Time | Time | meas | m_fet_tvDecTime | r_DopTime | unilateral | 125196\20217-6\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | TR Vmax | Max PG | meas | m_fet_trVmax | r_MaxPG | unilateral | 125196\20247-3\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | TV Vmax | Max PG | meas | m_fet_tvVmax | r_MaxPG | unilateral | 125196\20247-3\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | TR VTI | Mean PG | meas | m_fet_trVTI | r_CardPGmean | unilateral | 125196\20256-4\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | TV VTI | Mean PG | meas | m_fet_tvVTI | r_CardPGmean | unilateral | 125196\20256-4\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | TR Vmax | Vmax | meas | m_fet_trVmax | r_Vmax | unilateral | 125196\20351-3\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | TV Vmax | Vmax | meas | m_fet_tvVmax | r_Vmax | unilateral | 125196\20351-3\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | TR VTI | VTI | meas | m_fet_trVTI | r_VTI | unilateral | 125196\20354-7\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | TV VTI | VTI | meas | m_fet_tvVTI | r_VTI | unilateral | 125196\20354-7\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | TV Peak E | Vel | meas | m_fet_tvPeakE | r_Vel | unilateral | 125196\59080-2\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59080-2, E-Wave Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | TV Peak A | Vel | meas | m_fet_tvPeakA | r_Vel | unilateral | 125196\59081-0\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59081-0, A-Wave Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | TV D-E Exc | Dist | meas | m_fet_TV_DE_Exc | r_Distance | unilateral | 125196\59091-9\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59091-9, D-E Excursion |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | TV A-C Interval | Time | meas | m_fet_TV_AC_Int | r_MMTime | unilateral | 125196\59103-2\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59103-2, A-C Interval |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | TV D-E Slope | Slope | meas | m_fet_TV_DE_Slope | r_MMSlope | unilateral | 125196\59127-1\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59127-1, D-E Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | TV E-F Slope | Slope | meas | m_fet_TV_EF_Slope | r_MMSlope | unilateral | 125196\59128-9\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59128-9, E-F Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | HR-TV | HR | meas | m_fet_HRTvDop | r_HeartRate | unilateral | 125196\8867-4\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | TV Annul Diam | Dist | meas | m_fet_tvAnn | r_Distance | unilateral | 125196\G-03A2\M-02550\T-3500E\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-3500E, Cardiac valve annulus | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | PV VTI | Vmean | meas | m_fet_pvVTI | r_CardTAPV | unilateral | 125196\11692-1\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | PV Acc Time | Time | meas | m_fet_pvAccTime | r_DopTime | unilateral | 125196\20168-1\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | Late Dias Slope | Slope | meas | m_fet_LateDiasSlope | r_MMSlope | unilateral | 125196\20216-8\F-32011\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20216-8, Deceleration Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Fetal Cardiac Ultrasound Report | Fetal Echo | PV Vmax | Max PG | meas | m_fet_pvVmax | r_MaxPG | unilateral | 125196\20247-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | PV VTI | Mean PG | meas | m_fet_pvVTI | r_CardPGmean | unilateral | 125196\20256-4\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | PV Vmax | Vmax | meas | m_fet_pvVmax | r_Vmax | unilateral | 125196\20351-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | PV VTI | VTI | meas | m_fet_pvVTI | r_VTI | unilateral | 125196\20354-7\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | A Wave Amp | Dist | meas | m_fet_AWaveAmp | r_Distance | unilateral | 125196\59100-8\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 59100-8, A-Wave Amplitude |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | B-C Slope | Slope | meas | m_fet_BC_Slope | r_MMSlope | unilateral | 125196\59126-3\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 59126-3, B-C Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | HR-PV | HR | meas | m_fet_HRPvDop | r_HeartRate | unilateral | 125196\8867-4\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | HR-PV (MM) | HR | meas | m_fet_HRPvMM | r_HeartRate | unilateral | 125196\8867-4\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | PV Annul Diam | Dist | meas | m_fet_pvAnn | r_Distance | unilateral | 125196\G-03A2\M-02550\T-3500E\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-3500E, Cardiac valve annulus | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | MR VTI | Vmean | meas | m_fet_mrVTI | r_CardTAPV | unilateral | 125196\11692-1\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | MV VTI | Vmean | meas | m_fet_mvVTI | r_CardTAPV | unilateral | 125196\11692-1\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | MV EPSS | Dist | meas | m_fet_MV_EPSS | r_Distance | unilateral | 125196\18036-4\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 18036-4, Mitral Valve EPSS, E wave |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | MV Acc Time | Time | meas | m_fet_mvAccTime | r_DopTime | unilateral | 125196\20168-1\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | MV Dec Slope | Slope | meas | m_fet_mvDecSlope | r_DopSlope | unilateral | 125196\20216-8\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20216-8, Deceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | MV Dec Time | Time | meas | m_fet_mvDecTime | r_DopTime | unilateral | 125196\20217-6\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | MR Vmax | Max PG | meas | m_fet_mrVmax | r_MaxPG | unilateral | 125196\20247-3\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | MV Vmax | Max PG | meas | m_fet_mvVmax | r_MaxPG | unilateral | 125196\20247-3\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | MR VTI | Mean PG | meas | m_fet_mrVTI | r_CardPGmean | unilateral | 125196\20256-4\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | MV VTI | Mean PG | meas | m_fet_mvVTI | r_CardPGmean | unilateral | 125196\20256-4\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | MV P½t | P½t | meas | m_fet_mvPHalfT | r_PHalftime | unilateral | 125196\20280-4\T-35300 | SRT, T-35300, Mitral Valve | LN, 20280-4, Pressure Half-Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | MR Vmax | Vmax | meas | m_fet_mrVmax | r_Vmax | unilateral | 125196\20351-3\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | MV Vmax | Vmax | meas | m_fet_mvVmax | r_Vmax | unilateral | 125196\20351-3\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | MR VTI | VTI | meas | m_fet_mrVTI | r_VTI | unilateral | 125196\20354-7\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | MV VTI | VTI | meas | m_fet_mvVTI | r_VTI | unilateral | 125196\20354-7\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | MV Peak E | Vel | meas | m_fet_mvPeakE | r_Vel | unilateral | 125196\59080-2\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 59080-2, E-Wave Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | MV Peak A | Vel | meas | m_fet_mvPeakA | r_Vel | unilateral | 125196\59081-0\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 59081-0, A-Wave Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | MV D-E Exc | Dist | meas | m_fet_MV_DE_Exc | r_Distance | unilateral | 125196\59091-9\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 59091-9, D-E Excursion |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | MV E-E Sep | Dist | meas | m_fet_MV_EE_Separ | r_Distance | unilateral | 125196\59098-4\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 59098-4, Mitral Valve E-septal Separation |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | MV A-C Interval | Time | meas | m_fet_MV_AC_Int | r_MMTime | unilateral | 125196\59103-2\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 59103-2, A-C Interval |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | MV D-E Slope | Slope | meas | m_fet_MV_DE_Slope | r_MMSlope | unilateral | 125196\59127-1\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 59127-1, D-E Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | MV E-F Slope | Slope | meas | m_fet_MV_EF_Slope | r_MMSlope | unilateral | 125196\59128-9\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 59128-9, E-F Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | HR-MV | HR | meas | m_fet_HRMvDop | r_HeartRate | unilateral | 125196\8867-4\T-35300 | SRT, T-35300, Mitral Valve | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | MV Annul Diam | Dist | meas | m_fet_mvAnn | r_Distance | unilateral | 125196\G-03A2\M-02550\T-3500E\T-35300 | SRT, T-35300, Mitral Valve | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-3500E, Cardiac valve annulus | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | MV Annul Circ | Circ | meas | m_fet_mvAnnCirc | r_Circumference | unilateral | 125196\G-03A2\M-02560\T-3500E\T-35300 | SRT, T-35300, Mitral Valve | SRT, M-02560, Circumference |  |  | UCUM, mm, mm | SRT, T-3500E, Cardiac valve annulus | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | AI End Dias Vel | PG | meas | m_fet_aiEDVel | r_PG | unilateral | 125196\C12122-10\F-32011\G-0367\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12122-10, Pressure Gradient |  |  | UCUM, mmHg, mmHg |  |  |  | SRT, F-32011, End Diastole | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | AI End Dias Vel | Vel | meas | m_fet_aiEDVel | r_Vel | unilateral | 125196\11653-3\F-32011\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, F-32011, End Diastole | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | AI VTI | Vmean | meas | m_fet_aiVTI | r_CardTAPV | unilateral | 125196\11692-1\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | AV VTI | Vmean | meas | m_fet_avVTI | r_CardTAPV | unilateral | 125196\11692-1\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | AV Cusp Sep | Dist | meas | m_fet_AV_Cusp_Sep | r_Distance | unilateral | 125196\17996-0\G-0394\T-35400 | SRT, T-35400, Aortic Valve | LN, 17996-0, Aortic Valve Cusp Separation |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | AI Acc Time | Time | meas | m_fet_aiAccTime | r_DopTime | unilateral | 125196\20168-1\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | AV Acc Time | Time | meas | m_fet_avAccTime | r_DopTime | unilateral | 125196\20168-1\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | AI Vmax | Max PG | meas | m_fet_aiVmax | r_MaxPG | unilateral | 125196\20247-3\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | AV Vmax | Max PG | meas | m_fet_avVmax | r_MaxPG | unilateral | 125196\20247-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | AI VTI | Mean PG | meas | m_fet_aiVTI | r_CardPGmean | unilateral | 125196\20256-4\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | AV VTI | Mean PG | meas | m_fet_avVTI | r_CardPGmean | unilateral | 125196\20256-4\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | AI P½t | P½t | meas | m_fet_aiPHalfT | r_PHalftime | unilateral | 125196\20280-4\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20280-4, Pressure Half-Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | AI Vmax | Vmax | meas | m_fet_aiVmax | r_Vmax | unilateral | 125196\20351-3\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | AV Vmax | Vmax | meas | m_fet_avVmax | r_Vmax | unilateral | 125196\20351-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | AI VTI | VTI | meas | m_fet_aiVTI | r_VTI | unilateral | 125196\20354-7\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | AV VTI | VTI | meas | m_fet_avVTI | r_VTI | unilateral | 125196\20354-7\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | HR-AV | HR | meas | m_fet_HRAvDop | r_HeartRate | unilateral | 125196\8867-4\T-35400 | SRT, T-35400, Aortic Valve | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | HR-AV (MM) | HR | meas | m_fet_HRAvMM | r_HeartRate | unilateral | 125196\8867-4\G-0394\T-35400 | SRT, T-35400, Aortic Valve | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Ao Annul Diam | Dist | meas | m_fet_aoAnn | r_Distance | unilateral | 125196\G-03A2\M-02550\T-35400\T-35410 | SRT, T-35400, Aortic Valve | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-35410, Aortic Valve Ring | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Ao Annul Circ | Circ | meas | m_fet_aoAnnCirc | r_Circumference | unilateral | 125196\G-03A2\M-02560\T-35400\T-35410 | SRT, T-35400, Aortic Valve | SRT, M-02560, Circumference |  |  | UCUM, mm, mm | SRT, T-35410, Aortic Valve Ring | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Thoracic Ao | ? | meas | m_fet_thrAo | r_DopAngle | unilateral | 125196\C12122-01\T-42070 | SRT, T-42070, Thoracic Aorta | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | Thoracic Ao | EDV | meas | m_fet_thrAo | r_EDV | unilateral | 125196\11653-3\T-42070 | SRT, T-42070, Thoracic Aorta | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Thoracic Ao | MDV | meas | m_fet_thrAo | r_MDV | unilateral | 125196\11665-7\T-42070 | SRT, T-42070, Thoracic Aorta | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Thoracic Ao | TAPV | meas | m_fet_thrAo | r_TAPV | unilateral | 125196\11692-1\T-42070 | SRT, T-42070, Thoracic Aorta | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Thoracic Ao | PSV | meas | m_fet_thrAo | r_PSV | unilateral | 125196\11726-7\T-42070 | SRT, T-42070, Thoracic Aorta | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Thoracic Ao | PI | meas | m_fet_thrAo | r_PI | unilateral | 125196\12008-9\T-42070 | SRT, T-42070, Thoracic Aorta | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Thoracic Ao | RI | meas | m_fet_thrAo | r_RI | unilateral | 125196\12023-8\T-42070 | SRT, T-42070, Thoracic Aorta | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Thoracic Ao | S/D | meas | m_fet_thrAo | r_SD | unilateral | 125196\12144-2\T-42070 | SRT, T-42070, Thoracic Aorta | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Thoracic Ao | AI | meas | m_fet_thrAo | r_AccelIndex | unilateral | 125196\20167-3\T-42070 | SRT, T-42070, Thoracic Aorta | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | Thoracic Ao | AT | meas | m_fet_thrAo | r_AccelTime | unilateral | 125196\20168-1\T-42070 | SRT, T-42070, Thoracic Aorta | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Thoracic Ao | DT | meas | m_fet_thrAo | r_DecelTime | unilateral | 125196\20217-6\T-42070 | SRT, T-42070, Thoracic Aorta | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Desc Ao | Max PG | meas | m_fet_descAoDop | r_MaxPG | unilateral | 125196\20247-3\T-42070 | SRT, T-42070, Thoracic Aorta | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Fetal Cardiac Ultrasound Report | Fetal Echo | Desc Ao | Vmax | meas | m_fet_descAoDop | r_Vmax | unilateral | 125196\20351-3\T-42070 | SRT, T-42070, Thoracic Aorta | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Thoracic Ao | TAMV | meas | m_fet_thrAo | r_TAMV | unilateral | 125196\20352-1\T-42070 | SRT, T-42070, Thoracic Aorta | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Thoracic Ao | VTI | meas | m_fet_thrAo | r_VTI | unilateral | 125196\20354-7\T-42070 | SRT, T-42070, Thoracic Aorta | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Thoracic Ao | HR | meas | m_fet_thrAo | r_HeartRate | unilateral | 125196\8867-4\T-42070 | SRT, T-42070, Thoracic Aorta | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | Thoracic Ao | Dist | meas | m_fet_thrAo | r_Distance | unilateral | 125196\G-0364\T-42070 | SRT, T-42070, Thoracic Aorta | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Desc Ao Diam (MM) | Dist | meas | m_fet_descAoMM | r_Distance | unilateral | 125196\G-0394\M-02550\T-42070 | SRT, T-42070, Thoracic Aorta | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Desc Ao Diam (2D) | Dist | meas | m_fet_descAo2D | r_Distance | unilateral | 125196\G-03A2\M-02550\T-42070 | SRT, T-42070, Thoracic Aorta | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Asc Ao | Max PG | meas | m_fet_ascAoDop | r_MaxPG | unilateral | 125196\20247-3\T-42100 | SRT, T-42100, Ascending aorta | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Fetal Cardiac Ultrasound Report | Fetal Echo | Asc Ao | Vmax | meas | m_fet_ascAoDop | r_Vmax | unilateral | 125196\20351-3\T-42100 | SRT, T-42100, Ascending aorta | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Asc Ao Diam (MM) | Dist | meas | m_fet_ascAoMM | r_Distance | unilateral | 125196\G-0394\M-02550\T-42100 | SRT, T-42100, Ascending aorta | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Asc Ao Diam (2D) | Dist | meas | m_fet_ascAo2D | r_Distance | unilateral | 125196\G-03A2\M-02550\T-42100 | SRT, T-42100, Ascending aorta | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | AoR | Max PG | meas | m_fet_aoRDop | r_MaxPG | unilateral | 125196\20247-3\T-42110 | SRT, T-42110, Root of Aorta | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Fetal Cardiac Ultrasound Report | Fetal Echo | AoR | Vmax | meas | m_fet_aoRDop | r_Vmax | unilateral | 125196\20351-3\T-42110 | SRT, T-42110, Root of Aorta | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | AoR Diam (MM) | Dist | meas | m_fet_aoRMM | r_Distance | unilateral | 125196\G-0394\M-02550\T-42110 | SRT, T-42110, Root of Aorta | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | AoR Diam (2D) | Dist | meas | m_fet_aoR2D | r_Distance | unilateral | 125196\G-03A2\M-02550\T-42110 | SRT, T-42110, Root of Aorta | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Ao Arch | Max PG | meas | m_fet_aoArchDop | r_MaxPG | unilateral | 125196\20247-3\T-42300 | SRT, T-42300, Aortic arch | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Fetal Cardiac Ultrasound Report | Fetal Echo | Ao Arch | Vmax | meas | m_fet_aoArchDop | r_Vmax | unilateral | 125196\20351-3\T-42300 | SRT, T-42300, Aortic arch | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Ao Arch Diam (MM) | Dist | meas | m_fet_aoArchMM | r_Distance | unilateral | 125196\G-0394\M-02550\T-42300 | SRT, T-42300, Aortic arch | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Ao Arch Diam (2D) | Dist | meas | m_fet_aoArch2D | r_Distance | unilateral | 125196\G-03A2\M-02550\T-42300 | SRT, T-42300, Aortic arch | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Ao Isthmus | Max PG | meas | m_fet_isthmusDop | r_MaxPG | unilateral | 125196\20247-3\T-42310 | SRT, T-42310, Aortic Isthmus | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Fetal Cardiac Ultrasound Report | Fetal Echo | Ao Isthmus | Vmax | meas | m_fet_isthmusDop | r_Vmax | unilateral | 125196\20351-3\T-42310 | SRT, T-42310, Aortic Isthmus | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Ao Isthmus Diam (MM) | Dist | meas | m_fet_isthmusMM | r_Distance | unilateral | 125196\G-0394\M-02550\T-42310 | SRT, T-42310, Aortic Isthmus | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Ao Isthmus Diam (2D) | Dist | meas | m_fet_isthmus2D | r_Distance | unilateral | 125196\G-03A2\M-02550\T-42310 | SRT, T-42310, Aortic Isthmus | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | RPA | Max PG | meas | m_fet_rpaDop | r_MaxPG | unilateral | 125196\20247-3\T-44000\T-44200 | SRT, T-44000, Pulmonary artery | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-44200, Right Pulmonary Artery
Fetal Cardiac Ultrasound Report | Fetal Echo | LPA | Max PG | meas | m_fet_lpaDop | r_MaxPG | unilateral | 125196\20247-3\T-44000\T-44400 | SRT, T-44000, Pulmonary artery | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-44400, Left Pulmonary Artery
Fetal Cardiac Ultrasound Report | Fetal Echo | RPA | Vmax | meas | m_fet_rpaDop | r_Vmax | unilateral | 125196\20351-3\T-44000\T-44200 | SRT, T-44000, Pulmonary artery | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, T-44200, Right Pulmonary Artery
Fetal Cardiac Ultrasound Report | Fetal Echo | LPA | Vmax | meas | m_fet_lpaDop | r_Vmax | unilateral | 125196\20351-3\T-44000\T-44400 | SRT, T-44000, Pulmonary artery | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, T-44400, Left Pulmonary Artery
Fetal Cardiac Ultrasound Report | Fetal Echo | RPA Diam (MM) | Dist | meas | m_fet_rpaMM | r_Distance | unilateral | 125196\G-0394\M-02550\T-44000\T-44200 | SRT, T-44000, Pulmonary artery | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-44200, Right Pulmonary Artery | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LPA Diam (MM) | Dist | meas | m_fet_lpaMM | r_Distance | unilateral | 125196\G-0394\M-02550\T-44000\T-44400 | SRT, T-44000, Pulmonary artery | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-44400, Left Pulmonary Artery | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | RPA Diam (2D) | Dist | meas | m_fet_rpa2D | r_Distance | unilateral | 125196\G-03A2\M-02550\T-44000\T-44200 | SRT, T-44000, Pulmonary artery | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-44200, Right Pulmonary Artery | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LPA Diam (2D) | Dist | meas | m_fet_lpa2D | r_Distance | unilateral | 125196\G-03A2\M-02550\T-44000\T-44400 | SRT, T-44000, Pulmonary artery | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-44400, Left Pulmonary Artery | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | MPA | Max PG | meas | m_fet_mpaDop | r_MaxPG | unilateral | 125196\20247-3\T-44100 | SRT, T-44100, Pulmonary Trunk | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Fetal Cardiac Ultrasound Report | Fetal Echo | MPA | Vmax | meas | m_fet_mpaDop | r_Vmax | unilateral | 125196\20351-3\T-44100 | SRT, T-44100, Pulmonary Trunk | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | MPA Diam (MM) | Dist | meas | m_fet_mpaMM | r_Distance | unilateral | 125196\G-0394\M-02550\T-44100 | SRT, T-44100, Pulmonary Trunk | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | MPA Diam (2D) | Dist | meas | m_fet_mpa2D | r_Distance | unilateral | 125196\G-03A2\M-02550\T-44100 | SRT, T-44100, Pulmonary Trunk | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic A | ? | meas | m_fet_hepArt | r_DopAngle | unilateral | 125196\C12122-01\T-46420 | SRT, T-46420, Hepatic Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic A | EDV | meas | m_fet_hepArt | r_EDV | unilateral | 125196\11653-3\T-46420 | SRT, T-46420, Hepatic Artery | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic A | MDV | meas | m_fet_hepArt | r_MDV | unilateral | 125196\11665-7\T-46420 | SRT, T-46420, Hepatic Artery | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic A | TAPV | meas | m_fet_hepArt | r_TAPV | unilateral | 125196\11692-1\T-46420 | SRT, T-46420, Hepatic Artery | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic A | PSV | meas | m_fet_hepArt | r_PSV | unilateral | 125196\11726-7\T-46420 | SRT, T-46420, Hepatic Artery | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic A | PI | meas | m_fet_hepArt | r_PI | unilateral | 125196\12008-9\T-46420 | SRT, T-46420, Hepatic Artery | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic A | RI | meas | m_fet_hepArt | r_RI | unilateral | 125196\12023-8\T-46420 | SRT, T-46420, Hepatic Artery | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic A | S/D | meas | m_fet_hepArt | r_SD | unilateral | 125196\12144-2\T-46420 | SRT, T-46420, Hepatic Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic A | AI | meas | m_fet_hepArt | r_AccelIndex | unilateral | 125196\20167-3\T-46420 | SRT, T-46420, Hepatic Artery | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic A | AT | meas | m_fet_hepArt | r_AccelTime | unilateral | 125196\20168-1\T-46420 | SRT, T-46420, Hepatic Artery | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic A | DT | meas | m_fet_hepArt | r_DecelTime | unilateral | 125196\20217-6\T-46420 | SRT, T-46420, Hepatic Artery | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic A | TAMV | meas | m_fet_hepArt | r_TAMV | unilateral | 125196\20352-1\T-46420 | SRT, T-46420, Hepatic Artery | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic A | VTI | meas | m_fet_hepArt | r_VTI | unilateral | 125196\20354-7\T-46420 | SRT, T-46420, Hepatic Artery | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic A | HR | meas | m_fet_hepArt | r_HeartRate | unilateral | 125196\8867-4\T-46420 | SRT, T-46420, Hepatic Artery | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic A | Dist | meas | m_fet_hepArt | r_Distance | unilateral | 125196\G-0364\T-46420 | SRT, T-46420, Hepatic Artery | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Splenic A | ? | meas | m_fet_splenArt | r_DopAngle | unilateral | 125196\C12122-01\T-46460 | SRT, T-46460, Splenic Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | Splenic A | EDV | meas | m_fet_splenArt | r_EDV | unilateral | 125196\11653-3\T-46460 | SRT, T-46460, Splenic Artery | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Splenic A | MDV | meas | m_fet_splenArt | r_MDV | unilateral | 125196\11665-7\T-46460 | SRT, T-46460, Splenic Artery | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Splenic A | TAPV | meas | m_fet_splenArt | r_TAPV | unilateral | 125196\11692-1\T-46460 | SRT, T-46460, Splenic Artery | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Splenic A | PSV | meas | m_fet_splenArt | r_PSV | unilateral | 125196\11726-7\T-46460 | SRT, T-46460, Splenic Artery | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Splenic A | PI | meas | m_fet_splenArt | r_PI | unilateral | 125196\12008-9\T-46460 | SRT, T-46460, Splenic Artery | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Splenic A | RI | meas | m_fet_splenArt | r_RI | unilateral | 125196\12023-8\T-46460 | SRT, T-46460, Splenic Artery | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Splenic A | S/D | meas | m_fet_splenArt | r_SD | unilateral | 125196\12144-2\T-46460 | SRT, T-46460, Splenic Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Splenic A | AI | meas | m_fet_splenArt | r_AccelIndex | unilateral | 125196\20167-3\T-46460 | SRT, T-46460, Splenic Artery | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | Splenic A | AT | meas | m_fet_splenArt | r_AccelTime | unilateral | 125196\20168-1\T-46460 | SRT, T-46460, Splenic Artery | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Splenic A | DT | meas | m_fet_splenArt | r_DecelTime | unilateral | 125196\20217-6\T-46460 | SRT, T-46460, Splenic Artery | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Splenic A | TAMV | meas | m_fet_splenArt | r_TAMV | unilateral | 125196\20352-1\T-46460 | SRT, T-46460, Splenic Artery | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Splenic A | VTI | meas | m_fet_splenArt | r_VTI | unilateral | 125196\20354-7\T-46460 | SRT, T-46460, Splenic Artery | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Splenic A | HR | meas | m_fet_splenArt | r_HeartRate | unilateral | 125196\8867-4\T-46460 | SRT, T-46460, Splenic Artery | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | Splenic A | Dist | meas | m_fet_splenArt | r_Distance | unilateral | 125196\G-0364\T-46460 | SRT, T-46460, Splenic Artery | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Pulm V | Vel | meas | m_fet_pulmVeins | r_Vel | unilateral | 125196\20351-3\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | IVC | PG | meas | m_fet_ivcDop | r_PG | unilateral | 125196\C12122-10\T-48710 | SRT, T-48710, Inferior vena cava | 99PMSBLUS, C12122-10, Pressure Gradient |  |  | UCUM, mmHg, mmHg
Fetal Cardiac Ultrasound Report | Fetal Echo | IVC | Vel | meas | m_fet_ivcDop | r_Vel | unilateral | 125196\20351-3\T-48710 | SRT, T-48710, Inferior vena cava | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | IVC Rev VTI | VTI | meas | m_fet_ivcRevVTI | r_VTI | unilateral | 125196\20354-7\G-0367\T-48710 | SRT, T-48710, Inferior vena cava | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, G-0367, Regurgitant Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | IVC Forw VTI | VTI | meas | m_fet_ivcForwVTI | r_VTI | unilateral | 125196\20354-7\R-42047\T-48710 | SRT, T-48710, Inferior vena cava | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Fetal Cardiac Ultrasound Report | Fetal Echo | IVC Diam (2D) | Dist | meas | m_fet_ivc2D | r_Distance | unilateral | 125196\G-03A2\M-02550\T-48710 | SRT, T-48710, Inferior vena cava | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic V | ? | meas | m_fet_hepVein | r_DopAngle | unilateral | 125196\C12122-01\T-48720 | SRT, T-48720, Hepatic Vein | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic V | EDV | meas | m_fet_hepVein | r_EDV | unilateral | 125196\11653-3\T-48720 | SRT, T-48720, Hepatic Vein | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic V | MDV | meas | m_fet_hepVein | r_MDV | unilateral | 125196\11665-7\T-48720 | SRT, T-48720, Hepatic Vein | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic V | TAPV | meas | m_fet_hepVein | r_TAPV | unilateral | 125196\11692-1\T-48720 | SRT, T-48720, Hepatic Vein | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic V | PSV | meas | m_fet_hepVein | r_PSV | unilateral | 125196\11726-7\T-48720 | SRT, T-48720, Hepatic Vein | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic V | PI | meas | m_fet_hepVein | r_PI | unilateral | 125196\12008-9\T-48720 | SRT, T-48720, Hepatic Vein | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic V | RI | meas | m_fet_hepVein | r_RI | unilateral | 125196\12023-8\T-48720 | SRT, T-48720, Hepatic Vein | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic V | S/D | meas | m_fet_hepVein | r_SD | unilateral | 125196\12144-2\T-48720 | SRT, T-48720, Hepatic Vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic V | AI | meas | m_fet_hepVein | r_AccelIndex | unilateral | 125196\20167-3\T-48720 | SRT, T-48720, Hepatic Vein | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic V | AT | meas | m_fet_hepVein | r_AccelTime | unilateral | 125196\20168-1\T-48720 | SRT, T-48720, Hepatic Vein | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic V | DT | meas | m_fet_hepVein | r_DecelTime | unilateral | 125196\20217-6\T-48720 | SRT, T-48720, Hepatic Vein | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic V | TAMV | meas | m_fet_hepVein | r_TAMV | unilateral | 125196\20352-1\T-48720 | SRT, T-48720, Hepatic Vein | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic V | VTI | meas | m_fet_hepVein | r_VTI | unilateral | 125196\20354-7\T-48720 | SRT, T-48720, Hepatic Vein | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic V | HR | meas | m_fet_hepVein | r_HeartRate | unilateral | 125196\8867-4\T-48720 | SRT, T-48720, Hepatic Vein | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | Hepatic V | Dist | meas | m_fet_hepVein | r_Distance | unilateral | 125196\G-0364\T-48720 | SRT, T-48720, Hepatic Vein | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Placenta | ? | meas | m_fet_placenta | r_DopAngle | unilateral | 125196\C12122-01\T-F1412 | SRT, T-F1412, Vitelline Artery of Placenta | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | Placenta | EDV | meas | m_fet_placenta | r_EDV | unilateral | 125196\11653-3\T-F1412 | SRT, T-F1412, Vitelline Artery of Placenta | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Placenta | MDV | meas | m_fet_placenta | r_MDV | unilateral | 125196\11665-7\T-F1412 | SRT, T-F1412, Vitelline Artery of Placenta | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Placenta | TAPV | meas | m_fet_placenta | r_TAPV | unilateral | 125196\11692-1\T-F1412 | SRT, T-F1412, Vitelline Artery of Placenta | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Placenta | PSV | meas | m_fet_placenta | r_PSV | unilateral | 125196\11726-7\T-F1412 | SRT, T-F1412, Vitelline Artery of Placenta | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Placenta | PI | meas | m_fet_placenta | r_PI | unilateral | 125196\12008-9\T-F1412 | SRT, T-F1412, Vitelline Artery of Placenta | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Placenta | RI | meas | m_fet_placenta | r_RI | unilateral | 125196\12023-8\T-F1412 | SRT, T-F1412, Vitelline Artery of Placenta | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Placenta | S/D | meas | m_fet_placenta | r_SD | unilateral | 125196\12144-2\T-F1412 | SRT, T-F1412, Vitelline Artery of Placenta | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Placenta | AI | meas | m_fet_placenta | r_AccelIndex | unilateral | 125196\20167-3\T-F1412 | SRT, T-F1412, Vitelline Artery of Placenta | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | Placenta | AT | meas | m_fet_placenta | r_AccelTime | unilateral | 125196\20168-1\T-F1412 | SRT, T-F1412, Vitelline Artery of Placenta | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Placenta | DT | meas | m_fet_placenta | r_DecelTime | unilateral | 125196\20217-6\T-F1412 | SRT, T-F1412, Vitelline Artery of Placenta | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Placenta | TAMV | meas | m_fet_placenta | r_TAMV | unilateral | 125196\20352-1\T-F1412 | SRT, T-F1412, Vitelline Artery of Placenta | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Placenta | VTI | meas | m_fet_placenta | r_VTI | unilateral | 125196\20354-7\T-F1412 | SRT, T-F1412, Vitelline Artery of Placenta | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Placenta | HR | meas | m_fet_placenta | r_HeartRate | unilateral | 125196\8867-4\T-F1412 | SRT, T-F1412, Vitelline Artery of Placenta | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | Placenta | Dist | meas | m_fet_placenta | r_Distance | unilateral | 125196\G-0364\T-F1412 | SRT, T-F1412, Vitelline Artery of Placenta | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical A | ? | meas | m_fet_umbArt | r_DopAngle | unilateral | 125196\C12122-01\T-F1810 | SRT, T-F1810, Umbilical Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical A | EDV | meas | m_fet_umbArt | r_EDV | unilateral | 125196\11653-3\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical A | MDV | meas | m_fet_umbArt | r_MDV | unilateral | 125196\11665-7\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical A | TAPV | meas | m_fet_umbArt | r_TAPV | unilateral | 125196\11692-1\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical A | PSV | meas | m_fet_umbArt | r_PSV | unilateral | 125196\11726-7\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical A | PI | meas | m_fet_umbArt | r_PI | unilateral | 125196\12008-9\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical A | RI | meas | m_fet_umbArt | r_RI | unilateral | 125196\12023-8\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical A | S/D | meas | m_fet_umbArt | r_SD | unilateral | 125196\12144-2\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical A | AI | meas | m_fet_umbArt | r_AccelIndex | unilateral | 125196\20167-3\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical A | AT | meas | m_fet_umbArt | r_AccelTime | unilateral | 125196\20168-1\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical A | DT | meas | m_fet_umbArt | r_DecelTime | unilateral | 125196\20217-6\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical A | TAMV | meas | m_fet_umbArt | r_TAMV | unilateral | 125196\20352-1\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical A | VTI | meas | m_fet_umbArt | r_VTI | unilateral | 125196\20354-7\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical A | HR | meas | m_fet_umbArt | r_HeartRate | unilateral | 125196\8867-4\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical A | Dist | meas | m_fet_umbArt | r_Distance | unilateral | 125196\G-0364\T-F1810 | SRT, T-F1810, Umbilical Artery | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical V | ? | meas | m_fet_umbVein | r_DopAngle | unilateral | 125196\C12122-01\T-F1820 | SRT, T-F1820, Umbilical Vein | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | Umbilical V | Vel | meas | m_fet_umbVein | r_Vel | unilateral | 125196\20351-3\T-F1820 | SRT, T-F1820, Umbilical Vein | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Ven | ? | meas | m_fet_ductVen | r_DopAngle | unilateral | 125196\C12122-01\T-F6806 | SRT, T-F6806, Ductus Venosus | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Ven | EDV | meas | m_fet_ductVen | r_EDV | unilateral | 125196\11653-3\T-F6806 | SRT, T-F6806, Ductus Venosus | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Ven | MDV | meas | m_fet_ductVen | r_MDV | unilateral | 125196\11665-7\T-F6806 | SRT, T-F6806, Ductus Venosus | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Ven | TAPV | meas | m_fet_ductVen | r_TAPV | unilateral | 125196\11692-1\T-F6806 | SRT, T-F6806, Ductus Venosus | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Ven | PSV | meas | m_fet_ductVen | r_PSV | unilateral | 125196\11726-7\T-F6806 | SRT, T-F6806, Ductus Venosus | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Ven | PI | meas | m_fet_ductVen | r_PI | unilateral | 125196\12008-9\T-F6806 | SRT, T-F6806, Ductus Venosus | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Ven | RI | meas | m_fet_ductVen | r_RI | unilateral | 125196\12023-8\T-F6806 | SRT, T-F6806, Ductus Venosus | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Ven | S/D | meas | m_fet_ductVen | r_SD | unilateral | 125196\12144-2\T-F6806 | SRT, T-F6806, Ductus Venosus | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Ven | AI | meas | m_fet_ductVen | r_AccelIndex | unilateral | 125196\20167-3\T-F6806 | SRT, T-F6806, Ductus Venosus | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Ven | AT | meas | m_fet_ductVen | r_AccelTime | unilateral | 125196\20168-1\T-F6806 | SRT, T-F6806, Ductus Venosus | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Ven | DT | meas | m_fet_ductVen | r_DecelTime | unilateral | 125196\20217-6\T-F6806 | SRT, T-F6806, Ductus Venosus | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Ven | TAMV | meas | m_fet_ductVen | r_TAMV | unilateral | 125196\20352-1\T-F6806 | SRT, T-F6806, Ductus Venosus | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Ven | VTI | meas | m_fet_ductVen | r_VTI | unilateral | 125196\20354-7\T-F6806 | SRT, T-F6806, Ductus Venosus | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Ven | HR | meas | m_fet_ductVen | r_HeartRate | unilateral | 125196\8867-4\T-F6806 | SRT, T-F6806, Ductus Venosus | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Ven | Dist | meas | m_fet_ductVen | r_Distance | unilateral | 125196\G-0364\T-F6806 | SRT, T-F6806, Ductus Venosus | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art | ? | meas | m_fet_ductArtDop | r_DopAngle | unilateral | 125196\C12122-01\T-F6845 | SRT, T-F6845, Ductus arteriosus | 99PMSBLUS, C12122-01, Doppler Correction Angle |  |  | UCUM, deg, deg
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art Diam (MM) | Dist | meas | m_fet_ductArtMM | r_Distance | unilateral | 125196\G-0394\P5000-01-24\T-F6845 | SRT, T-F6845, Ductus arteriosus | 99PMSBLUS, P5000-01-24, Ductus Arteriosus Diameter |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art Diam (2D) | Dist | meas | m_fet_ductArt2D | r_Distance | unilateral | 125196\G-03A2\P5000-01-24\T-F6845 | SRT, T-F6845, Ductus arteriosus | 99PMSBLUS, P5000-01-24, Ductus Arteriosus Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art | EDV | meas | m_fet_ductArtDop | r_EDV | unilateral | 125196\11653-3\T-F6845 | SRT, T-F6845, Ductus arteriosus | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art | MDV | meas | m_fet_ductArtDop | r_MDV | unilateral | 125196\11665-7\T-F6845 | SRT, T-F6845, Ductus arteriosus | LN, 11665-7, Minimum Diastolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art | TAPV | meas | m_fet_ductArtDop | r_TAPV | unilateral | 125196\11692-1\T-F6845 | SRT, T-F6845, Ductus arteriosus | LN, 11692-1, Time averaged peak velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art | PSV | meas | m_fet_ductArtDop | r_PSV | unilateral | 125196\11726-7\T-F6845 | SRT, T-F6845, Ductus arteriosus | LN, 11726-7, Peak Systolic Velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art | PI | meas | m_fet_ductArtDop | r_PI | unilateral | 125196\12008-9\T-F6845 | SRT, T-F6845, Ductus arteriosus | LN, 12008-9, Pulsatility Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art | RI | meas | m_fet_ductArtDop | r_RI | unilateral | 125196\12023-8\T-F6845 | SRT, T-F6845, Ductus arteriosus | LN, 12023-8, Resistivity Index |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art | S/D | meas | m_fet_ductArtDop | r_SD | unilateral | 125196\12144-2\T-F6845 | SRT, T-F6845, Ductus arteriosus | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art | AI | meas | m_fet_ductArtDop | r_AccelIndex | unilateral | 125196\20167-3\T-F6845 | SRT, T-F6845, Ductus arteriosus | LN, 20167-3, Acceleration Index |  |  | UCUM, mm/s2, mm/s2
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art | AT | meas | m_fet_ductArtDop | r_AccelTime | unilateral | 125196\20168-1\T-F6845 | SRT, T-F6845, Ductus arteriosus | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art | DT | meas | m_fet_ductArtDop | r_DecelTime | unilateral | 125196\20217-6\T-F6845 | SRT, T-F6845, Ductus arteriosus | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art | TAMV | meas | m_fet_ductArtDop | r_TAMV | unilateral | 125196\20352-1\T-F6845 | SRT, T-F6845, Ductus arteriosus | LN, 20352-1, Time averaged mean velocity |  |  | UCUM, mm/s, mm/s
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art | VTI | meas | m_fet_ductArtDop | r_VTI | unilateral | 125196\20354-7\T-F6845 | SRT, T-F6845, Ductus arteriosus | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art | HR | meas | m_fet_ductArtDop | r_HeartRate | unilateral | 125196\8867-4\T-F6845 | SRT, T-F6845, Ductus arteriosus | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | Duct Art | Dist | meas | m_fet_ductArtDop | r_Distance | unilateral | 125196\G-0364\T-F6845 | SRT, T-F6845, Ductus arteriosus | SRT, G-0364, Vessel lumen diameter |  |  | UCUM, mm, mm
Fetal Cardiac Ultrasound Report | Fetal Echo | LA/RA | none | calc | c_fet_laRa2D | none | unilateral | 125196\99999\P5000-01-10 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-10, Ratio of LA Dimension to RA Dimension |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Fetal HR 2 Bts(DOP) | none | calc | c_DOP_2BeatPKToPK | none | unilateral | 125196\11948-7\99999 | 99PMSBLUS, 99999, Fetal Heart | LN, 11948-7, Fetal Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Fetal Cardiac Ultrasound Report | Fetal Echo | Fetal HR 2 Bts(MM) | none | calc | c_MM_2BeatPKToPK | none | unilateral | 125196\11948-7\99999\G-0394 | 99PMSBLUS, 99999, Fetal Heart | LN, 11948-7, Fetal Heart Rate |  |  | UCUM, {H.B.}/min, bpm |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Qp/Qs | none | calc | c_fet_qpqs | none | unilateral | 125196\29462-9\P5-30031 | SRT, P5-30031, Cardiac Shunt Study | LN, 29462-9, Pulmonary-to-Systemic Shunt Flow Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | LA/Ao (MM) | none | calc | c_fet_laAoMM | none | unilateral | 125196\17985-3\G-0394\T-32300 | SRT, T-32300, Left Atrium | LN, 17985-3, Left Atrium to Aortic Root Ratio |  |  | UCUM, 1, no units |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LA/Ao (2D) | none | calc | c_fet_laAo2D | none | unilateral | 125196\17985-3\G-03A2\T-32300 | SRT, T-32300, Left Atrium | LN, 17985-3, Left Atrium to Aortic Root Ratio |  |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | IVS% (MM) | none | calc | c_fet_ivsMM | none | unilateral | 125196\59092-7\G-0394\T-32410 | SRT, T-32410, Interventricular septum | LN, 59092-7, % Thickening |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | IVS% (2D) | none | calc | c_fet_ivs | none | unilateral | 125196\59092-7\G-03A2\T-32410 | SRT, T-32410, Interventricular septum | LN, 59092-7, % Thickening |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | RV Eject Force | none | calc | c_fet_rvotVef | none | unilateral | 125196\C12252-01\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12252-01, Ejection Force |  |  | 99PMSBLUS, kdynes, kdynes
Fetal Cardiac Ultrasound Report | Fetal Echo | RVPEPc | none | calc | c_fet_RVPEPc | none | unilateral | 125196\59087-7\T-32500 | SRT, T-32500, Right Ventricle | LN, 59087-7, Heart Rate-Corrected Pre-Ejection Period |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | RV PEP/ET | none | calc | c_fet_RV_PEP_ET | none | unilateral | 125196\59088-5\T-32500 | SRT, T-32500, Right Ventricle | LN, 59088-5, Pre-Ejection Period/Ejection Time Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | CO (RVOT) | none | calc | c_fet_rvotCo | none | unilateral | 125196\F-32100\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s | SRT, T-32550, Right Ventricle Outflow Tract
Fetal Cardiac Ultrasound Report | Fetal Echo | SV (RVOT) | none | calc | c_fet_rvotSv | none | unilateral | 125196\F-32120\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 | SRT, T-32550, Right Ventricle Outflow Tract
Fetal Cardiac Ultrasound Report | Fetal Echo | RVOT Area | none | calc | c_fet_rvotArea | none | unilateral | 125196\G-038E\G-03A2\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | SRT, T-32550, Right Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LV Eject Force | none | calc | c_fet_lvotVef | none | unilateral | 125196\C12252-01\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12252-01, Ejection Force |  |  | 99PMSBLUS, kdynes, kdynes
Fetal Cardiac Ultrasound Report | Fetal Echo | LV Mass (Cubed) | none | calc | c_fet_LV_Mass | none | unilateral | 125196\125206\18087-7\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18087-7, Left Ventricle Mass |  |  | UCUM, g, grams |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Fetal Cardiac Ultrasound Report | Fetal Echo | IVS/LVPW (MM) | none | calc | c_fet_ivsLvMM | none | unilateral | 125196\18155-2\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18155-2, Interventricular Septum to Posterior Wall Thickness Ratio |  |  | UCUM, 1, no units |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | IVS/LVPW (2D) | none | calc | c_fet_ivsLvpw | none | unilateral | 125196\18155-2\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18155-2, Interventricular Septum to Posterior Wall Thickness Ratio |  |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LV/RV | none | calc | c_fet_lvRv2D | none | unilateral | 125196\59078-6\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 59078-6, Left Ventricle/Right Ventricle Diameter Ratio |  |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LVETc | none | calc | c_fet_LVETc | none | unilateral | 125196\59086-9\T-32600 | SRT, T-32600, Left Ventricle | LN, 59086-9, Heart Rate-Corrected Ejection Time |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | LVPEPc | none | calc | c_fet_LVPEPc | none | unilateral | 125196\59087-7\T-32600 | SRT, T-32600, Left Ventricle | LN, 59087-7, Heart Rate-Corrected Pre-Ejection Period |  |  | UCUM, ms, ms
Fetal Cardiac Ultrasound Report | Fetal Echo | LV PEP/ET | none | calc | c_fet_LV_PEP_ET | none | unilateral | 125196\59088-5\T-32600 | SRT, T-32600, Left Ventricle | LN, 59088-5, Pre-Ejection Period/Ejection Time Ratio |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | LVPW% (MM) | none | calc | c_fet_lvpwMM | none | unilateral | 125196\59092-7\G-0394\R-42175\T-32600 | SRT, T-32600, Left Ventricle | LN, 59092-7, % Thickening |  |  | UCUM, %, Percent | SRT, R-42175, Posterior Wall | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | LVPW% (2D) | none | calc | c_fet_lvpw | none | unilateral | 125196\59092-7\G-03A2\R-42175\T-32600 | SRT, T-32600, Left Ventricle | LN, 59092-7, % Thickening |  |  | UCUM, %, Percent | SRT, R-42175, Posterior Wall | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Mean VCF | none | calc | c_fet_Mean_VCF | none | unilateral | 125196\59117-2\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59117-2, Mean Velocity of Circumferential Fiber Shortening (Mean VcFv) |  |  | 99PMSBLUS, circ/s, circ/s |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | Mean VCFc | none | calc | c_fet_Mean_VCFc | none | unilateral | 125196\59118-0\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59118-0, HR-Corrected Mean Velocity of Circumferential Fiber Shortening |  |  | 99PMSBLUS, circ/s, circ/s |  | SRT, G-0394, M mode
Fetal Cardiac Ultrasound Report | Fetal Echo | E`/A` Medial | none | calc | c_f_MedE_A | none | unilateral | 125196\59129-7\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59129-7, Left Ventricle E to A Tissue Velocity Ratio |  |  | UCUM, 1, no units | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | E`/A` Lateral | none | calc | c_f_LatE_A | none | unilateral | 125196\59129-7\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59129-7, Left Ventricle E to A Tissue Velocity Ratio |  |  | UCUM, 1, no units | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Fetal Cardiac Ultrasound Report | Fetal Echo | FS (MM-Teich) | none | calc | c_fet_fsMM | none | unilateral | 125196\125209\59132-1\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59132-1, Fractional Shortening |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Fetal Cardiac Ultrasound Report | Fetal Echo | FS (2D-Teich) | none | calc | c_fet_fs | none | unilateral | 125196\125209\59132-1\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 59132-1, Fractional Shortening |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Fetal Cardiac Ultrasound Report | Fetal Echo | EF (MM-Teich) | none | calc | c_fet_efMM | none | unilateral | 125196\125209\F-32070\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Fetal Cardiac Ultrasound Report | Fetal Echo | EF (2D-Teich) | none | calc | c_fet_ef | none | unilateral | 125196\125209\F-32070\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Fetal Cardiac Ultrasound Report | Fetal Echo | CO (MM-Teich) | none | calc | c_fet_coMM | none | unilateral | 125196\125209\F-32100\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Fetal Cardiac Ultrasound Report | Fetal Echo | CO (2D-Teich) | none | calc | c_fet_co | none | unilateral | 125196\125209\F-32100\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Fetal Cardiac Ultrasound Report | Fetal Echo | CO (LVOT) | none | calc | c_fet_lvotCo | none | unilateral | 125196\F-32100\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s | SRT, T-32650, Left Ventricle Outflow Tract
Fetal Cardiac Ultrasound Report | Fetal Echo | SV (MM-Teich) | none | calc | c_fet_svMM | none | unilateral | 125196\125209\F-32120\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Fetal Cardiac Ultrasound Report | Fetal Echo | SV (2D-Teich) | none | calc | c_fet_sv | none | unilateral | 125196\125209\F-32120\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Fetal Cardiac Ultrasound Report | Fetal Echo | SV (LVOT) | none | calc | c_fet_lvotSv | none | unilateral | 125196\F-32120\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 | SRT, T-32650, Left Ventricle Outflow Tract
Fetal Cardiac Ultrasound Report | Fetal Echo | E/E` Medial | none | calc | c_f_E_MedE | none | unilateral | 125196\G-037B\G-0391\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037B, Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave |  |  | UCUM, 1, no units | SRT, G-0391, Medial Mitral Annulus
Fetal Cardiac Ultrasound Report | Fetal Echo | E/E` Lateral | none | calc | c_f_E_LatE | none | unilateral | 125196\G-037B\G-0392\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037B, Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave |  |  | UCUM, 1, no units | SRT, G-0392, Lateral Mitral Annulus
Fetal Cardiac Ultrasound Report | Fetal Echo | LVOT Area | none | calc | c_fet_lvotArea | none | unilateral | 125196\G-038E\G-03A2\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | SRT, T-32650, Left Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | EDV (MM-Teich) | none | calc | c_fet_edvMM | none | unilateral | 125196\125209\F-32011\G-0394\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole |  | DCM, 125209, Teichholz
Fetal Cardiac Ultrasound Report | Fetal Echo | ESV (MM-Teich) | none | calc | c_fet_esvMM | none | unilateral | 125196\125209\G-0394\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125209, Teichholz
Fetal Cardiac Ultrasound Report | Fetal Echo | EDV (2D-Teich) | none | calc | c_fet_edv | none | unilateral | 125196\125209\F-32011\G-03A2\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | DCM, 125209, Teichholz
Fetal Cardiac Ultrasound Report | Fetal Echo | ESV (2D-Teich) | none | calc | c_fet_esv | none | unilateral | 125196\125209\G-03A2\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125209, Teichholz
Fetal Cardiac Ultrasound Report | Fetal Echo | TV E/A | none | calc | c_fet_tvEa | none | unilateral | 125196\59104-0\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59104-0, Peak E wave/Peak A wave by US |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | MV E/A | none | calc | c_fet_mvEa | none | unilateral | 125196\59104-0\T-35300 | SRT, T-35300, Mitral Valve | LN, 59104-0, Peak E wave/Peak A wave by US |  |  | UCUM, 1, no units
Fetal Cardiac Ultrasound Report | Fetal Echo | Ao/PA | none | calc | c_fet_aoPa2D | none | unilateral | 125196\C12211-08\G-03A2\T-42110 | SRT, T-42110, Root of Aorta | 99PMSBLUS, C12211-08, Aortic Root to Pulmonary Trunk ratio |  |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Fetal Cardiac Ultrasound Report | Fetal Echo | % Rev Flow | none | calc | c_fet_revFlow | none | unilateral | 125196\C12215-01\T-48710 | SRT, T-48710, Inferior vena cava | 99PMSBLUS, C12215-01, Reverse to Forward Velocity Time Integral (VTI) Ratio |  |  | UCUM, %, Percent
OB-GYN Ultrasound Procedure Report | GYN | Fibroid H1 | Dist | meas | m_Fibroidht1 | r_Distance | unilateral | 125000\C4-02\C7470-01\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | 99PMSBLUS, C7470-01, Height |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid H2 | Dist | meas | m_Fibroidht2 | r_Distance | unilateral | 125000\C4-02\C7470-01\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | 99PMSBLUS, C7470-01, Height |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid H3 | Dist | meas | m_Fibroidht3 | r_Distance | unilateral | 125000\C4-02\C7470-01\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | 99PMSBLUS, C7470-01, Height |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid H4 | Dist | meas | m_Fibroidht4 | r_Distance | unilateral | 125000\C4-02\C7470-01\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | 99PMSBLUS, C7470-01, Height |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid H5 | Dist | meas | m_Fibroidht5 | r_Distance | unilateral | 125000\C4-02\C7470-01\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | 99PMSBLUS, C7470-01, Height |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid H6 | Dist | meas | m_Fibroidht6 | r_Distance | unilateral | 125000\C4-02\C7470-01\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | 99PMSBLUS, C7470-01, Height |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid H7 | Dist | meas | m_Fibroidht7 | r_Distance | unilateral | 125000\C4-02\C7470-01\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | 99PMSBLUS, C7470-01, Height |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid H8 | Dist | meas | m_Fibroidht8 | r_Distance | unilateral | 125000\C4-02\C7470-01\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | 99PMSBLUS, C7470-01, Height |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid W1 | Dist | meas | m_Fibroidw1 | r_Distance | unilateral | 125000\C4-02\G-A220\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | SNM3, G-A220, Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid W2 | Dist | meas | m_Fibroidw2 | r_Distance | unilateral | 125000\C4-02\G-A220\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | SNM3, G-A220, Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid W3 | Dist | meas | m_Fibroidw3 | r_Distance | unilateral | 125000\C4-02\G-A220\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | SNM3, G-A220, Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid W4 | Dist | meas | m_Fibroidw4 | r_Distance | unilateral | 125000\C4-02\G-A220\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | SNM3, G-A220, Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid W5 | Dist | meas | m_Fibroidw5 | r_Distance | unilateral | 125000\C4-02\G-A220\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | SNM3, G-A220, Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid W6 | Dist | meas | m_Fibroidw6 | r_Distance | unilateral | 125000\C4-02\G-A220\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | SNM3, G-A220, Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid W7 | Dist | meas | m_Fibroidw7 | r_Distance | unilateral | 125000\C4-02\G-A220\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | SNM3, G-A220, Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid W8 | Dist | meas | m_Fibroidw8 | r_Distance | unilateral | 125000\C4-02\G-A220\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | SNM3, G-A220, Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid L1 | Dist | meas | m_Fibroidlen1 | r_Distance | unilateral | 125000\C4-02\G-A22A\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | SRT, G-A22A, Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid L2 | Dist | meas | m_Fibroidlen2 | r_Distance | unilateral | 125000\C4-02\G-A22A\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | SRT, G-A22A, Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid L3 | Dist | meas | m_Fibroidlen3 | r_Distance | unilateral | 125000\C4-02\G-A22A\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | SRT, G-A22A, Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid L4 | Dist | meas | m_Fibroidlen4 | r_Distance | unilateral | 125000\C4-02\G-A22A\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | SRT, G-A22A, Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid L5 | Dist | meas | m_Fibroidlen5 | r_Distance | unilateral | 125000\C4-02\G-A22A\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | SRT, G-A22A, Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid L6 | Dist | meas | m_Fibroidlen6 | r_Distance | unilateral | 125000\C4-02\G-A22A\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | SRT, G-A22A, Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid L7 | Dist | meas | m_Fibroidlen7 | r_Distance | unilateral | 125000\C4-02\G-A22A\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | SRT, G-A22A, Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid L8 | Dist | meas | m_Fibroidlen8 | r_Distance | unilateral | 125000\C4-02\G-A22A\xxx | 99PMSBLUS, C4-02, Uterine Fibroid | SRT, G-A22A, Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | L Ov Width | Dist | meas | m_ltOvwidth | r_Distance | unilateral | 125000\11829-9\T-87000 | SRT, T-87000, Ovary | LN, 11829-9, Left Ovary Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | R Ov Width | Dist | meas | m_ob_rtOvwidth | r_Distance | unilateral | 125000\11830-7\T-87000 | SRT, T-87000, Ovary | LN, 11830-7, Right Ovary Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | L Ov Length | Dist | meas | m_ltOvlen | r_Distance | unilateral | 125000\11840-6\T-87000 | SRT, T-87000, Ovary | LN, 11840-6, Left Ovary Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | R Ov Length | Dist | meas | m_ob_rtOvlen | r_Distance | unilateral | 125000\11841-4\T-87000 | SRT, T-87000, Ovary | LN, 11841-4, Right Ovary Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | L Ov Height | Dist | meas | m_ltOvheight | r_Distance | unilateral | 125000\11857-0\T-87000 | SRT, T-87000, Ovary | LN, 11857-0, Left Ovary Height |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | R Ov Height | Dist | meas | m_ob_rtOvheight | r_Distance | unilateral | 125000\11858-8\T-87000 | SRT, T-87000, Ovary | LN, 11858-8, Right Ovary Height |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Foll1 | Dist | meas | m_follicle1 | r_Distance | bilateral | 125000\11793-7\T-87600\xxx | SRT, T-87600, Ovarian Follicle | LN, 11793-7, Follicle Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Foll10 | Dist | meas | m_follicle10 | r_Distance | bilateral | 125000\11793-7\T-87600\xxx | SRT, T-87600, Ovarian Follicle | LN, 11793-7, Follicle Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Foll2 | Dist | meas | m_follicle2 | r_Distance | bilateral | 125000\11793-7\T-87600\xxx | SRT, T-87600, Ovarian Follicle | LN, 11793-7, Follicle Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Foll3 | Dist | meas | m_follicle3 | r_Distance | bilateral | 125000\11793-7\T-87600\xxx | SRT, T-87600, Ovarian Follicle | LN, 11793-7, Follicle Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Foll4 | Dist | meas | m_follicle4 | r_Distance | bilateral | 125000\11793-7\T-87600\xxx | SRT, T-87600, Ovarian Follicle | LN, 11793-7, Follicle Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Foll5 | Dist | meas | m_follicle5 | r_Distance | bilateral | 125000\11793-7\T-87600\xxx | SRT, T-87600, Ovarian Follicle | LN, 11793-7, Follicle Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Foll6 | Dist | meas | m_follicle6 | r_Distance | bilateral | 125000\11793-7\T-87600\xxx | SRT, T-87600, Ovarian Follicle | LN, 11793-7, Follicle Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Foll7 | Dist | meas | m_follicle7 | r_Distance | bilateral | 125000\11793-7\T-87600\xxx | SRT, T-87600, Ovarian Follicle | LN, 11793-7, Follicle Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Foll8 | Dist | meas | m_follicle8 | r_Distance | bilateral | 125000\11793-7\T-87600\xxx | SRT, T-87600, Ovarian Follicle | LN, 11793-7, Follicle Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Foll9 | Dist | meas | m_follicle9 | r_Distance | bilateral | 125000\11793-7\T-87600\xxx | SRT, T-87600, Ovarian Follicle | LN, 11793-7, Follicle Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | ? | meas | m_ob_rtUterArt | r_DopAngle | unilateral | 125000\C12122-01\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | EDV | meas | m_ob_rtUterArt | r_EDV | unilateral | 125000\11653-3\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | MDV | meas | m_ob_rtUterArt | r_MDV | unilateral | 125000\11665-7\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | TAPV | meas | m_ob_rtUterArt | r_TAPV | unilateral | 125000\11692-1\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | PSV | meas | m_ob_rtUterArt | r_PSV | unilateral | 125000\11726-7\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | PI | meas | m_ob_rtUterArt | r_PI | unilateral | 125000\12008-9\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | RI | meas | m_ob_rtUterArt | r_RI | unilateral | 125000\12023-8\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | S/D | meas | m_ob_rtUterArt | r_SD | unilateral | 125000\12144-2\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | AI | meas | m_ob_rtUterArt | r_AccelIndex | unilateral | 125000\20167-3\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | AT | meas | m_ob_rtUterArt | r_AccelTime | unilateral | 125000\20168-1\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | DT | meas | m_ob_rtUterArt | r_DecelTime | unilateral | 125000\20217-6\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | PG | meas | m_ob_rtUterArt | r_PG | unilateral | 125000\20247-3\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | MG | meas | m_ob_rtUterArt | r_PGmean | unilateral | 125000\20256-4\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | TAMV | meas | m_ob_rtUterArt | r_TAMV | unilateral | 125000\20352-1\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | VTI | meas | m_ob_rtUterArt | r_VTI | unilateral | 125000\20354-7\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | HR | meas | m_ob_rtUterArt | r_HeartRate | unilateral | 125000\8867-4\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine A | Dist | meas | m_ob_rtUterArt | r_Distance | unilateral | 125000\G-0364\G-A100\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [10 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | ? | meas | m_ob_ltUterArt | r_DopAngle | unilateral | 125000\C12122-01\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | EDV | meas | m_ob_ltUterArt | r_EDV | unilateral | 125000\11653-3\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | MDV | meas | m_ob_ltUterArt | r_MDV | unilateral | 125000\11665-7\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | TAPV | meas | m_ob_ltUterArt | r_TAPV | unilateral | 125000\11692-1\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | PSV | meas | m_ob_ltUterArt | r_PSV | unilateral | 125000\11726-7\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | PI | meas | m_ob_ltUterArt | r_PI | unilateral | 125000\12008-9\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | RI | meas | m_ob_ltUterArt | r_RI | unilateral | 125000\12023-8\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | S/D | meas | m_ob_ltUterArt | r_SD | unilateral | 125000\12144-2\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | AI | meas | m_ob_ltUterArt | r_AccelIndex | unilateral | 125000\20167-3\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | AT | meas | m_ob_ltUterArt | r_AccelTime | unilateral | 125000\20168-1\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | DT | meas | m_ob_ltUterArt | r_DecelTime | unilateral | 125000\20217-6\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | PG | meas | m_ob_ltUterArt | r_PG | unilateral | 125000\20247-3\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | MG | meas | m_ob_ltUterArt | r_PGmean | unilateral | 125000\20256-4\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | TAMV | meas | m_ob_ltUterArt | r_TAMV | unilateral | 125000\20352-1\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | VTI | meas | m_ob_ltUterArt | r_VTI | unilateral | 125000\20354-7\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | HR | meas | m_ob_ltUterArt | r_HeartRate | unilateral | 125000\8867-4\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine A | Dist | meas | m_ob_ltUterArt | r_Distance | unilateral | 125000\G-0364\G-A101\T-46820\T-D6007 | SRT, T-D6007, Pelvic Vascular Structure | SRT, T-46820, Uterine Artery | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [10 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | RLQ | Dist | meas | m_rlq | r_Distance | unilateral | 125000\11623-6\T-F1300 | SRT, T-F1300, Amniotic Sac | LN, 11623-6, Fourth Quadrant Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | RUQ | Dist | meas | m_ruq | r_Distance | unilateral | 125000\11624-4\T-F1300 | SRT, T-F1300, Amniotic Sac | LN, 11624-4, First Quadrant Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | LLQ | Dist | meas | m_llq | r_Distance | unilateral | 125000\11625-1\T-F1300 | SRT, T-F1300, Amniotic Sac | LN, 11625-1, Third Quadrant Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | LUQ | Dist | meas | m_luq | r_Distance | unilateral | 125000\11626-9\T-F1300 | SRT, T-F1300, Amniotic Sac | LN, 11626-9, Second Quadrant Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | ? | meas | m_ob_ductVen | r_DopAngle | unilateral | 125000\C12122-01\C12141-01\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | EDV | meas | m_ob_ductVen | r_EDV | unilateral | 125000\11653-3\C12141-01\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | MDV | meas | m_ob_ductVen | r_MDV | unilateral | 125000\11665-7\C12141-01\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | TAPV | meas | m_ob_ductVen | r_TAPV | unilateral | 125000\11692-1\C12141-01\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | PSV | meas | m_ob_ductVen | r_PSV | unilateral | 125000\11726-7\C12141-01\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | PI | meas | m_ob_ductVen | r_PI | unilateral | 125000\12008-9\C12141-01\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | RI | meas | m_ob_ductVen | r_RI | unilateral | 125000\12023-8\C12141-01\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | S/D | meas | m_ob_ductVen | r_SD | unilateral | 125000\12144-2\C12141-01\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | AI | meas | m_ob_ductVen | r_AccelIndex | unilateral | 125000\20167-3\C12141-01\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | AT | meas | m_ob_ductVen | r_AccelTime | unilateral | 125000\20168-1\C12141-01\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | DT | meas | m_ob_ductVen | r_DecelTime | unilateral | 125000\20217-6\C12141-01\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | PG | meas | m_ob_ductVen | r_PG | unilateral | 125000\20247-3\C12141-01\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | MG | meas | m_ob_ductVen | r_PGmean | unilateral | 125000\20256-4\C12141-01\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | TAMV | meas | m_ob_ductVen | r_TAMV | unilateral | 125000\20352-1\C12141-01\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | VTI | meas | m_ob_ductVen | r_VTI | unilateral | 125000\20354-7\C12141-01\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | HR | meas | m_ob_ductVen | r_HeartRate | unilateral | 125000\8867-4\C12141-01\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm
OB-GYN Ultrasound Procedure Report | OB | Duct Ven | Dist | meas | m_ob_ductVen | r_Distance | unilateral | 125000\C12141-01\G-0364\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | 99PMSBLUS, C12141-01, Ductus Venosus | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Aorta | ? | meas | m_ob_Ao | r_DopAngle | unilateral | 125000\C12122-01\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg
OB-GYN Ultrasound Procedure Report | OB | Aorta | EDV | meas | m_ob_Ao | r_EDV | unilateral | 125000\11653-3\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Aorta | MDV | meas | m_ob_Ao | r_MDV | unilateral | 125000\11665-7\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Aorta | TAPV | meas | m_ob_Ao | r_TAPV | unilateral | 125000\11692-1\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Aorta | PSV | meas | m_ob_Ao | r_PSV | unilateral | 125000\11726-7\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Aorta | PI | meas | m_ob_Ao | r_PI | unilateral | 125000\12008-9\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Aorta | RI | meas | m_ob_Ao | r_RI | unilateral | 125000\12023-8\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Aorta | S/D | meas | m_ob_Ao | r_SD | unilateral | 125000\12144-2\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Aorta | AI | meas | m_ob_Ao | r_AccelIndex | unilateral | 125000\20167-3\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2
OB-GYN Ultrasound Procedure Report | OB | Aorta | AT | meas | m_ob_Ao | r_AccelTime | unilateral | 125000\20168-1\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | Aorta | DT | meas | m_ob_Ao | r_DecelTime | unilateral | 125000\20217-6\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | Aorta | PG | meas | m_ob_Ao | r_PG | unilateral | 125000\20247-3\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg
OB-GYN Ultrasound Procedure Report | OB | Aorta | MG | meas | m_ob_Ao | r_PGmean | unilateral | 125000\20256-4\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg
OB-GYN Ultrasound Procedure Report | OB | Aorta | TAMV | meas | m_ob_Ao | r_TAMV | unilateral | 125000\20352-1\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Aorta | VTI | meas | m_ob_Ao | r_VTI | unilateral | 125000\20354-7\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Aorta | HR | meas | m_ob_Ao | r_HeartRate | unilateral | 125000\8867-4\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm
OB-GYN Ultrasound Procedure Report | OB | Aorta | Dist | meas | m_ob_Ao | r_Distance | unilateral | 125000\G-0364\T-42000\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42000, Aorta | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | ? | meas | m_ob_thrAo | r_DopAngle | unilateral | 125000\C12122-01\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | EDV | meas | m_ob_thrAo | r_EDV | unilateral | 125000\11653-3\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | MDV | meas | m_ob_thrAo | r_MDV | unilateral | 125000\11665-7\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | TAPV | meas | m_ob_thrAo | r_TAPV | unilateral | 125000\11692-1\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | PSV | meas | m_ob_thrAo | r_PSV | unilateral | 125000\11726-7\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | PI | meas | m_ob_thrAo | r_PI | unilateral | 125000\12008-9\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | RI | meas | m_ob_thrAo | r_RI | unilateral | 125000\12023-8\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | S/D | meas | m_ob_thrAo | r_SD | unilateral | 125000\12144-2\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | AI | meas | m_ob_thrAo | r_AccelIndex | unilateral | 125000\20167-3\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | AT | meas | m_ob_thrAo | r_AccelTime | unilateral | 125000\20168-1\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | DT | meas | m_ob_thrAo | r_DecelTime | unilateral | 125000\20217-6\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | PG | meas | m_ob_thrAo | r_PG | unilateral | 125000\20247-3\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | MG | meas | m_ob_thrAo | r_PGmean | unilateral | 125000\20256-4\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | TAMV | meas | m_ob_thrAo | r_TAMV | unilateral | 125000\20352-1\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | VTI | meas | m_ob_thrAo | r_VTI | unilateral | 125000\20354-7\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | HR | meas | m_ob_thrAo | r_HeartRate | unilateral | 125000\8867-4\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm
OB-GYN Ultrasound Procedure Report | OB | Thoracic Ao | Dist | meas | m_ob_thrAo | r_Distance | unilateral | 125000\G-0364\T-42070\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-42070, Thoracic aorta | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | MCA | ? | meas | m_mca | r_DopAngle | unilateral | 125000\C12122-01\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg
OB-GYN Ultrasound Procedure Report | OB | MCA | EDV | meas | m_mca | r_EDV | unilateral | 125000\11653-3\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | MCA | MDV | meas | m_mca | r_MDV | unilateral | 125000\11665-7\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | MCA | TAPV | meas | m_mca | r_TAPV | unilateral | 125000\11692-1\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | MCA | PSV | meas | m_mca | r_PSV | unilateral | 125000\11726-7\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | MCA | PI | meas | m_mca | r_PI | unilateral | 125000\12008-9\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | MCA | RI | meas | m_mca | r_RI | unilateral | 125000\12023-8\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | MCA | S/D | meas | m_mca | r_SD | unilateral | 125000\12144-2\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | MCA | AI | meas | m_mca | r_AccelIndex | unilateral | 125000\20167-3\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2
OB-GYN Ultrasound Procedure Report | OB | MCA | AT | meas | m_mca | r_AccelTime | unilateral | 125000\20168-1\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | MCA | DT | meas | m_mca | r_DecelTime | unilateral | 125000\20217-6\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | MCA | PG | meas | m_mca | r_PG | unilateral | 125000\20247-3\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg
OB-GYN Ultrasound Procedure Report | OB | MCA | MG | meas | m_mca | r_PGmean | unilateral | 125000\20256-4\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg
OB-GYN Ultrasound Procedure Report | OB | MCA | TAMV | meas | m_mca | r_TAMV | unilateral | 125000\20352-1\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | MCA | VTI | meas | m_mca | r_VTI | unilateral | 125000\20354-7\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | MCA | HR | meas | m_mca | r_HeartRate | unilateral | 125000\8867-4\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm
OB-GYN Ultrasound Procedure Report | OB | MCA | Dist | meas | m_mca | r_Distance | unilateral | 125000\G-0364\T-45600\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-45600, Middle Cerebral Artery | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | ? | meas | m_ob_descAo | r_DopAngle | unilateral | 125000\C12122-01\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | EDV | meas | m_ob_descAo | r_EDV | unilateral | 125000\11653-3\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | MDV | meas | m_ob_descAo | r_MDV | unilateral | 125000\11665-7\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | TAPV | meas | m_ob_descAo | r_TAPV | unilateral | 125000\11692-1\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | PSV | meas | m_ob_descAo | r_PSV | unilateral | 125000\11726-7\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | PI | meas | m_ob_descAo | r_PI | unilateral | 125000\12008-9\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | RI | meas | m_ob_descAo | r_RI | unilateral | 125000\12023-8\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | S/D | meas | m_ob_descAo | r_SD | unilateral | 125000\12144-2\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | AI | meas | m_ob_descAo | r_AccelIndex | unilateral | 125000\20167-3\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | AT | meas | m_ob_descAo | r_AccelTime | unilateral | 125000\20168-1\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | DT | meas | m_ob_descAo | r_DecelTime | unilateral | 125000\20217-6\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | PG | meas | m_ob_descAo | r_PG | unilateral | 125000\20247-3\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | MG | meas | m_ob_descAo | r_PGmean | unilateral | 125000\20256-4\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | TAMV | meas | m_ob_descAo | r_TAMV | unilateral | 125000\20352-1\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | VTI | meas | m_ob_descAo | r_VTI | unilateral | 125000\20354-7\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | HR | meas | m_ob_descAo | r_HeartRate | unilateral | 125000\8867-4\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm
OB-GYN Ultrasound Procedure Report | OB | Desc Ao | Dist | meas | m_ob_descAo | r_Distance | unilateral | 125000\G-0364\T-D0765\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-D0765, Descending Aorta | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | ? | meas | m_umbArt | r_DopAngle | unilateral | 125000\C12122-01\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | EDV | meas | m_umbArt | r_EDV | unilateral | 125000\11653-3\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | MDV | meas | m_umbArt | r_MDV | unilateral | 125000\11665-7\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | TAPV | meas | m_umbArt | r_TAPV | unilateral | 125000\11692-1\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | PSV | meas | m_umbArt | r_PSV | unilateral | 125000\11726-7\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | PI | meas | m_umbArt | r_PI | unilateral | 125000\12008-9\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | RI | meas | m_umbArt | r_RI | unilateral | 125000\12023-8\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | S/D | meas | m_umbArt | r_SD | unilateral | 125000\12144-2\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | AI | meas | m_umbArt | r_AccelIndex | unilateral | 125000\20167-3\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | AT | meas | m_umbArt | r_AccelTime | unilateral | 125000\20168-1\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | DT | meas | m_umbArt | r_DecelTime | unilateral | 125000\20217-6\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | PG | meas | m_umbArt | r_PG | unilateral | 125000\20247-3\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | MG | meas | m_umbArt | r_PGmean | unilateral | 125000\20256-4\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | TAMV | meas | m_umbArt | r_TAMV | unilateral | 125000\20352-1\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | VTI | meas | m_umbArt | r_VTI | unilateral | 125000\20354-7\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | HR | meas | m_umbArt | r_HeartRate | unilateral | 125000\8867-4\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm
OB-GYN Ultrasound Procedure Report | OB | Umbilical A | Dist | meas | m_umbArt | r_Distance | unilateral | 125000\G-0364\T-F1810\T-F6800 | SRT, T-F6800, Embryonic Vascular Structure | SRT, T-F1810, Umbilical Artery | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | RV Dimension | Dist | meas | m_ob_rvDimen2D | r_Distance | unilateral | 125000\99999\P5000-01-01 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-01, Right Ventricle Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | RA Dimension | Dist | meas | m_ob_raDimen2D | r_Distance | unilateral | 125000\99999\P5000-01-02 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-02, Right Atrium Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | RVOT Diam | Dist | meas | m_ob_rvotDiam | r_Distance | unilateral | 125000\99999\P5000-01-03 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-03, RV Outflow Tract Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | LV Dimension | Dist | meas | m_ob_lvDimen2D | r_Distance | unilateral | 125000\99999\P5000-01-04 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-04, Left Ventricle Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | LA Dimension | Dist | meas | m_ob_laDimen2D | r_Distance | unilateral | 125000\99999\P5000-01-05 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-05, Left Atrium Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | LVOT Diam | Dist | meas | m_ob_lvotDiam | r_Distance | unilateral | 125000\99999\P5000-01-06 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-06, LV Outflow Tract Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | IVS (2D) | Dist | meas | m_ob_ivs2D | r_Distance | unilateral | 125000\99999\P5000-01-08 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-08, Interventricular Septum Thickness |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | LA Length | Dist | meas | m_ob_laLen | r_Distance | unilateral | 125000\99999\P5000-01-13 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-13, Left Atrium Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | LA Width | Dist | meas | m_ob_laWidth | r_Distance | unilateral | 125000\99999\P5000-01-14 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-14, Left Atrium Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | RA Length | Dist | meas | m_ob_raLen | r_Distance | unilateral | 125000\99999\P5000-01-15 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-15, Right Atrium Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | RA Width | Dist | meas | m_ob_raWidth | r_Distance | unilateral | 125000\99999\P5000-01-16 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-16, Right Atrium Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | LV Length | Dist | meas | m_ob_lvLen | r_Distance | unilateral | 125000\99999\P5000-01-17 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-17, Left Ventricle Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | LV Width | Dist | meas | m_ob_lvWidth | r_Distance | unilateral | 125000\99999\P5000-01-18 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-18, Left Ventricle Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | RV Length | Dist | meas | m_ob_rvLen | r_Distance | unilateral | 125000\99999\P5000-01-19 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-19, Right Ventricle Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | RV Width | Dist | meas | m_ob_rvWidth | r_Distance | unilateral | 125000\99999\P5000-01-20 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-20, Right Ventricle Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Ao Annul Diam | Dist | meas | m_ob_aoAnn | r_Distance | unilateral | 125000\99999\P5000-01-21 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-21, Aortic Annulus Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | MV Annul Diam | Dist | meas | m_ob_mvAnn | r_Distance | unilateral | 125000\99999\P5000-01-22 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-22, Mitral Annulus Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | TV Annul Diam | Dist | meas | m_ob_tvAnn | r_Distance | unilateral | 125000\99999\P5000-01-23 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-23, Tricuspid Annulus Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Duct Art Diam (2D) | Dist | meas | m_ob_ductArt2D | r_Distance | unilateral | 125000\99999\P5000-01-24 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-24, Ductus Arteriosus Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Cardiac Axis | Angle | meas | m_ob_Cardiac_Axis | r_Angle | unilateral | 125000\99999\P5000-01-25 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-25, Cardiac Axis |  |  | UCUM, deg, deg
OB-GYN Ultrasound Procedure Report | OB | LVET | Time | meas | m_ob_LVET | r_DopTime | unilateral | 125000\122211\99999 | 99PMSBLUS, 99999, Fetal Heart | DCM, 122211, Left Ventricular ejection time |  |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | RVET | Time | meas | m_ob_RVET | r_DopTime | unilateral | 125000\122213\99999 | 99PMSBLUS, 99999, Fetal Heart | DCM, 122213, Right Ventricular ejection time |  |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | Asc Ao Diam (2D) | Dist | meas | m_ob_ascAo2D | r_Distance | unilateral | 125000\18012-5\99999 | 99PMSBLUS, 99999, Fetal Heart | LN, 18012-5, Ascending Aortic Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Desc Ao Diam (2D) | Dist | meas | m_ob_descAo2D | r_Distance | unilateral | 125000\18013-3\99999 | 99PMSBLUS, 99999, Fetal Heart | LN, 18013-3, Descending Aortic Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | AoR Diam (2D) | Dist | meas | m_ob_aoR2D | r_Distance | unilateral | 125000\18015-8\99999 | 99PMSBLUS, 99999, Fetal Heart | LN, 18015-8, Aortic Root Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | MPA Diam (2D) | Dist | meas | m_ob_mpa2D | r_Distance | unilateral | 125000\18020-8\99999 | 99PMSBLUS, 99999, Fetal Heart | LN, 18020-8, Main Pulmonary Artery Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | IVRT | Time | meas | m_ob_IVRT | r_DopTime | unilateral | 125000\59083-6\99999 | 99PMSBLUS, 99999, Fetal Heart | LN, 59083-6, Isovolumic Relaxation Time |  |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | IVCT | Time | meas | m_ob_IVCT | r_DopTime | unilateral | 125000\59084-4\99999 | 99PMSBLUS, 99999, Fetal Heart | LN, 59084-4, Isovolumic Contraction Time |  |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | PR Interval | Time | meas | m_ob_PR_Interval | r_DopTime | unilateral | 125000\5.13.5-7\99999 | 99PMSBLUS, 99999, Fetal Heart | SCPECG, 5.13.5-7, PR Interval |  |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | MCOT | Time | meas | m_ob_MCOT | r_DopTime | unilateral | 125000\99999\G-0387 | 99PMSBLUS, 99999, Fetal Heart | SRT, G-0387, Mitral Valve Closure to Opening Time |  |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | TCOT | Time | meas | m_ob_TCOT | r_DopTime | unilateral | 125000\99999\G-0389 | 99PMSBLUS, 99999, Fetal Heart | SRT, G-0389, Tricuspid Valve Closure to Opening Time |  |  | UCUM, ms, ms
OB-GYN Ultrasound Procedure Report | OB | Ear | Dist | meas | m_ear | r_Distance | unilateral | 125000\125002\C12005-01 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-01, Ear length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | FTA | Area | meas | m_fta | r_Area | unilateral | 125000\125002\C12005-02 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-02, Fetal Trunk Cross Sectional Area |  |  | UCUM, mm2, mm2
OB-GYN Ultrasound Procedure Report | OB | Heart Circ (2D) | Circ | meas | m_ob_hrtCirc | r_Circumference | unilateral | 125000\125002\C12005-03 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-03, Heart Circumference |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | M Phalanx 5 | Dist | meas | m_mphalanx5 | r_Distance | unilateral | 125000\125002\C12005-04 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-04, Length of middle Phalanx of the 5th Digit |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | APThD | Dist | meas | m_aptd | r_Distance | unilateral | 125000\125002\C12005-07 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-07, Anterior-Posterior Thoracic Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | TTD | Dist | meas | m_ttd | r_Distance | unilateral | 125000\125002\C12005-08 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-08, Transverse Trunk Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Heart Area (2D) | Area | meas | m_ob_hrtArea | r_Area | unilateral | 125000\125002\C12005-09 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-09, Heart Area |  |  | UCUM, mm2, mm2
OB-GYN Ultrasound Procedure Report | OB | FMF Angle | Angle | meas | m_FMFAngle | r_Angle | unilateral | 125000\125002\C12005-11 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-11, Frontomaxillary Facial Angle |  |  | UCUM, deg, deg
OB-GYN Ultrasound Procedure Report | OB | Inf Fac Angle | Angle | meas | m_IFAngle | r_Angle | unilateral | 125000\125002\C12005-13 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-13, Inferior Facial Angle |  |  | UCUM, deg, deg
OB-GYN Ultrasound Procedure Report | OB | Iliac Angle | Angle | meas | m_IliacAngle | r_Angle | unilateral | 125000\125002\C12005-15 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-15, Iliac Angle |  |  | UCUM, deg, deg
OB-GYN Ultrasound Procedure Report | OB | Mandible | Dist | meas | m_mandible | r_Distance | unilateral | 125000\125002\C12005-23 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-23, Mandible Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | L Lung Diam | Dist | meas | m_leftLungDiam | r_Distance | unilateral | 125000\125002\C12005-24 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-24, Left Lung Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | R Lung Diam | Dist | meas | m_rtLungDiam | r_Distance | unilateral | 125000\125002\C12005-25 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-25, Right Lung Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Post Fossa | Dist | meas | m_postfossa | r_Distance | unilateral | 125000\125002\C12005-26 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-26, Post Fossa Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Scapula | Dist | meas | m_scapula | r_Distance | unilateral | 125000\125002\C12005-27 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-27, Scapula Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Iliac Crest | Dist | meas | m_iliacCrest | r_Distance | unilateral | 125000\125002\C12005-28 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-28, Iliac Crest Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Renal L | Dist | meas | m_renalLat | r_Distance | unilateral | 125000\125002\C12011-09 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12011-09, Renal Longitudinal Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Renal AP | Dist | meas | m_renalAp | r_Distance | unilateral | 125000\125002\C12011-10 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12011-10, Renal Antero-posterior Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Renal Tr | Dist | meas | m_renalTr | r_Distance | unilateral | 125000\125002\C12011-11 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12011-11, Renal Transverse Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Renal Pelvis | Dist | meas | m_renalPelvis | r_Distance | unilateral | 125000\125002\C12011-12 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12011-12, Renal Pelvis Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Pelvis L | Dist | meas | m_pelvisLat | r_Distance | unilateral | 125000\125002\C12011-13 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12011-13, Pelvis Longitudinal Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Pelvis AP | Dist | meas | m_pelvisAp | r_Distance | unilateral | 125000\125002\C12011-14 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12011-14, Pelvis Antero-posterior Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Pelvis Tr | Dist | meas | m_pelvisTr | r_Distance | unilateral | 125000\125002\C12011-15 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12011-15, Pelvis Transverse Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Ureter AP | Dist | meas | m_ureterAp | r_Distance | unilateral | 125000\125002\C12011-16 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12011-16, Ureter Antero-posterior Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Ureter Tr | Dist | meas | m_ureterTr | r_Distance | unilateral | 125000\125002\C12011-17 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12011-17, Ureter Transverse Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Bladder L | Dist | meas | m_bladderLat | r_Distance | unilateral | 125000\125002\C12011-18 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12011-18, Bladder Longitudinal Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Bladder AP | Dist | meas | m_bladderAp | r_Distance | unilateral | 125000\125002\C12011-19 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12011-19, Bladder Antero-posterior Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Bladder Tr | Dist | meas | m_bladderTr | r_Distance | unilateral | 125000\125002\C12011-20 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12011-20, Bladder Transverse Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Adr Gland L | Dist | meas | m_adrGlandLat | r_Distance | unilateral | 125000\125002\C12011-21 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12011-21, Adrenal Gland Longitudinal Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Adr Gland AP | Dist | meas | m_adrGlandAp | r_Distance | unilateral | 125000\125002\C12011-22 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12011-22, Adrenal Gland Antero-posterior Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Adr Gland Tr | Dist | meas | m_adrGlandTr | r_Distance | unilateral | 125000\125002\C12011-23 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12011-23, Adrenal Gland Transverse Dimension |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Head Angle | Angle | meas | m_HeadAngle | r_Angle | unilateral | 125000\125002\C12019-06 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12019-06, Head Angle |  |  | UCUM, deg, deg
OB-GYN Ultrasound Procedure Report | OB | APAD | Dist | meas | m_apd | r_Distance | unilateral | 125000\11818-2\125002 | DCM, 125002, Fetal Biometry | LN, 11818-2, Anterior-Posterior Abdominal Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | APTD | Dist | meas | m_ap_trunk_diam | r_Distance | unilateral | 125000\11819-0\125002 | DCM, 125002, Fetal Biometry | LN, 11819-0, Anterior-Posterior Trunk Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | BPD | Dist | meas | m_bpd | r_Distance | unilateral | 125000\11820-8\125002 | DCM, 125002, Fetal Biometry | LN, 11820-8, Biparietal Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | OFD | Dist | meas | m_ofd | r_Distance | unilateral | 125000\11851-3\125002 | DCM, 125002, Fetal Biometry | LN, 11851-3, Occipital-Frontal Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | TAD | Dist | meas | m_tad | r_Distance | unilateral | 125000\11862-0\125002 | DCM, 125002, Fetal Biometry | LN, 11862-0, Transverse Abdominal Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Cerebellum | Dist | meas | m_cerebellum | r_Distance | unilateral | 125000\11863-8\125002 | DCM, 125002, Fetal Biometry | LN, 11863-8, Trans Cerebellar Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | TThD | Dist | meas | m_TrThD | r_Distance | unilateral | 125000\11864-6\125002 | DCM, 125002, Fetal Biometry | LN, 11864-6, Transverse Thoracic Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | FL | Dist | meas | m_fl | r_Distance | unilateral | 125000\11963-6\125002 | DCM, 125002, Fetal Biometry | LN, 11963-6, Femur Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Foot | Dist | meas | m_foot | r_Distance | unilateral | 125000\11965-1\125002 | DCM, 125002, Fetal Biometry | LN, 11965-1, Foot length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | AC | Circ | meas | m_ac | r_Circumference | unilateral | 125000\11979-2\125002 | DCM, 125002, Fetal Biometry | LN, 11979-2, Abdominal Circumference |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | HC | Circ | meas | m_hc | r_Circumference | unilateral | 125000\11984-2\125002 | DCM, 125002, Fetal Biometry | LN, 11984-2, Head Circumference |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | ThC | Circ | meas | m_tc | r_Circumference | unilateral | 125000\11988-3\125002 | DCM, 125002, Fetal Biometry | LN, 11988-3, Thoracic Circumference |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Nasal | Dist | meas | m_nasal | r_Distance | unilateral | 125000\125003\C12006-01 | DCM, 125003, Fetal Long Bones | 99PMSBLUS, C12006-01, Nasal Bone Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Clavicle | Dist | meas | m_clavicle | r_Distance | unilateral | 125000\11962-8\125003 | DCM, 125003, Fetal Long Bones | LN, 11962-8, Clavicle length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Fibula | Dist | meas | m_fibula | r_Distance | unilateral | 125000\11964-4\125003 | DCM, 125003, Fetal Long Bones | LN, 11964-4, Fibula length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Humerus | Dist | meas | m_humerus | r_Distance | unilateral | 125000\11966-9\125003 | DCM, 125003, Fetal Long Bones | LN, 11966-9, Humerus length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Radius | Dist | meas | m_radius | r_Distance | unilateral | 125000\11967-7\125003 | DCM, 125003, Fetal Long Bones | LN, 11967-7, Radius length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Tibia | Dist | meas | m_tibia | r_Distance | unilateral | 125000\11968-5\125003 | DCM, 125003, Fetal Long Bones | LN, 11968-5, Tibia length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Ulna | Dist | meas | m_ulna | r_Distance | unilateral | 125000\11969-3\125003 | DCM, 125003, Fetal Long Bones | LN, 11969-3, Ulna length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Orbit1 | Dist | meas | m_orbit1 | r_Distance | unilateral | 125000\125004\C12007-01 | DCM, 125004, Fetal Cranium | 99PMSBLUS, C12007-01, Diameter of the First Orbit |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Orbit2 | Dist | meas | m_orbit2 | r_Distance | unilateral | 125000\125004\C12007-02 | DCM, 125004, Fetal Cranium | 99PMSBLUS, C12007-02, Diameter of the Second Orbit |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | OOD | Dist | meas | m_oodiam | r_Distance | unilateral | 125000\11629-3\125004 | DCM, 125004, Fetal Cranium | LN, 11629-3, Outer Orbital Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Cist Mag | Dist | meas | m_cistmag | r_Distance | unilateral | 125000\11860-4\125004 | DCM, 125004, Fetal Cranium | LN, 11860-4, Cisterna Magna length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Nuch Fold | Dist | meas | m_nuchfold | r_Distance | unilateral | 125000\12146-7\125004 | DCM, 125004, Fetal Cranium | LN, 12146-7, Nuchal Fold thickness |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Lat Vent | Dist | meas | m_latventricle | r_Distance | unilateral | 125000\12171-5\125004 | DCM, 125004, Fetal Cranium | LN, 12171-5, Lateral Ventricle width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | IOD | Dist | meas | m_iod | r_Distance | unilateral | 125000\125004\33070-4 | DCM, 125004, Fetal Cranium | LN, 33070-4, Inner Orbital Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Fluid Volume | bpp | meas | m_afv | r_bpp | unilateral | 125000\11630-1\125006 | DCM, 125006, Biophysical Profile | LN, 11630-1, Amniotic Fluid Volume |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Movement | bpp | meas | m_fetalmovement | r_bpp | unilateral | 125000\11631-9\125006 | DCM, 125006, Biophysical Profile | LN, 11631-9, Gross Body Movement |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Breathing | bpp | meas | m_fetalbreathing | r_bpp | unilateral | 125000\11632-7\125006 | DCM, 125006, Biophysical Profile | LN, 11632-7, Fetal Breathing |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Tone | bpp | meas | m_fetaltone | r_bpp | unilateral | 125000\11635-0\125006 | DCM, 125006, Biophysical Profile | LN, 11635-0, Fetal Tone |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Heart Rate | HR | meas | m_fetalbiohr | r_HeartRate | unilateral | 125000\11948-7\125008 | DCM, 125008, Fetus Summary | LN, 11948-7, Fetal Heart Rate |  |  | UCUM, {H.B.}/min, bpm
OB-GYN Ultrasound Procedure Report | OB | EMM | Dist | meas | m_emm | r_Distance | unilateral | 125000\125009\C12009-05 | DCM, 125009, Early Gestation | 99PMSBLUS, C12009-05, Endomyometrial Mantle |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Yolk Sac | Dist | meas | m_yolksac | r_Distance | unilateral | 125000\11816-6\125009 | DCM, 125009, Early Gestation | LN, 11816-6, Yolk Sac length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Gest Sac | Dist | meas | m_gestsac | r_Distance | unilateral | 125000\11850-5\125009 | DCM, 125009, Early Gestation | LN, 11850-5, Gestational Sac Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Sac Diam1 | Dist | meas | m_sacdiam1 | r_Distance | unilateral | 125000\11850-5\125009 | DCM, 125009, Early Gestation | LN, 11850-5, Gestational Sac Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Sac Diam2 | Dist | meas | m_sacdiam2 | r_Distance | unilateral | 125000\11850-5\125009 | DCM, 125009, Early Gestation | LN, 11850-5, Gestational Sac Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Sac Diam3 | Dist | meas | m_sacdiam3 | r_Distance | unilateral | 125000\11850-5\125009 | DCM, 125009, Early Gestation | LN, 11850-5, Gestational Sac Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | CRL | Dist | meas | m_crl | r_Distance | unilateral | 125000\11957-8\125009 | DCM, 125009, Early Gestation | LN, 11957-8, Crown Rump Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | Nuch Luc | Dist | meas | m_nuchluc | r_Distance | unilateral | 125000\125009\33069-6 | DCM, 125009, Early Gestation | LN, 33069-6, Nuchal Translucency |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | SL | Dist | meas | m_sl | r_Distance | unilateral | 125000\125009\33071-2 | DCM, 125009, Early Gestation | LN, 33071-2, Spine Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Bladder L | Dist | meas | m_gyn_blad_len | r_Distance | unilateral | 125000\125011\C12011-01 | DCM, 125011, Pelvis and Uterus | 99PMSBLUS, C12011-01, Bladder Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Bladder W | Dist | meas | m_gyn_blad_wid | r_Distance | unilateral | 125000\125011\C12011-02 | DCM, 125011, Pelvis and Uterus | 99PMSBLUS, C12011-02, Bladder Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Bladder H | Dist | meas | m_gyn_blad_ht | r_Distance | unilateral | 125000\125011\C12011-03 | DCM, 125011, Pelvis and Uterus | 99PMSBLUS, C12011-03, Bladder Height |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | PV Bladder L | Dist | meas | m_gyn_pvblad_len | r_Distance | unilateral | 125000\125011\C12011-05 | DCM, 125011, Pelvis and Uterus | 99PMSBLUS, C12011-05, Post Void Bladder Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | PV Bladder W | Dist | meas | m_gyn_pvblad_wid | r_Distance | unilateral | 125000\125011\C12011-06 | DCM, 125011, Pelvis and Uterus | 99PMSBLUS, C12011-06, Post Void Bladder Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | PV Bladder H | Dist | meas | m_gyn_pvblad_ht | r_Distance | unilateral | 125000\125011\C12011-07 | DCM, 125011, Pelvis and Uterus | 99PMSBLUS, C12011-07, Post Void Bladder Height |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Cerv Length | Dist | meas | m_cervixlen | r_Distance | unilateral | 125000\11961-0\125011 | DCM, 125011, Pelvis and Uterus | LN, 11961-0, Cervix Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Endo Thick | Dist | meas | m_endomthick | r_Distance | unilateral | 125000\12145-9\125011 | DCM, 125011, Pelvis and Uterus | LN, 12145-9, Endometrium Thickness |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Uterus Length | Dist | meas | m_uteruslen | r_Distance | unilateral | 125000\11842-2\125011 | DCM, 125011, Pelvis and Uterus | LN, 11842-2, Uterus Length |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Uterus Height | Dist | meas | m_uterusheight | r_Distance | unilateral | 125000\11859-6\125011 | DCM, 125011, Pelvis and Uterus | LN, 11859-6, Uterus Height |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Uterus Width | Dist | meas | m_uteruswidth | r_Distance | unilateral | 125000\11865-3\125011 | DCM, 125011, Pelvis and Uterus | LN, 11865-3, Uterus Width |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Fibroid Vol 1 | none | calc | c_fibroidvol1 | none | unilateral | 125000\C4-02\G-D705\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Fibroid Vol 2 | none | calc | c_fibroidvol2 | none | unilateral | 125000\C4-02\G-D705\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Fibroid Vol 3 | none | calc | c_fibroidvol3 | none | unilateral | 125000\C4-02\G-D705\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Fibroid Vol 4 | none | calc | c_fibroidvol4 | none | unilateral | 125000\C4-02\G-D705\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Fibroid Vol 5 | none | calc | c_fibroidvol5 | none | unilateral | 125000\C4-02\G-D705\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Fibroid Vol 6 | none | calc | c_fibroidvol6 | none | unilateral | 125000\C4-02\G-D705\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Fibroid Vol 7 | none | calc | c_fibroidvol7 | none | unilateral | 125000\C4-02\G-D705\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Fibroid Vol 8 | none | calc | c_fibroidvol8 | none | unilateral | 125000\C4-02\G-D705\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | L Ov Volume | none | calc | c_ltOvVol | none | unilateral | 125000\12164-0\T-87000 |  | LN, 12164-0, Left Ovary Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | OB | R Ov Volume | none | calc | c_ob_rtOvVol | none | unilateral | 125000\12165-7\T-87000 |  | LN, 12165-7, Right Ovary Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Foll1 Volume | none | calc | c_foll1vol | none | bilateral | 125000\G-D705\T-87600\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Foll10 Volume | none | calc | c_follavol | none | bilateral | 125000\G-D705\T-87600\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Foll2 Volume | none | calc | c_foll2vol | none | bilateral | 125000\G-D705\T-87600\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Foll3 Volume | none | calc | c_foll3vol | none | bilateral | 125000\G-D705\T-87600\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Foll4 Volume | none | calc | c_foll4vol | none | bilateral | 125000\G-D705\T-87600\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Foll5 Volume | none | calc | c_foll5vol | none | bilateral | 125000\G-D705\T-87600\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Foll6 Volume | none | calc | c_foll6vol | none | bilateral | 125000\G-D705\T-87600\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Foll7 Volume | none | calc | c_foll7vol | none | bilateral | 125000\G-D705\T-87600\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Foll8 Volume | none | calc | c_foll8vol | none | bilateral | 125000\G-D705\T-87600\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Foll9 Volume | none | calc | c_foll9vol | none | bilateral | 125000\G-D705\T-87600\xxx |  | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | OB | AFI | none | calc | c_afi | none | unilateral | 125000\11627-7\T-F1300 |  | LN, 11627-7, Amniotic Fluid Index |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | MCA RI | none | calc | c_mca_ri | none | unilateral | 125000\C12121-02\T-F6800 |  | 99PMSBLUS, C12121-02, Middle Cerebral Artery Resistivity Index |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | MCA PI | none | calc | c_mca_pi | none | unilateral | 125000\C12121-03\T-F6800 |  | 99PMSBLUS, C12121-03, Middle Cerebral Artery Pulsatility Index |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Umb RI | none | calc | c_umb_ri | none | unilateral | 125000\C12121-04\T-F6800 |  | 99PMSBLUS, C12121-04, Umbilical Artery Resistivity Index |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Umb PI | none | calc | c_umb_pi | none | unilateral | 125000\C12121-05\T-F6800 |  | 99PMSBLUS, C12121-05, Umbilical Artery Pulsatility Index |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | LV/RV | none | calc | c_ob_lvRv2D | none | unilateral | 125000\99999\P5000-01-09 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-09, Ratio of LV Dimension to RV Dimension |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | LA/RA | none | calc | c_ob_laRa2D | none | unilateral | 125000\99999\P5000-01-10 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-10, Ratio of LA Dimension to RA Dimension |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Ao/PA | none | calc | c_ob_aoPa2D | none | unilateral | 125000\99999\P5000-01-11 | 99PMSBLUS, 99999, Fetal Heart | 99PMSBLUS, P5000-01-11, Ratio of Aortic Root Diameter to Main PA Diameter |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | MPI (Tei) | none | calc | c_ob_TeiIndex | none | unilateral | 125000\59099-2\99999 | 99PMSBLUS, 99999, Fetal Heart | LN, 59099-2, Myocardial Performance Index (Tei) |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | LV MPI | none | calc | c_ob_LV_MPI | none | unilateral | 125000\99999\G-037F | 99PMSBLUS, 99999, Fetal Heart | SRT, G-037F, Left Ventricular Index of Myocardial Performance |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | RV MPI | none | calc | c_ob_RV_MPI | none | unilateral | 125000\99999\G-0381 | 99PMSBLUS, 99999, Fetal Heart | SRT, G-0381, Right Ventricular Index of Myocardial Performance |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | CI | none | calc | c_ci | none | unilateral | 125000\11823-2\125001 | DCM, 125001, Fetal Biometry Ratios | LN, 11823-2, Cephalic Index |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | FL/AC | none | calc | c_flac | none | unilateral | 125000\11871-1\125001 | DCM, 125001, Fetal Biometry Ratios | LN, 11871-1, FL/AC |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | FL/BPD | none | calc | c_flbpd | none | unilateral | 125000\11872-9\125001 | DCM, 125001, Fetal Biometry Ratios | LN, 11872-9, FL/BPD |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | HC/AC | none | calc | c_hcac | none | unilateral | 125000\11947-9\125001 | DCM, 125001, Fetal Biometry Ratios | LN, 11947-9, HC/AC |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | HrtC/ThrC | none | calc | c_ob_hrtThrcirc | none | unilateral | 125000\125002\C12004-01 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12004-01, HrtC/TC (Heart Circumference/Thoracic Circumference) |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | APTDxTTD | none | calc | c_aptd_ttd | none | unilateral | 125000\125002\C12005-10 | DCM, 125002, Fetal Biometry | 99PMSBLUS, C12005-10, APTDxTTD |  |  | UCUM, mm2, mm2
OB-GYN Ultrasound Procedure Report | OB | LA/Ao | none | calc | c_ob_laAo2D | none | unilateral | 125000\125002\P5000-01-12 | DCM, 125002, Fetal Biometry | 99PMSBLUS, P5000-01-12, Ratio of LA Dimension to Aortic Root Diameter |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | BPDa | none | calc | c_bpda | none | unilateral | 125000\11824-0\125002 | DCM, 125002, Fetal Biometry | LN, 11824-0, BPD area corrected |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | OB | AC(c) | none | calc | c_ac | none | unilateral | 125000\11979-2\125002\R-41D2D | DCM, 125002, Fetal Biometry | LN, 11979-2, Abdominal Circumference |  |  | UCUM, mm, mm | [9 empty] | SRT, R-41D2D, Calculated
OB-GYN Ultrasound Procedure Report | OB | HC(c) | none | calc | c_hc | none | unilateral | 125000\11984-2\125002\R-41D2D | DCM, 125002, Fetal Biometry | LN, 11984-2, Head Circumference |  |  | UCUM, mm, mm | [9 empty] | SRT, R-41D2D, Calculated
OB-GYN Ultrasound Procedure Report | OB | ThC(c) | none | calc | c_tc | none | unilateral | 125000\11988-3\125002\R-41D2D | DCM, 125002, Fetal Biometry | LN, 11988-3, Thoracic Circumference |  |  | UCUM, mm, mm | [9 empty] | SRT, R-41D2D, Calculated
OB-GYN Ultrasound Procedure Report | OB | FTA(c) | none | calc | c_fta | none | unilateral | 125000\125002\33068-8\R-41D2D | DCM, 125002, Fetal Biometry | LN, 33068-8, Thoracic Area |  |  | UCUM, mm2, mm2 | [9 empty] | SRT, R-41D2D, Calculated
OB-GYN Ultrasound Procedure Report | OB | BPP Score | none | calc | c_bppscore | none | unilateral | 125000\11634-3\125006 | DCM, 125006, Biophysical Profile | LN, 11634-3, Biophysical Profile Sum Score |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | GA(Con) | none | calc | c_gacon | none | unilateral | 125000\125008\C12019-02 | DCM, 125008, Fetus Summary | 99PMSBLUS, C12019-02, Gestational Age by conception date |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | GA(EDD) | none | calc | c_gaedd | none | unilateral | 125000\125008\C12019-03 | DCM, 125008, Fetus Summary | 99PMSBLUS, C12019-03, Gestational Age by EDD |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | GA(LMP) | none | calc | c_galmp | none | unilateral | 125000\11885-1\125008 | DCM, 125008, Fetus Summary | LN, 11885-1, Gestational Age by LMP |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | AUA | none | calc | c_aua | none | unilateral | 125000\11888-5\125008 | DCM, 125008, Fetus Summary | LN, 11888-5, Composite Ultrasound Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Mean Sac Diam | none | calc | c_msd | none | unilateral | 125000\125009\C12009-04 | DCM, 125009, Early Gestation | 99PMSBLUS, C12009-04, Mean Sac Diameter |  |  | UCUM, mm, mm
OB-GYN Ultrasound Procedure Report | GYN | Bladder Vol | none | calc | c_gyn_blad_vol | none | unilateral | 125000\125011\C12011-04 | DCM, 125011, Pelvis and Uterus | 99PMSBLUS, C12011-04, Bladder Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | PV Bladder Vol | none | calc | c_gyn_pvblad_vol | none | unilateral | 125000\125011\C12011-08 | DCM, 125011, Pelvis and Uterus | 99PMSBLUS, C12011-08, Post Void Bladder Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | GYN | Uterus Volume | none | calc | c_uterusvol | none | unilateral | 125000\125011\33192-6 | DCM, 125011, Pelvis and Uterus | LN, 33192-6, Uterus Volume |  |  | UCUM, mm3, mm3
OB-GYN Ultrasound Procedure Report | OB | Alexander | none |  | efwAlexander_pc | none | unilateral | 125000\11767-1\121424\C12016-09 |  | LN, 11767-1, EFW percentile rank |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Brenner | none |  | efwAuaBrenner_pc | none | unilateral | 125000\11767-1\121424\33189-2 |  | LN, 11767-1, EFW percentile rank |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Doubilet | none |  | efwAuaDoubilet_pc | none | unilateral | 125000\11767-1\121424\C12016-10 |  | LN, 11767-1, EFW percentile rank |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Hadlock | none |  | efwAuaHad_pc | none | unilateral | 125000\11767-1\121424\33183-5 |  | LN, 11767-1, EFW percentile rank |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Philips Custom | none |  | efwAtl_pc | none | unilateral | 125000\11767-1\121424\C12016-04 |  | LN, 11767-1, EFW percentile rank |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Shinozuka | none |  | efwAuaShino_pc | none | unilateral | 125000\11767-1\121424\C12016-12 |  | LN, 11767-1, EFW percentile rank |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Williams (Female) | none |  | efwAuaWilliamsFem_pc | none | unilateral | 125000\11767-1\121424\C12016-06 |  | LN, 11767-1, EFW percentile rank |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Williams (Male) | none |  | efwAuaWilliamsMale_pc | none | unilateral | 125000\11767-1\121424\C12016-05 |  | LN, 11767-1, EFW percentile rank |  |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | ASUM (2001) | none |  | acAsum_ga | none | unilateral | 125000\121424\18185-9\C12013-01 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | BMUS | none |  | crlBMUS_ga | none | unilateral | 125000\121424\18185-9\C12013-42 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Chitty | none |  | cerebChitty_ga | none | unilateral | 125000\121420\18185-9\C12013-15 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Chitty (Derived) | none |  | hcChittyDer_ga | none | unilateral | 125000\121420\18185-9\33111-6 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Chitty (Measured) | none |  | hcChittyMeas_ga | none | unilateral | 125000\121420\18185-9\33110-8 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Chitty (Outer Inner) | none |  | bpdaChittyOI_ga | none | unilateral | 125000\121420\18185-9\C12013-03 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Chitty (Outer Outer) | none |  | bpdaChittyOO_ga | none | unilateral | 125000\121420\18185-9\C12013-04 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Hadlock | none |  | acHad_ga | none | unilateral | 125000\11892-7\121420\18185-9 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Hansmann | none |  | acHans_ga | none | unilateral | 125000\121424\18185-9\C12013-13 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Hellman | none |  | msdHellman_ga | none | unilateral | 125000\121420\18185-9\C12013-10 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Hill | none |  | cerebHill_ga | none | unilateral | 125000\121420\18185-9\33134-8 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | INTERGROWTH-21?? | none |  | crlInterGrowth21_ga | none | unilateral | 125000\121420\18185-9\C12013-47 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Jeanty | none |  | bpdJeanty_ga | none | unilateral | 125000\11905-7\121424\18185-9 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | JSUM (2001) | none |  | crlJSUM2001_ga | none | unilateral | 125000\121424\18185-9\C12013-22 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Nimrod | none |  | tcNimrod_ga | none | unilateral | 125000\121420\18185-9\C12013-23 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Nyberg | none |  | gestsacNyb_ga | none | unilateral | 125000\121420\18185-9\33107-4 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Osaka | none |  | bpdOsaka_ga | none | unilateral | 125000\121424\18185-9\C12013-31 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Robinson | none |  | crlRob_ga | none | unilateral | 125000\11914-9\121420\18185-9 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Shinozuka | none |  | acShino_ga | none | unilateral | 125000\121424\18185-9\33076-1 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Tokyo | none |  | slTokyo_ga | none | unilateral | 125000\121424\18185-9\33127-2 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Verburg | none |  | acVerb_ga | none | unilateral | 125000\121424\18185-9\C12013-43 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Yarkoni | none |  | clavYark_ga | none | unilateral | 125000\121424\18185-9\33088-6 |  | LN, 18185-9, Gestational Age |  |  | UCUM, d, days
OB-GYN Ultrasound Procedure Report | OB | Campbell | none |  | efwCamp_fw | none | unilateral | 125000\11727-5\11756-4\121420 |  | LN, 11727-5, Estimated Weight |  |  | UCUM, g, grams
OB-GYN Ultrasound Procedure Report | OB | Hadlock (AC | none |  | efwHad1_fw | none | unilateral | 125000\11727-5\11751-5\121420 |  | LN, 11727-5, Estimated Weight |  |  | UCUM, g, grams
OB-GYN Ultrasound Procedure Report | OB | Osaka(BPD | none |  | efwOsaka3_fw | none | unilateral | 125000\11727-5\121420\33140-5 |  | LN, 11727-5, Estimated Weight |  |  | UCUM, g, grams
OB-GYN Ultrasound Procedure Report | OB | Shino(APTDxTTD | none |  | efwShino4_fw | none | unilateral | 125000\11727-5\121420\33143-9 |  | LN, 11727-5, Estimated Weight |  |  | UCUM, g, grams
OB-GYN Ultrasound Procedure Report | OB | Shinozuka(AC | none |  | efwShino3_fw | none | unilateral | 125000\11727-5\121420\C12014-01 |  | LN, 11727-5, Estimated Weight |  |  | UCUM, g, grams
OB-GYN Ultrasound Procedure Report | OB | Tokyo(BPD | none |  | efwTokyo1986_fw | none | unilateral | 125000\11727-5\121420\C12014-02 |  | LN, 11727-5, Estimated Weight |  |  | UCUM, g, grams
Patient Characteristics | patientcharacteristics | Subject Name | none |  | DMM_EXAM_PATIENT_NAME | none | n/a | 121029 |  |  |  | DCM, 121029, Subject Name
Patient Characteristics | patientcharacteristics | Subject ID | none |  | DMM_EXAM_PATIENT_ID | none | n/a | 121030 |  |  |  | DCM, 121030, Subject ID
Patient Characteristics | patientcharacteristics | Subject Sex | none |  | DMM_EXAM_PATIENT_SEX | none | n/a | 121032 |  |  |  | DCM, 121032, Subject Sex
Patient Characteristics | patientcharacteristics | Subject Birth Date | none |  | DMM_EXAM_PATIENT_DOB | none | n/a | 121031 |  |  |  | DCM, 121031, Subject Birth Date
Patient Characteristics | patientcharacteristics | Comment | none |  | DMM_EXAM_PATIENT_COMMENTS | none | n/a | 121106 |  |  |  | DCM, 121106, Comment
Patient Characteristics | patientcharacteristics | Patient Height | none |  | DMM_STUDY_HEIGHT | none | n/a | 2338315 |  |  |  | LN, 8302-2, Patient Height
Patient Characteristics | patientcharacteristics | Patient Weight | none |  | DMM_STUDY_WEIGHT | none | n/a | 29463-7 |  |  |  | LN, 29463-7, Patient Weight
Patient Characteristics | patientcharacteristics | Accession Number | none |  | DMM_EXAM_PATIENT_ACCESSION_NUMBER | none | n/a | 121022 |  |  |  | DCM, 121022, Accession Number
Patient Characteristics | patientcharacteristics | Sonographer | none |  | DMM_STUDY_SONOGRAPHER | none | n/a | 121093 |  |  |  | DCM, 121093, Sonographer
Patient Characteristics | patientcharacteristics | Study Description | none |  | DMM_EXAM_STUDY_DESCRIPTION | none | n/a | T9910-07 |  |  |  | 99PMSBLUS, T9910-07, Study Description
Patient Characteristics | patientcharacteristics | Referring Physician | none |  | DMM_EXAM_REFERRING_PHYSICIAN_NAME | none | n/a | T9910-08 |  |  |  | 99PMSBLUS, T9910-08, Referring Physician
Patient Characteristics | patientcharacteristics | Exam date | none |  | DICOM_PERFORMED_PROCEDURE_STEP_START_DATE | none | n/a | T9910-09 |  |  |  | 99PMSBLUS, T9910-09, Exam date
Patient Characteristics | patientcharacteristics | Study date | none |  | DMM_STUDY_DATE | none | n/a | T9910-105 |  |  |  | 99PMSBLUS, T9910-105, Study date
Patient Characteristics | patientcharacteristics | Diabetic | none |  | DMM_STUDY_OB_DIABETIC | none | n/a | T9910-103 |  |  |  | 99PMSBLUS, T9910-103, Diabetic
Patient Characteristics | patientcharacteristics | Diabetic Type | none |  | DMM_STUDY_OB_TYPE | none | n/a | T9910-104 |  |  |  | 99PMSBLUS, T9910-104, Diabetic Type
Patient Characteristics | patientcharacteristics | Gravida | none |  | DMM_STUDY_GYN_GRAVIDA | none | n/a | 11996-6 |  |  |  | LN, 11996-6, Gravida
Patient Characteristics | patientcharacteristics | Para | none |  | DMM_STUDY_GYN_PARA | none | n/a | 11977-6 |  |  |  | LN, 11977-6, Para
Patient Characteristics | patientcharacteristics | Aborta | none |  | DMM_STUDY_GYN_ABORTA | none | n/a | 11612-9 |  |  |  | LN, 11612-9, Aborta
Patient Characteristics | patientcharacteristics | Ectopic Pregnancies | none |  | DMM_STUDY_GYN_ECTOPIC | none | n/a | 33065-4 |  |  |  | LN, 33065-4, Ectopic Pregnancies
Patient Characteristics | patientcharacteristics | Expected Ovulation Date | none |  | DMM_STUDY_GYN_EXPECTED_OVULATION | none | n/a | T9910-10 |  |  |  | 99PMSBLUS, T9910-10, Expected Ovulation Date
Patient Characteristics | patientcharacteristics | Abnormal Cycles | none |  | DMM_STUDY_GYN_ABNORMAL_CYCLES | none | n/a | T9910-11 |  |  |  | 99PMSBLUS, T9910-11, Abnormal Cycles
Patient Characteristics | patientcharacteristics | Birth Control Medication | none |  | DMM_STUDY_GYN_BIRTH_CONTROL_PILLS | none | n/a | T9910-12 |  |  |  | 99PMSBLUS, T9910-12, Birth Control Medication
Patient Characteristics | patientcharacteristics | Birth Control Medication Duration | none |  | DMM_STUDY_GYN_BIRTH_CONTROL_MEDS_DURATION | none | n/a | T9910-13 |  |  |  | 99PMSBLUS, T9910-13, Birth Control Medication Duration
Patient Characteristics | patientcharacteristics | Hormone Replacement Therapy | none |  | DMM_STUDY_BREAST_HOMRMONE_REPLACEMENT | none | n/a | T9910-100 |  |  |  | 99PMSBLUS, T9910-100, Hormone Replacement Therapy
Patient Characteristics | patientcharacteristics | Hormone Replacement Therapy Year Started | none |  | DMM_STUDY_BREAST_YEAR_STARTED | none | n/a | T9910-101 |  |  |  | 99PMSBLUS, T9910-101, Hormone Replacement Therapy Year Started
Patient Characteristics | patientcharacteristics | Menopause | none |  | DMM_STUDY_BREAST_MENOPAUSE | none | n/a | T9910-102 |  |  |  | 99PMSBLUS, T9910-102, Menopause
Patient Characteristics | patientcharacteristics | Pelvic Pain Right | none |  | DMM_STUDY_GYN_PELVIC_PAIN_RIGHT | none | n/a | T9910-14 |  |  |  | 99PMSBLUS, T9910-14, Pelvic Pain Right
Patient Characteristics | patientcharacteristics | Pelvic Pain Left | none |  | DMM_STUDY_GYN_PELVIC_PAIN_LEFT | none | n/a | T9910-15 |  |  |  | 99PMSBLUS, T9910-15, Pelvic Pain Left
Patient Characteristics | patientcharacteristics | Bleeding | none |  | DMM_STUDY_GYN_BLEEDING | none | n/a | T9910-16 |  |  |  | 99PMSBLUS, T9910-16, Bleeding
Patient Characteristics | patientcharacteristics | Bleeding Duration | none |  | DMM_STUDY_GYN_BLEEDING_DURATION | none | n/a | T9910-17 |  |  |  | 99PMSBLUS, T9910-17, Bleeding Duration
Patient Characteristics | patientcharacteristics | Complete Hysterectomy | none |  | DMM_STUDY_GYN_COMPLETE | none | n/a | T9910-18 |  |  |  | 99PMSBLUS, T9910-18, Complete Hysterectomy
Patient Characteristics | patientcharacteristics | Partial Hysterectomy | none |  | DMM_STUDY_GYN_PARTIAL | none | n/a | T9910-19 |  |  |  | 99PMSBLUS, T9910-19, Partial Hysterectomy
Patient Characteristics | patientcharacteristics | Ovary Surgery Right | none |  | DMM_STUDY_GYN_RT | none | n/a | T9910-20 |  |  |  | 99PMSBLUS, T9910-20, Ovary Surgery Right
Patient Characteristics | patientcharacteristics | Ovary Surgery Left | none |  | DMM_STUDY_GYN_LT | none | n/a | T9910-21 |  |  |  | 99PMSBLUS, T9910-21, Ovary Surgery Left
Patient Characteristics | patientcharacteristics | Ovary Surgery Bilateral | none |  | DMM_STUDY_GYN_BILATERAL | none | n/a | T9910-22 |  |  |  | 99PMSBLUS, T9910-22, Ovary Surgery Bilateral
Patient Characteristics | patientcharacteristics | Endometrium | none |  | DMM_STUDY_GYN_ENDOMETRIUM | none | n/a | T9910-23 |  |  |  | 99PMSBLUS, T9910-23, Endometrium
Patient Characteristics | patientcharacteristics | LMP | none |  | DMM_STUDY_OB_LMP | none | n/a | 11955-2 |  |  |  | LN, 11955-2, LMP
Patient Characteristics | patientcharacteristics | Prior right breast surgery | none |  | DMM_STUDY_BREAST_PRIOR_BREAST_SURGERY_RIGHT | none | n/a | T9910-188 |  |  |  | 99PMSBLUS, T9910-188, Prior right breast surgery
Patient Characteristics | patientcharacteristics | Prior right breast biopsy | none |  | DMM_STUDY_BREAST_PRIOR_BREAST_BIOPSY_RIGHT | none | n/a | T9910-189 |  |  |  | 99PMSBLUS, T9910-189, Prior right breast biopsy
Patient Characteristics | patientcharacteristics | Prior left breast surgery | none |  | DMM_STUDY_BREAST_PRIOR_BREAST_SURGERY_LEFT | none | n/a | T9910-190 |  |  |  | 99PMSBLUS, T9910-190, Prior left breast surgery
Patient Characteristics | patientcharacteristics | Prior left breast biopsy | none |  | DMM_STUDY_BREAST_PRIOR_BREAST_BIOPSY_LEFT | none | n/a | T9910-191 |  |  |  | 99PMSBLUS, T9910-191, Prior left breast biopsy
Patient Characteristics | patientcharacteristics | History of uterus cancer | none |  | DMM_STUDY_BREAST_UTERUS | none | n/a | T9910-192 |  |  |  | 99PMSBLUS, T9910-192, History of uterus cancer
Patient Characteristics | patientcharacteristics | History of ovaries cancer | none |  | DMM_STUDY_BREAST_OVARIES | none | n/a | T9910-193 |  |  |  | 99PMSBLUS, T9910-193, History of ovaries cancer
Patient Characteristics | patientcharacteristics | History of colon cancer | none |  | DMM_STUDY_BREAST_COLON | none | n/a | T9910-194 |  |  |  | 99PMSBLUS, T9910-194, History of colon cancer
Patient Characteristics | patientcharacteristics | Mother had breast cancer | none |  | DMM_STUDY_BREAST_MOTHER | none | n/a | T9910-195 |  |  |  | 99PMSBLUS, T9910-195, Mother had breast cancer
Patient Characteristics | patientcharacteristics | Daughter had breast cancer | none |  | DMM_STUDY_BREAST_DAUGHTER | none | n/a | T9910-196 |  |  |  | 99PMSBLUS, T9910-196, Daughter had breast cancer
Patient Characteristics | patientcharacteristics | Sister had breast cancer | none |  | DMM_STUDY_BREAST_SISTER | none | n/a | T9910-197 |  |  |  | 99PMSBLUS, T9910-197, Sister had breast cancer
Patient Characteristics | patientcharacteristics | Grandmother had breast cancer | none |  | DMM_STUDY_BREAST_GRANDMOTHER | none | n/a | T9910-198 |  |  |  | 99PMSBLUS, T9910-198, Grandmother had breast cancer
Patient Characteristics | patientcharacteristics | Aunt had breast cancer | none |  | DMM_STUDY_BREAST_AUNT | none | n/a | T9910-199 |  |  |  | 99PMSBLUS, T9910-199, Aunt had breast cancer
Patient Characteristics | patientcharacteristics | The age of Mother at diagnosis of cancer | none |  | DMM_STUDY_BREAST_MOTHER_AGE | none | n/a | T9910-200 |  |  |  | 99PMSBLUS, T9910-200, The age of Mother at diagnosis of cancer
Patient Characteristics | patientcharacteristics | The age of Daughter at diagnosis of cancer | none |  | DMM_STUDY_BREAST_DAUGHTER_AGE | none | n/a | T9910-201 |  |  |  | 99PMSBLUS, T9910-201, The age of Daughter at diagnosis of cancer
Patient Characteristics | patientcharacteristics | The age of Sister at diagnosis of cancer | none |  | DMM_STUDY_BREAST_SISTER_AGE | none | n/a | T9910-202 |  |  |  | 99PMSBLUS, T9910-202, The age of Sister at diagnosis of cancer
Patient Characteristics | patientcharacteristics | The age of Grandmother at diagnosis of cancer | none |  | DMM_STUDY_BREAST_GRANDMOTHER_AGE | none | n/a | T9910-203 |  |  |  | 99PMSBLUS, T9910-203, The age of Grandmother at diagnosis of cancer
Patient Characteristics | patientcharacteristics | The age of Aunt at diagnosis of cancer | none |  | DMM_STUDY_BREAST_AUNT_AGE | none | n/a | T9910-204 |  |  |  | 99PMSBLUS, T9910-204, The age of Aunt at diagnosis of cancer
Patient Characteristics | patientcharacteristics | Right palpable mass | none |  | DMM_STUDY_BREAST_PALPABLE_MASS_RT | none | n/a | T9910-205 |  |  |  | 99PMSBLUS, T9910-205, Right palpable mass
Patient Characteristics | patientcharacteristics | Left palpable mass | none |  | DMM_STUDY_BREAST_PALPABLE_MASS_LT | none | n/a | T9910-206 |  |  |  | 99PMSBLUS, T9910-206, Left palpable mass
Patient Characteristics | patientcharacteristics | Right nipple discharge | none |  | DMM_STUDY_BREAST_NIPPLE_DISCHARGE_RT | none | n/a | T9910-207 |  |  |  | 99PMSBLUS, T9910-207, Right nipple discharge
Patient Characteristics | patientcharacteristics | Left nipple discharge | none |  | DMM_STUDY_BREAST_NIPPLE_DISCHARGE_LT | none | n/a | T9910-208 |  |  |  | 99PMSBLUS, T9910-208, Left nipple discharge
Patient Characteristics | patientcharacteristics | Breast Pain Right | none |  | DMM_STUDY_BREAST_PAIN_RT | none | n/a | T9910-209 |  |  |  | 99PMSBLUS, T9910-209, Breast Pain Right
Patient Characteristics | patientcharacteristics | Breast Pain Left | none |  | DMM_STUDY_BREAST_PAIN_LT | none | n/a | T9910-210 |  |  |  | 99PMSBLUS, T9910-210, Breast Pain Left
Patient Characteristics | patientcharacteristics | Mammogram Findings Right | none |  | DMM_STUDY_BREAST_MAMMO_FINDING_RT | none | n/a | T9910-211 |  |  |  | 99PMSBLUS, T9910-211, Mammogram Findings Right
Patient Characteristics | patientcharacteristics | Mammogram Findings Left | none |  | DMM_STUDY_BREAST_MAMMO_FINDINGS_LT | none | n/a | T9910-212 |  |  |  | 99PMSBLUS, T9910-212, Mammogram Findings Left
Patient Characteristics | patientcharacteristics | Mammogram Location of Findings | none |  | DMM_STUDY_BREAST_MAMMO_LOCATION | none | n/a | T9910-213 |  |  |  | 99PMSBLUS, T9910-213, Mammogram Location of Findings
Patient Characteristics | patientcharacteristics | Mammogram Date | none |  | DMM_STUDY_BREAST_MAMMOGRAM_DATE | none | n/a | T9910-214 |  |  |  | 99PMSBLUS, T9910-214, Mammogram Date
Patient Characteristics | patientcharacteristics | History of breast cancer | none |  | DMM_STUDY_BREAST_CANCER | none | n/a | T9910-215 |  |  |  | 99PMSBLUS, T9910-215, History of breast cancer
Patient Characteristics | patientcharacteristics | Breast feeding history | none |  | DMM_STUDY_BREAST_BREAST_FEEDING | none | n/a | T9910-216 |  |  |  | 99PMSBLUS, T9910-216, Breast feeding history
Patient Characteristics | patientcharacteristics | Prostate specific antigen | none |  | DMM_STUDY_UROLOGY_PSA | none | n/a | T9910-217 |  |  |  | 99PMSBLUS, T9910-217, Prostate specific antigen
Patient Characteristics | patientcharacteristics | Smoker | none |  | DMM_STUDY_CARDIAC_SMOKER | none | n/a | T9910-74 |  |  |  | 99PMSBLUS, T9910-74, Smoker
Patient Characteristics | patientcharacteristics | Hypertension | none |  | DMM_STUDY_CARDIAC_HYPERTENSION | none | n/a | T9910-75 |  |  |  | 99PMSBLUS, T9910-75, Hypertension
Patient Characteristics | patientcharacteristics | Recent Injury | none |  | DMM_STUDY_VASCULAR_RECENT_INJURY | none | n/a | T9910-24 |  |  |  | 99PMSBLUS, T9910-24, Recent Injury
Patient Characteristics | patientcharacteristics | Surgeries | none |  | DMM_STUDY_VASCULAR_SURGERIES | none | n/a | T9910-25 |  |  |  | 99PMSBLUS, T9910-25, Surgeries
Patient Characteristics | patientcharacteristics | Aphasia | none |  | DMM_STUDY_VASCULAR_APHASIA | none | n/a | T9910-26 |  |  |  | 99PMSBLUS, T9910-26, Aphasia
Patient Characteristics | patientcharacteristics | Double Vision | none |  | DMM_STUDY_VASCULAR_DOUBLE_VISION | none | n/a | T9910-27 |  |  |  | 99PMSBLUS, T9910-27, Double Vision
Patient Characteristics | patientcharacteristics | Memory Loss | none |  | DMM_STUDY_VASCULAR_MEMORY_LOSS | none | n/a | T9910-28 |  |  |  | 99PMSBLUS, T9910-28, Memory Loss
Patient Characteristics | patientcharacteristics | Syncope | none |  | DMM_STUDY_VASCULAR_SYNCOPE | none | n/a | T9910-29 |  |  |  | 99PMSBLUS, T9910-29, Syncope
Patient Characteristics | patientcharacteristics | Confusion | none |  | DMM_STUDY_VASCULAR_CONFUSION | none | n/a | T9910-30 |  |  |  | 99PMSBLUS, T9910-30, Confusion
Patient Characteristics | patientcharacteristics | Bruit | none |  | DMM_STUDY_VASCULAR_BRUIT | none | n/a | T9910-31 |  |  |  | 99PMSBLUS, T9910-31, Bruit
Patient Characteristics | patientcharacteristics | Stroke | none |  | DMM_STUDY_VASCULAR_STROKE | none | n/a | T9910-32 |  |  |  | 99PMSBLUS, T9910-32, Stroke
Patient Characteristics | patientcharacteristics | Stroke Date | none |  | DMM_STUDY_VASCULAR_STROKE_DATE | none | n/a | T9910-33 |  |  |  | 99PMSBLUS, T9910-33, Stroke Date
Patient Characteristics | patientcharacteristics | Endarterectomy Right | none |  | DMM_STUDY_VASCULAR_CAROTID_ENDART_RT | none | n/a | T9910-34 |  |  |  | 99PMSBLUS, T9910-34, Endarterectomy Right
Patient Characteristics | patientcharacteristics | Endarterectomy Left | none |  | DMM_STUDY_VASCULAR_CAROTID_ENDART_LT | none | n/a | T9910-35 |  |  |  | 99PMSBLUS, T9910-35, Endarterectomy Left
Patient Characteristics | patientcharacteristics | Endarterectomy Date | none |  | DMM_STUDY_VASCULAR_CAROTID_ENDART | none | n/a | T9910-36 |  |  |  | 99PMSBLUS, T9910-36, Endarterectomy Date
Patient Characteristics | patientcharacteristics | Hemiparesis Right | none |  | DMM_STUDY_VASCULAR_HEMIPARESIS_RT | none | n/a | T9910-37 |  |  |  | 99PMSBLUS, T9910-37, Hemiparesis Right
Patient Characteristics | patientcharacteristics | Hemiparesis Left | none |  | DMM_STUDY_VASCULAR_HEMIPARESIS_LT | none | n/a | T9910-38 |  |  |  | 99PMSBLUS, T9910-38, Hemiparesis Left
Patient Characteristics | patientcharacteristics | Weakness Right | none |  | DMM_STUDY_VASCULAR_WEAKNESS_RT | none | n/a | T9910-39 |  |  |  | 99PMSBLUS, T9910-39, Weakness Right
Patient Characteristics | patientcharacteristics | Weakness Left | none |  | DMM_STUDY_VASCULAR_WEAKNESS_LT | none | n/a | T9910-40 |  |  |  | 99PMSBLUS, T9910-40, Weakness Left
Patient Characteristics | patientcharacteristics | Amaurosis Fugax Right | none |  | DMM_STUDY_VASCULAR_FUGAX_RT | none | n/a | T9910-41 |  |  |  | 99PMSBLUS, T9910-41, Amaurosis Fugax Right
Patient Characteristics | patientcharacteristics | Amaurosis Fugax Left | none |  | DMM_STUDY_VASCULAR_FUGAX_LT | none | n/a | T9910-42 |  |  |  | 99PMSBLUS, T9910-42, Amaurosis Fugax Left
Patient Characteristics | patientcharacteristics | Obesity | none |  | DMM_STUDY_VASCULAR_OBESITY | none | n/a | T9910-43 |  |  |  | 99PMSBLUS, T9910-43, Obesity
Patient Characteristics | patientcharacteristics | Edema | none |  | DMM_STUDY_VASCULAR_EDEMA | none | n/a | T9910-44 |  |  |  | 99PMSBLUS, T9910-44, Edema
Patient Characteristics | patientcharacteristics | Previous DVT Right | none |  | DMM_STUDY_VASCULAR_DVT_RT | none | n/a | T9910-45 |  |  |  | 99PMSBLUS, T9910-45, Previous DVT Right
Patient Characteristics | patientcharacteristics | Previous DVT Left | none |  | DMM_STUDY_VASCULAR_DVT_LT | none | n/a | T9910-46 |  |  |  | 99PMSBLUS, T9910-46, Previous DVT Left
Patient Characteristics | patientcharacteristics | History of Pulmonary Thrombus | none |  | DMM_STUDY_VASCULAR_HX_OF_PULM | none | n/a | T9910-47 |  |  |  | 99PMSBLUS, T9910-47, History of Pulmonary Thrombus
Patient Characteristics | patientcharacteristics | History of Malignancy | none |  | DMM_STUDY_VASCULAR_HX_OF_MAL | none | n/a | T9910-48 |  |  |  | 99PMSBLUS, T9910-48, History of Malignancy
Patient Characteristics | patientcharacteristics | Pregnant | none |  | DMM_STUDY_VASCULAR_PREGNANT | none | n/a | T9910-49 |  |  |  | 99PMSBLUS, T9910-49, Pregnant
Patient Characteristics | patientcharacteristics | Claudication | none |  | DMM_STUDY_VASCULAR_CLAUDICATION | none | n/a | T9910-50 |  |  |  | 99PMSBLUS, T9910-50, Claudication
Patient Characteristics | patientcharacteristics | RUQ Pain | none |  | DMM_STUDY_ABDOMEN_RUQ | none | n/a | T9910-51 |  |  |  | 99PMSBLUS, T9910-51, RUQ Pain
Patient Characteristics | patientcharacteristics | LUQ Pain | none |  | DMM_STUDY_ABDOMEN_LUQ | none | n/a | T9910-52 |  |  |  | 99PMSBLUS, T9910-52, LUQ Pain
Patient Characteristics | patientcharacteristics | Midline Pain | none |  | DMM_STUDY_ABDOMEN_MIDLINE | none | n/a | T9910-53 |  |  |  | 99PMSBLUS, T9910-53, Midline Pain
Patient Characteristics | patientcharacteristics | RLQ Pain | none |  | DMM_STUDY_ABDOMEN_RLQ | none | n/a | T9910-54 |  |  |  | 99PMSBLUS, T9910-54, RLQ Pain
Patient Characteristics | patientcharacteristics | LLQ Pain | none |  | DMM_STUDY_ABDOMEN_LLQ | none | n/a | T9910-55 |  |  |  | 99PMSBLUS, T9910-55, LLQ Pain
Patient Characteristics | patientcharacteristics | Periumbilical Pain | none |  | DMM_STUDY_ABDOMEN_PERIUMBILICAL | none | n/a | T9910-56 |  |  |  | 99PMSBLUS, T9910-56, Periumbilical Pain
Patient Characteristics | patientcharacteristics | Nausea | none |  | DMM_STUDY_ABDOMEN_NAUSEA | none | n/a | T9910-57 |  |  |  | 99PMSBLUS, T9910-57, Nausea
Patient Characteristics | patientcharacteristics | Nausea Duration | none |  | DMM_STUDY_ABDOMEN_NAUSA_DURATION | none | n/a | T9910-58 |  |  |  | 99PMSBLUS, T9910-58, Nausea Duration
Patient Characteristics | patientcharacteristics | Vomiting | none |  | DMM_STUDY_ABDOMEN_VOMITING | none | n/a | T9910-59 |  |  |  | 99PMSBLUS, T9910-59, Vomiting
Patient Characteristics | patientcharacteristics | Vomiting Duration | none |  | DMM_STUDY_ABDOMEN_VOMITING_DURATION | none | n/a | T9910-60 |  |  |  | 99PMSBLUS, T9910-60, Vomiting Duration
Patient Characteristics | patientcharacteristics | Diarrhea | none |  | DMM_STUDY_ABDOMEN_DIARRHEA | none | n/a | T9910-61 |  |  |  | 99PMSBLUS, T9910-61, Diarrhea
Patient Characteristics | patientcharacteristics | Diarrhea Duration | none |  | DMM_STUDY_ABDOMEN_DIARRHEA_DURATION | none | n/a | T9910-62 |  |  |  | 99PMSBLUS, T9910-62, Diarrhea Duration
Patient Characteristics | patientcharacteristics | Weight Loss | none |  | DMM_STUDY_ABDOMEN_WEIGHT_LOSS | none | n/a | T9910-63 |  |  |  | 99PMSBLUS, T9910-63, Weight Loss
Patient Characteristics | patientcharacteristics | Weight Loss Duration | none |  | DMM_STUDY_ABDOMEN_WEIGHT_LOSS_DURATION | none | n/a | T9910-64 |  |  |  | 99PMSBLUS, T9910-64, Weight Loss Duration
Patient Characteristics | patientcharacteristics | Abnormal Lab Values | none |  | DMM_STUDY_ABDOMEN_ABNORMAL_LAB_RESULTS | none | n/a | T9910-65 |  |  |  | 99PMSBLUS, T9910-65, Abnormal Lab Values
Patient Characteristics | patientcharacteristics | History of Aortic Aneurysm | none |  | DMM_STUDY_ABDOMEN_HX_OF_AORTIC_ANEURYSM | none | n/a | T9910-66 |  |  |  | 99PMSBLUS, T9910-66, History of Aortic Aneurysm
Patient Characteristics | patientcharacteristics | Previous Measurement | none |  | DMM_STUDY_ABDOMEN_PREVIOUS_MEASUREMENT | none | n/a | T9910-67 |  |  |  | 99PMSBLUS, T9910-67, Previous Measurement
Patient Characteristics | patientcharacteristics | Aortic Aneurysm Date | none |  | DMM_STUDY_ABDOMEN_AORTIC_DATE | none | n/a | T9910-68 |  |  |  | 99PMSBLUS, T9910-68, Aortic Aneurysm Date
Patient Characteristics | patientcharacteristics | Cholesystectomy | none |  | DMM_STUDY_ABDOMEN_CHOLECYSTOSTOMY | none | n/a | T9910-69 |  |  |  | 99PMSBLUS, T9910-69, Cholesystectomy
Patient Characteristics | patientcharacteristics | Cholesystectomy Date | none |  | DMM_STUDY_ABDOMEN_CHOLE_DATE | none | n/a | T9910-70 |  |  |  | 99PMSBLUS, T9910-70, Cholesystectomy Date
Patient Characteristics | patientcharacteristics | TIPSS | none |  | DMM_STUDY_ABDOMEN_TIPSS | none | n/a | T9910-71 |  |  |  | 99PMSBLUS, T9910-71, TIPSS
Patient Characteristics | patientcharacteristics | TIPSS Date | none |  | DMM_STUDY_ABDOMEN_TIPSS_DATE | none | n/a | T9910-72 |  |  |  | 99PMSBLUS, T9910-72, TIPSS Date
Patient Characteristics | patientcharacteristics | Other Surgeries | none |  | DMM_STUDY_ABDOMEN_OTHER_SURGERIES | none | n/a | T9910-73 |  |  |  | 99PMSBLUS, T9910-73, Other Surgeries
Patient Characteristics | patientcharacteristics | Stroke | none |  | DMM_STUDY_TCD_STROKE | none | n/a | T9910-106 |  |  |  | 99PMSBLUS, T9910-106, Stroke
Patient Characteristics | patientcharacteristics | Speech Distrubance | none |  | DMM_STUDY_TCD_SPEECH | none | n/a | T9910-107 |  |  |  | 99PMSBLUS, T9910-107, Speech Distrubance
Patient Characteristics | patientcharacteristics | Transient Ischemic Attack | none |  | DMM_STUDY_TCD_TIA | none | n/a | T9910-108 |  |  |  | 99PMSBLUS, T9910-108, Transient Ischemic Attack
Patient Characteristics | patientcharacteristics | Visual Distrubance | none |  | DMM_STUDY_TCD_VISUAL | none | n/a | T9910-109 |  |  |  | 99PMSBLUS, T9910-109, Visual Distrubance
Patient Characteristics | patientcharacteristics | Headaches | none |  | DMM_STUDY_TCD_HEADACHES | none | n/a | T9910-110 |  |  |  | 99PMSBLUS, T9910-110, Headaches
Patient Characteristics | patientcharacteristics | Vertigo | none |  | DMM_STUDY_TCD_VERTIGO | none | n/a | T9910-111 |  |  |  | 99PMSBLUS, T9910-111, Vertigo
Patient Characteristics | patientcharacteristics | Disphasia | none |  | DMM_STUDY_TCD_DYSPHASIA | none | n/a | T9910-112 |  |  |  | 99PMSBLUS, T9910-112, Disphasia
Patient Characteristics | patientcharacteristics | Aphasia | none |  | DMM_STUDY_TCD_APHASIA | none | n/a | T9910-113 |  |  |  | 99PMSBLUS, T9910-113, Aphasia
Patient Characteristics | patientcharacteristics | Diabetes | none |  | DMM_STUDY_TCD_DIABETES | none | n/a | T9910-114 |  |  |  | 99PMSBLUS, T9910-114, Diabetes
Patient Characteristics | patientcharacteristics | Hypertension | none |  | DMM_STUDY_TCD_HYPERTENSION | none | n/a | T9910-115 |  |  |  | 99PMSBLUS, T9910-115, Hypertension
Patient Characteristics | patientcharacteristics | Carotid Artery Disease | none |  | DMM_STUDY_TCD_CAD | none | n/a | T9910-116 |  |  |  | 99PMSBLUS, T9910-116, Carotid Artery Disease
Patient Characteristics | patientcharacteristics | Peripheral Disease | none |  | DMM_STUDY_TCD_PERIPHERAL_DISEASE | none | n/a | T9910-117 |  |  |  | 99PMSBLUS, T9910-117, Peripheral Disease
Patient Characteristics | patientcharacteristics | Smoker | none |  | DMM_STUDY_TCD_SMOKER | none | n/a | T9910-118 |  |  |  | 99PMSBLUS, T9910-118, Smoker
Patient Characteristics | patientcharacteristics | Family History | none |  | DMM_STUDY_TCD_FAMILY_HISTORY | none | n/a | T9910-119 |  |  |  | 99PMSBLUS, T9910-119, Family History
Patient Characteristics | patientcharacteristics | Paralysis in Face | none |  | DMM_STUDY_TCD_PARALYSIS_FACE | none | n/a | T9910-120 |  |  |  | 99PMSBLUS, T9910-120, Paralysis in Face
Patient Characteristics | patientcharacteristics | Numbness in Face | none |  | DMM_STUDY_TCD_NUMBNESS_FACE | none | n/a | T9910-121 |  |  |  | 99PMSBLUS, T9910-121, Numbness in Face
Patient Characteristics | patientcharacteristics | Weakness in Face | none |  | DMM_STUDY_TCD_WEAKNESS_FACE | none | n/a | T9910-122 |  |  |  | 99PMSBLUS, T9910-122, Weakness in Face
Patient Characteristics | patientcharacteristics | Paralysis in Upper Extremities | none |  | DMM_STUDY_TCD_PARALYSIS_UE | none | n/a | T9910-123 |  |  |  | 99PMSBLUS, T9910-123, Paralysis in Upper Extremities
Patient Characteristics | patientcharacteristics | Numbness in Upper Extremities | none |  | DMM_STUDY_TCD_NUMBNESS_UE | none | n/a | T9910-124 |  |  |  | 99PMSBLUS, T9910-124, Numbness in Upper Extremities
Patient Characteristics | patientcharacteristics | Weakness in Upper Extremities | none |  | DMM_STUDY_TCD_WEAKNESS_UE | none | n/a | T9910-125 |  |  |  | 99PMSBLUS, T9910-125, Weakness in Upper Extremities
Patient Characteristics | patientcharacteristics | Paralysis in Lower Extremities | none |  | DMM_STUDY_TCD_PARALYSIS_LE | none | n/a | T9910-126 |  |  |  | 99PMSBLUS, T9910-126, Paralysis in Lower Extremities
Patient Characteristics | patientcharacteristics | Numbness in Lower Extremities | none |  | DMM_STUDY_TCD_NUMBNESS_LE | none | n/a | T9910-127 |  |  |  | 99PMSBLUS, T9910-127, Numbness in Lower Extremities
Patient Characteristics | patientcharacteristics | Weakness in Lower Extremities | none |  | DMM_STUDY_TCD_WEAKNESS_LE | none | n/a | T9910-128 |  |  |  | 99PMSBLUS, T9910-128, Weakness in Lower Extremities
Patient Characteristics | patientcharacteristics | Hematocrit | none |  | DMM_STUDY_TCD_HEMATOCRIT | none | n/a | T9910-129 |  |  |  | 99PMSBLUS, T9910-129, Hematocrit
Patient Characteristics | patientcharacteristics | Heart Rate | none |  | DMM_STUDY_TCD_HEART_RATE | none | n/a | T9910-130 |  |  |  | 99PMSBLUS, T9910-130, Heart Rate
Patient Characteristics | patientcharacteristics | Systolic Blood Pressure (Right) | none |  | DMM_STUDY_TCD_RT_ARM_BPH | none | n/a | T9910-131 |  |  |  | 99PMSBLUS, T9910-131, Systolic Blood Pressure (Right)
Patient Characteristics | patientcharacteristics | Diastolic Blood Pressure (Right) | none |  | DMM_STUDY_TCD_RT_ARM_BPL | none | n/a | T9910-132 |  |  |  | 99PMSBLUS, T9910-132, Diastolic Blood Pressure (Right)
Patient Characteristics | patientcharacteristics | Systolic Blood Pressure (Left) | none |  | DMM_STUDY_TCD_LT_ARM_BPH | none | n/a | T9910-133 |  |  |  | 99PMSBLUS, T9910-133, Systolic Blood Pressure (Left)
Patient Characteristics | patientcharacteristics | Diastolic Blood Pressure (Left) | none |  | DMM_STUDY_TCD_LT_ARM_BPL | none | n/a | T9910-134 |  |  |  | 99PMSBLUS, T9910-134, Diastolic Blood Pressure (Left)
Patient Characteristics | patientcharacteristics | Additional Clinical Information | none |  | DMM_STUDY_TCD_ADD_CLINICAL_INFO | none | n/a | T9910-135 |  |  |  | 99PMSBLUS, T9910-135, Additional Clinical Information
Patient Characteristics | patientcharacteristics | Systolic Blood Pressure | none |  | DMM_STUDY_PEDCARD_SYS_BP | none | n/a | F-008EC |  |  |  | SRT, F-008EC, Systolic Blood Pressure
Patient Characteristics | patientcharacteristics | Diastolic Blood Pressure | none |  | DMM_STUDY_PEDCARD_DIAS_BP | none | n/a | F-008ED |  |  |  | SRT, F-008ED, Diastolic Blood Pressure
Patient Characteristics | patientcharacteristics | History of Rheumatic Fever | none |  | DMM_STUDY_CARDIAC_HX_OF_RHEUMATIC_FEVER | none | n/a | T9910-76 |  |  |  | 99PMSBLUS, T9910-76, History of Rheumatic Fever
Patient Characteristics | patientcharacteristics | Congestive Heart Failure | none |  | DMM_STUDY_CARDIAC_CONGESTIVE_HEART_FAILURE | none | n/a | T9910-77 |  |  |  | 99PMSBLUS, T9910-77, Congestive Heart Failure
Patient Characteristics | patientcharacteristics | Surgeries | none |  | DMM_STUDY_CARDIAC_SURGERIES | none | n/a | T9910-78 |  |  |  | 99PMSBLUS, T9910-78, Surgeries
Patient Characteristics | patientcharacteristics | Murmur | none |  | DMM_STUDY_CARDIAC_MURMUR | none | n/a | T9910-79 |  |  |  | 99PMSBLUS, T9910-79, Murmur
Patient Characteristics | patientcharacteristics | Murmur Type | none |  | DMM_STUDY_CARDIAC_MURMUR_TYPE | none | n/a | T9910-80 |  |  |  | 99PMSBLUS, T9910-80, Murmur Type
Patient Characteristics | patientcharacteristics | Murmur Grade | none |  | DMM_STUDY_CARDIAC_MURMUR_GRADE | none | n/a | T9910-81 |  |  |  | 99PMSBLUS, T9910-81, Murmur Grade
Patient Characteristics | patientcharacteristics | Arrythmia | none |  | DMM_STUDY_CARDIAC_ARRHYTHMIA | none | n/a | T9910-82 |  |  |  | 99PMSBLUS, T9910-82, Arrythmia
Patient Characteristics | patientcharacteristics | Chest Pain | none |  | DMM_STUDY_CARDIAC_CHEST_PAIN | none | n/a | T9910-83 |  |  |  | 99PMSBLUS, T9910-83, Chest Pain
Patient Characteristics | patientcharacteristics | Jugular Venous Distention | none |  | DMM_STUDY_CARDIAC_JUGULAR_VENOUS_DISTENTION | none | n/a | T9910-84 |  |  |  | 99PMSBLUS, T9910-84, Jugular Venous Distention
Patient Characteristics | patientcharacteristics | Dyspnea | none |  | DMM_STUDY_CARDIAC_DYSPNEA | none | n/a | T9910-85 |  |  |  | 99PMSBLUS, T9910-85, Dyspnea
Patient Characteristics | patientcharacteristics | Peripheral Edema | none |  | DMM_STUDY_CARDIAC_PERIPHERAL_EDEMA | none | n/a | T9910-86 |  |  |  | 99PMSBLUS, T9910-86, Peripheral Edema
Patient Characteristics | patientcharacteristics | Fatigue | none |  | DMM_STUDY_CARDIAC_FATIGUE | none | n/a | T9910-87 |  |  |  | 99PMSBLUS, T9910-87, Fatigue
Patient Characteristics | patientcharacteristics | Ascites | none |  | DMM_STUDY_CARDIAC_ASCITES | none | n/a | T9910-88 |  |  |  | 99PMSBLUS, T9910-88, Ascites
Patient Characteristics | patientcharacteristics | Syncope | none |  | DMM_STUDY_CARDIAC_SYNCOPE | none | n/a | T9910-89 |  |  |  | 99PMSBLUS, T9910-89, Syncope
Patient Characteristics | patientcharacteristics | Infection | none |  | DMM_STUDY_CARDIAC_INFECTION | none | n/a | T9910-90 |  |  |  | 99PMSBLUS, T9910-90, Infection
Patient Characteristics | patientcharacteristics | Dizziness | none |  | DMM_STUDY_CARDIAC_DIZZINESS | none | n/a | T9910-91 |  |  |  | 99PMSBLUS, T9910-91, Dizziness
Patient Characteristics | patientcharacteristics | Fever of Unknown Origin | none |  | DMM_STUDY_CARDIAC_FEVER_OF_UNKNOWN_ORIGIN | none | n/a | T9910-92 |  |  |  | 99PMSBLUS, T9910-92, Fever of Unknown Origin
Patient Characteristics | patientcharacteristics | Hemoptysis | none |  | DMM_STUDY_CARDIAC_HEMOPTYSIS | none | n/a | T9910-93 |  |  |  | 99PMSBLUS, T9910-93, Hemoptysis
Patient Characteristics | patientcharacteristics | TIA / Stroke | none |  | DMM_STUDY_CARDIAC_TIA_STROKE | none | n/a | T9910-94 |  |  |  | 99PMSBLUS, T9910-94, TIA / Stroke
Patient Characteristics | patientcharacteristics | Bioprosthetic Value Replacement Type | none |  | DMM_STUDY_CARDIAC_BIOPROSTHETIC_TYPE | none | n/a | T9910-95 |  |  |  | 99PMSBLUS, T9910-95, Bioprosthetic Value Replacement Type
Patient Characteristics | patientcharacteristics | Bioprosthetic Value Replacement Date | none |  | DMM_STUDY_CARDIAC_BIOPROSTHETIC_DATE | none | n/a | T9910-96 |  |  |  | 99PMSBLUS, T9910-96, Bioprosthetic Value Replacement Date
Patient Characteristics | patientcharacteristics | Mechanical Value Replacement Type | none |  | DMM_STUDY_CARDIAC_MECHANICAL_TYPE | none | n/a | T9910-97 |  |  |  | 99PMSBLUS, T9910-97, Mechanical Value Replacement Type
Patient Characteristics | patientcharacteristics | Mechanical Value Replacement Date | none |  | DMM_STUDY_CARDIAC_MECHANICAL_DATE | none | n/a | T9910-98 |  |  |  | 99PMSBLUS, T9910-98, Mechanical Value Replacement Date
Patient Characteristics | patientcharacteristics | Pacemaker | none |  | DMM_STUDY_CARDIAC_PACEMAKER | none | n/a | T9910-99 |  |  |  | 99PMSBLUS, T9910-99, Pacemaker
Patient Characteristics | patientcharacteristics | Murmur | none |  | DMM_STUDY_PEDCARD_MURMUR | none | n/a | R-00302 |  |  |  | SRT, R-00302, Murmur
Patient Characteristics | patientcharacteristics | Arrhythmia | none |  | DMM_STUDY_PEDCARD_ARRHYTHMIA | none | n/a | D3-30000 |  |  |  | SRT, D3-30000, Arrhythmia
Patient Characteristics | patientcharacteristics | Chest Pain | none |  | DMM_STUDY_PEDCARD_CHEST_PAIN | none | n/a | F-37000 |  |  |  | SRT, F-37000, Chest Pain
Patient Characteristics | patientcharacteristics | Cardiomegaly | none |  | DMM_STUDY_PEDCARD_CARDIOMEGALY | none | n/a | D3-10008 |  |  |  | SRT, D3-10008, Cardiomegaly
Patient Characteristics | patientcharacteristics | Cyanosis | none |  | DMM_STUDY_PEDCARD_CYANOSIS | none | n/a | M-04100 |  |  |  | SRT, M-04100, Cyanosis
Patient Characteristics | patientcharacteristics | Dextrocardia | none |  | DMM_STUDY_PEDCARD_DEXTRACARDIA | none | n/a | D4-31B16 |  |  |  | SRT, D4-31B16, Dextrocardia
Patient Characteristics | patientcharacteristics | Dyspnea | none |  | DMM_STUDY_PEDCARD_DYSPNIA | none | n/a | F-201B3 |  |  |  | SRT, F-201B3, Dyspnea
Patient Characteristics | patientcharacteristics | Fever | none |  | DMM_STUDY_PEDCARD_FEVER | none | n/a | F-0A44A |  |  |  | SRT, F-0A44A, Fever
Patient Characteristics | patientcharacteristics | Hemoptysis | none |  | DMM_STUDY_PEDCARD_HEMOPTYSIS | none | n/a | F-24210 |  |  |  | SRT, F-24210, Hemoptysis
Patient Characteristics | patientcharacteristics | Hypertension | none |  | DMM_STUDY_PEDCARD_SYS_HYPERTENSION | none | n/a | D3-02000 |  |  |  | SRT, D3-02000, Hypertension
Patient Characteristics | patientcharacteristics | Insulin dependent mother (IDM) | none |  | DMM_STUDY_PEDCARD_IDM | none | n/a | G-0586 |  |  |  | SRT, G-0586, Insulin dependent mother (IDM)
Patient Characteristics | patientcharacteristics | Mesocardia | none |  | DMM_STUDY_PEDCARD_MYSOCARDIA | none | n/a | D4-31B24 |  |  |  | SRT, D4-31B24, Mesocardia
Patient Characteristics | patientcharacteristics | Syncope | none |  | DMM_STUDY_PEDCARD_SYNCOPE | none | n/a | D3-00006 |  |  |  | SRT, D3-00006, Syncope
Patient Characteristics | patientcharacteristics | Cardiac pacemaker | none |  | DMM_STUDY_PEDCARD_PACEMAKER | none | n/a | A-11100 |  |  |  | DCM, A-11100, Cardiac pacemaker
Patient Characteristics | patientcharacteristics | Ventricular Septal Defect | none |  | DMM_STUDY_PEDCARD_VSD | none | n/a | D4-31150 |  |  |  | SRT, D4-31150, Ventricular Septal Defect
Patient Characteristics | patientcharacteristics | Patent Ductus Arteriosus | none |  | DMM_STUDY_PEDCARD_PDA | none | n/a | D4-32012 |  |  |  | SRT, D4-32012, Patent Ductus Arteriosus
Patient Characteristics | patientcharacteristics | Pulmonic valve stenosis | none |  | DMM_STUDY_PEDCARD_PS | none | n/a | D3-29051 |  |  |  | SRT, D3-29051, Pulmonic valve stenosis
Patient Characteristics | patientcharacteristics | Aortic Stenosis | none |  | DMM_STUDY_PEDCARD_AS_CHD | none | n/a | D3-29021 |  |  |  | SRT, D3-29021, Aortic Stenosis
Patient Characteristics | patientcharacteristics | Coarctation of the Aorta | none |  | DMM_STUDY_PEDCARD_COA | none | n/a | D4-32014 |  |  |  | SRT, D4-32014, Coarctation of the Aorta
Patient Characteristics | patientcharacteristics | Tetralogy of Fallot | none |  | DMM_STUDY_PEDCARD_TOF | none | n/a | D4-31110 |  |  |  | SRT, D4-31110, Tetralogy of Fallot
Patient Characteristics | patientcharacteristics | Partial anomalous pulmonary venous connection | none |  | DMM_STUDY_PEDCARD_PAPVR | none | n/a | D4-33622 |  |  |  | SRT, D4-33622, Partial anomalous pulmonary venous connection
Patient Characteristics | patientcharacteristics | Common atrioventricular canal | none |  | DMM_STUDY_PEDCARD_COMP_ECD | none | n/a | D4-31303 |  |  |  | SRT, D4-31303, Common atrioventricular canal
Patient Characteristics | patientcharacteristics | Atrial septal defect with endocardial cushion defect, partial type | none |  | DMM_STUDY_PEDCARD_PART_ECD | none | n/a | D4-31310 |  |  |  | SRT, D4-31310, Atrial septal defect with endocardial cushion defect, partial type
Patient Characteristics | patientcharacteristics | Complete transposition of great vessels | none |  | DMM_STUDY_PEDCARD_COMP_TGA | none | n/a | D4-31010 |  |  |  | SRT, D4-31010, Complete transposition of great vessels
Patient Characteristics | patientcharacteristics | Interrupted Aortic Arch | none |  | DMM_STUDY_PEDCARD_INTERR_AO_ARCH | none | n/a | D3-83001 |  |  |  | SRT, D3-83001, Interrupted Aortic Arch
Patient Characteristics | patientcharacteristics | Mitral stenosis | none |  | DMM_STUDY_PEDCARD_MS | none | n/a | D3-29011 |  |  |  | SRT, D3-29011, Mitral stenosis
Patient Characteristics | patientcharacteristics | Mitral valve prolapse | none |  | DMM_STUDY_PEDCARD_MVP | none | n/a | D3-29013 |  |  |  | SRT, D3-29013, Mitral valve prolapse
Patient Characteristics | patientcharacteristics | HIV Positive | none |  | DMM_STUDY_PEDCARD_HIV | none | n/a | F-0331B |  |  |  | SRT, F-0331B, HIV Positive
Patient Characteristics | patientcharacteristics | Mitral regurgitation | none |  | DMM_STUDY_PEDCARD_MR | none | n/a | D3-29012 |  |  |  | SRT, D3-29012, Mitral regurgitation
Patient Characteristics | patientcharacteristics | Acute febrile mucocutaneous lymph node syndrome | none |  | DMM_STUDY_PEDCARD_KAWAZAKI | none | n/a | D3-81660 |  |  |  | SRT, D3-81660, Acute febrile mucocutaneous lymph node syndrome
Patient Characteristics | patientcharacteristics | Rheumatic Fever | none |  | DMM_STUDY_PEDCARD_RHEUMATIC_FEVER | none | n/a | D3-17100 |  |  |  | SRT, D3-17100, Rheumatic Fever
Patient Characteristics | patientcharacteristics | Aortic regurgitation | none |  | DMM_STUDY_PEDCARD_AR | none | n/a | D3-29022 |  |  |  | SRT, D3-29022, Aortic regurgitation
Patient Characteristics | patientcharacteristics | Arrhythmia Type | none |  | DMM_STUDY_PEDCARD_ARRHYTHMIA_TYPE | none | n/a | T9910-176 |  |  |  | 99PMSBLUS, T9910-176, Arrhythmia Type
Patient Characteristics | patientcharacteristics | Murmur Type | none |  | DMM_STUDY_PEDCARD_MURMUR_TYPE | none | n/a | T9910-178 |  |  |  | 99PMSBLUS, T9910-178, Murmur Type
Patient Characteristics | patientcharacteristics | Atrial Septal Defect | none |  | DMM_STUDY_PEDCARD_ASD | none | n/a | D4-31220 |  |  |  | SRT, D4-31220, Atrial Septal Defect
Patient Characteristics | patientcharacteristics | Cardiomyopathy Type | none |  | DMM_STUDY_PEDCARD_CARDIOMYOPATHY_TYPE | none | n/a | T9910-180 |  |  |  | 99PMSBLUS, T9910-180, Cardiomyopathy Type
Patient Characteristics | patientcharacteristics | Infections Type | none |  | DMM_STUDY_PEDCARD_INFECTIONS_TYPE | none | n/a | T9910-181 |  |  |  | 99PMSBLUS, T9910-181, Infections Type
Patient Characteristics | patientcharacteristics | Surgeries Type | none |  | DMM_STUDY_PEDCARD_SURGERIES_TYPE | none | n/a | T9910-182 |  |  |  | 99PMSBLUS, T9910-182, Surgeries Type
Patient Characteristics | patientcharacteristics | Comment | none |  | DMM_STUDY_PEDCARD_COMMENTS | none | n/a | 121106 |  |  |  | DCM, 121106, Comment
Patient Characteristics | patientcharacteristics | BSA = 0.0004688*WT^(0.8168-0.0154*log(WT)) | none |  | DMM_STUDY_PEDCARD_WEIGHT_ONLY | none | n/a | 122246 |  |  |  | DCM, 122246, BSA = 0.0004688*WT^(0.8168-0.0154*log(WT))
Patient Characteristics | patientcharacteristics | Aortic stenosis | none |  | DMM_STUDY_PEDCARD_AS_AHD | none | n/a | D3-29021 |  |  |  | SRT, D3-29021, Aortic stenosis
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SVC Dias Vel | Vel | meas | m_p_SVC_DiasVel | r_Vel | unilateral | 125195\122202\F-32011\M-2460D | SRT, M-2460D, Right Superior vena cava | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SVC Sys Vel | Vel | meas | m_p_SVC_SysVel | r_Vel | unilateral | 125195\122204\M-2460D\R-FAB5B | SRT, M-2460D, Right Superior vena cava | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SVC A Vel | Vel | meas | m_p_SVC_ARevsVel | r_Vel | unilateral | 125195\59079-4\M-2460D | SRT, M-2460D, Right Superior vena cava | LN, 59079-4, Peak Reversal Velocity during Atrial Contraction |  |  | UCUM, mm/s, mm/s
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SVC A Dur | Time | meas | m_p_SVC_ARevsDur | r_DopTime | unilateral | 125195\59105-7\M-2460D | SRT, M-2460D, Right Superior vena cava | LN, 59105-7, A-Wave Duration |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SVC Diam | Dist | meas | m_ped_SVC_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\M-2460D | SRT, M-2460D, Right Superior vena cava | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVC Dias Vel | Vel | meas | m_p_IVC_DiasVel | r_Vel | unilateral | 125195\122202\F-32011\T-48710 | SRT, T-48710, Inferior vena cava | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVC Sys Vel | Vel | meas | m_p_IVC_SysVel | r_Vel | unilateral | 125195\122204\R-FAB5B\T-48710 | SRT, T-48710, Inferior vena cava | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVC A Vel | Vel | meas | m_p_IVC_ARevsVel | r_Vel | unilateral | 125195\59079-4\T-48710 | SRT, T-48710, Inferior vena cava | LN, 59079-4, Peak Reversal Velocity during Atrial Contraction |  |  | UCUM, mm/s, mm/s
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVC A Dur | Time | meas | m_p_IVC_ARevsDur | r_DopTime | unilateral | 125195\59105-7\T-48710 | SRT, T-48710, Inferior vena cava | LN, 59105-7, A-Wave Duration |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVC Diam | Dist | meas | m_ped_IVC_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-48710 | SRT, T-48710, Inferior vena cava | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Hepatic Dias Vel | Vel | meas | m_p_HepDiasVel | r_Vel | unilateral | 125195\122202\F-32011\T-48720 | SRT, T-48720, Hepatic Vein | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Hepatic Sys Vel | Vel | meas | m_p_HepSysVel | r_Vel | unilateral | 125195\122204\R-FAB5B\T-48720 | SRT, T-48720, Hepatic Vein | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Hepatic A Vel | Vel | meas | m_p_HepARevsVel | r_Vel | unilateral | 125195\59079-4\T-48720 | SRT, T-48720, Hepatic Vein | LN, 59079-4, Peak Reversal Velocity during Atrial Contraction |  |  | UCUM, mm/s, mm/s
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Hepatic A Dur | Time | meas | m_p_HepARevsDur | r_DopTime | unilateral | 125195\59105-7\T-48720 | SRT, T-48720, Hepatic Vein | LN, 59105-7, A-Wave Duration |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Pulm Dias Vel | Vel | meas | m_p_PulmDiasVel | r_Vel | unilateral | 125195\122202\F-32011\T-48581 | SRT, T-48581, Pulmonary Vein | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RU PulmV Dias Vel | Vel | meas | m_p_RU_PulmDiasVel | r_Vel | unilateral | 125195\122202\F-32011\T-48510\T-48581 | SRT, T-48581, Pulmonary Vein | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s | SRT, T-48510, Right Superior Pulmonary Vein |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RL PulmV Dias Vel | Vel | meas | m_p_RL_PulmDiasVel | r_Vel | unilateral | 125195\122202\F-32011\T-48520\T-48581 | SRT, T-48581, Pulmonary Vein | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s | SRT, T-48520, Right Inferior Pulmonary Vein |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LU PulmV Dias Vel | Vel | meas | m_p_LU_PulmDiasVel | r_Vel | unilateral | 125195\122202\F-32011\T-48530\T-48581 | SRT, T-48581, Pulmonary Vein | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s | SRT, T-48530, Left Superior Pulmonary Vein |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LL PulmV Dias Vel | Vel | meas | m_p_LL_PulmDiasVel | r_Vel | unilateral | 125195\122202\F-32011\T-48540\T-48581 | SRT, T-48581, Pulmonary Vein | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s | SRT, T-48540, Left Inferior Pulmonary Vein |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Pulm Sys Vel | Vel | meas | m_p_PulmSysVel | r_Vel | unilateral | 125195\122204\R-FAB5B\T-48581 | SRT, T-48581, Pulmonary Vein | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RU PulmV Sys Vel | Vel | meas | m_p_RU_PulmSysVel | r_Vel | unilateral | 125195\122204\R-FAB5B\T-48510\T-48581 | SRT, T-48581, Pulmonary Vein | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s | SRT, T-48510, Right Superior Pulmonary Vein |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RL PulmV Sys Vel | Vel | meas | m_p_RL_PulmSysVel | r_Vel | unilateral | 125195\122204\R-FAB5B\T-48520\T-48581 | SRT, T-48581, Pulmonary Vein | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s | SRT, T-48520, Right Inferior Pulmonary Vein |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LU PulmV Sys Vel | Vel | meas | m_p_LU_PulmSysVel | r_Vel | unilateral | 125195\122204\R-FAB5B\T-48530\T-48581 | SRT, T-48581, Pulmonary Vein | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s | SRT, T-48530, Left Superior Pulmonary Vein |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LL PulmV Sys Vel | Vel | meas | m_p_LL_PulmSysVel | r_Vel | unilateral | 125195\122204\R-FAB5B\T-48540\T-48581 | SRT, T-48581, Pulmonary Vein | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s | SRT, T-48540, Left Inferior Pulmonary Vein |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Pulm A Vel | Vel | meas | m_p_PulmARevsVel | r_Vel | unilateral | 125195\59079-4\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59079-4, Peak Reversal Velocity during Atrial Contraction |  |  | UCUM, mm/s, mm/s
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RU PulmV A Vel | Vel | meas | m_p_RU_PulmARevsVel | r_Vel | unilateral | 125195\59079-4\T-48510\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59079-4, Peak Reversal Velocity during Atrial Contraction |  |  | UCUM, mm/s, mm/s | SRT, T-48510, Right Superior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RL PulmV A Vel | Vel | meas | m_p_RL_PulmARevsVel | r_Vel | unilateral | 125195\59079-4\T-48520\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59079-4, Peak Reversal Velocity during Atrial Contraction |  |  | UCUM, mm/s, mm/s | SRT, T-48520, Right Inferior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LU PulmV A Vel | Vel | meas | m_p_LU_PulmARevsVel | r_Vel | unilateral | 125195\59079-4\T-48530\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59079-4, Peak Reversal Velocity during Atrial Contraction |  |  | UCUM, mm/s, mm/s | SRT, T-48530, Left Superior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LL PulmV A Vel | Vel | meas | m_p_LL_PulmARevsVel | r_Vel | unilateral | 125195\59079-4\T-48540\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59079-4, Peak Reversal Velocity during Atrial Contraction |  |  | UCUM, mm/s, mm/s | SRT, T-48540, Left Inferior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Pulm A Dur | Time | meas | m_p_PulmARevsDur | r_DopTime | unilateral | 125195\59105-7\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59105-7, A-Wave Duration |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RU PulmV A Dur | Time | meas | m_p_RU_PulmARevsDur | r_DopTime | unilateral | 125195\59105-7\T-48510\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59105-7, A-Wave Duration |  |  | UCUM, ms, ms | SRT, T-48510, Right Superior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RL PulmV A Dur | Time | meas | m_p_RL_PulmARevsDur | r_DopTime | unilateral | 125195\59105-7\T-48520\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59105-7, A-Wave Duration |  |  | UCUM, ms, ms | SRT, T-48520, Right Inferior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LU PulmV A Dur | Time | meas | m_p_LU_PulmARevsDur | r_DopTime | unilateral | 125195\59105-7\T-48530\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59105-7, A-Wave Duration |  |  | UCUM, ms, ms | SRT, T-48530, Left Superior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LL PulmV A Dur | Time | meas | m_p_LL_PulmARevsDur | r_DopTime | unilateral | 125195\59105-7\T-48540\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59105-7, A-Wave Duration |  |  | UCUM, ms, ms | SRT, T-48540, Left Inferior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RU PulmV Diam | Dist | meas | m_ped_RU_PulmV_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-48510\T-48581 | SRT, T-48581, Pulmonary Vein | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-48510, Right Superior Pulmonary Vein | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RL PulmV Diam | Dist | meas | m_ped_RL_PulmV_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-48520\T-48581 | SRT, T-48581, Pulmonary Vein | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-48520, Right Inferior Pulmonary Vein | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LU PulmV Diam | Dist | meas | m_ped_LU_PulmV_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-48530\T-48581 | SRT, T-48581, Pulmonary Vein | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-48530, Left Superior Pulmonary Vein | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LL PulmV Diam | Dist | meas | m_ped_LL_PulmV_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-48540\T-48581 | SRT, T-48581, Pulmonary Vein | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-48540, Left Inferior Pulmonary Vein | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Coarctation Diam | Dist | meas | m_ped_Coarc_Diam | r_Distance | unilateral | 125195\D4-32014\G-03A2\M-02550 | SRT, D4-32014, Coarctation of aorta | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA Dimen (MM) | Dist | meas | m_ped_LA_Dimen_MM | r_Distance | unilateral | 125195\29469-4\G-0394\R-FAB5B\T-32300 | SRT, T-32300, Left Atrium | LN, 29469-4, Left Atrium Antero-posterior Systolic Dimension |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Desc Ao Max PG | Max PG | meas | m_ped_DSC_Ao_MaxPG | r_MaxPG | unilateral | 125195\20247-3\T-42070 | SRT, T-42070, Thoracic aorta | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Desc Ao Mean PG | Mean PG | meas | m_ped_DSC_Ao_PG | r_CardPGmean | unilateral | 125195\20256-4\T-42070 | SRT, T-42070, Thoracic aorta | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Desc Ao Max PG | Vmax | meas | m_ped_DSC_Ao_MaxPG | r_Vmax | unilateral | 125195\20351-3\T-42070 | SRT, T-42070, Thoracic aorta | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Desc Ao Diam | Dist | meas | m_ped_DescAorta_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-42070 | SRT, T-42070, Thoracic aorta | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASC Ao Max PG | Max PG | meas | m_ped_ASC_Ao_MaxPG | r_MaxPG | unilateral | 125195\20247-3\T-42100 | SRT, T-42100, Ascending aorta | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Asc Ao Mean PG | Mean PG | meas | m_ped_ASC_Ao_PG | r_CardPGmean | unilateral | 125195\20256-4\T-42100 | SRT, T-42100, Ascending aorta | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASC Ao Max PG | Vmax | meas | m_ped_ASC_Ao_MaxPG | r_Vmax | unilateral | 125195\20351-3\T-42100 | SRT, T-42100, Ascending aorta | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Asc Ao Diam | Dist | meas | m_ped_AscAorta_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-42100 | SRT, T-42100, Ascending aorta | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Ao ST Jx Diam | Dist | meas | m_p_ST_J_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-42102 | SRT, T-42102, Aortic Sinotubular Junction | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AoR Diam (MM) | Dist | meas | m_ped_AoR_Diam_MM | r_Distance | unilateral | 125195\G-0394\M-02550\T-42110 | SRT, T-42110, Root of Aorta | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AoR Diam (2D) | Dist | meas | m_ped_AoR_Diam_2D | r_Distance | unilateral | 125195\G-03A2\M-02550\T-42110 | SRT, T-42110, Root of Aorta | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Ao Sinus Diam | Dist | meas | m_ped_Ao_Sinus_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-42200 | SRT, T-42200, Structure Sinus of Valsalva | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Ao Arch Diam | Dist | meas | m_ped_AoArch_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-42300 | SRT, T-42300, Aortic arch | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Ao Arch Dist Diam | Dist | meas | m_p_AoArch_Dist_Diam | r_Distance | unilateral | 125195\G-03A2\G-A119\M-02550\T-42300 | SRT, T-42300, Aortic arch | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, G-A119, Distal | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Ao Isthmus Diam | Dist | meas | m_ped_AoIsth_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-42310 | SRT, T-42310, Aortic isthmus | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Left Main | Dist | meas | m_ped_L_Main_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-43107 | SRT, T-43107, Left Main Coronary Artery | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LAD | Dist | meas | m_ped_LAD_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-43110 | SRT, T-43110, Anterior Descending Branch of Left Coronary Artery | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Cx | Dist | meas | m_ped_LCx_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-43120 | SRT, T-43120, Circumflex Coronary Artery | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RCA | Dist | meas | m_ped_RCA_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-43203 | SRT, T-43203, Right Coronary Artery | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PD | Dist | meas | m_ped_PDC_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-D0878 | SRT, T-D0878, Posterior Descending Coronary Artery | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PDA Dias Vel | Vel | meas | m_p_PDA_DiasVel | r_Vel | unilateral | 125195\122202\D4-32012\F-32011 | SRT, D4-32012, Patent Ductus Arteriosus | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PDA Sys Vel | Vel | meas | m_p_PDA_SysVel | r_Vel | unilateral | 125195\122204\D4-32012\R-FAB5B | SRT, D4-32012, Patent Ductus Arteriosus | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PDA Diam | Dist | meas | m_ped_PDA_Diam | r_Distance | unilateral | 125195\D4-32012\G-03A2\M-02550 | SRT, D4-32012, Patent Ductus Arteriosus | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA Dimen (2D) | Dist | meas | m_ped_RA_Dimen_2D | r_Distance | unilateral | 125195\121206\G-03A2\T-32200 | SRT, T-32200, Right Atrium | DCM, 121206, Distance |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA Pressure | Pressure | meas | m_ped_RA_Press | r_UserPressure | unilateral | 125195\F-31000\R-FAB5B\T-32200 | SRT, T-32200, Right Atrium | SRT, F-31000, Blood Pressure |  |  | UCUM, mmHg, mmHg |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RAed Major - A4C | Dist | meas | m_ped_RAED_MAJOR_A4C | r_Distance | unilateral | 125195\F-32011\G-A193\G-A19C\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RAes Major - A4C | Dist | meas | m_ped_RAES_MAJOR_A4C | r_Distance | unilateral | 125195\G-A193\G-A19C\R-FAB5B\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RAed Minor - A4C | Dist | meas | m_ped_RAED_MINOR_A4C | r_Distance | unilateral | 125195\F-32011\G-A194\G-A19C\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RAes Minor - A4C | Dist | meas | m_ped_RAES_MINOR_A4C | r_Distance | unilateral | 125195\G-A194\G-A19C\R-FAB5B\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA Dimen (MM) | Dist | meas | m_ped_RA_Dimen_MM | r_Distance | unilateral | 125195\G-0394\M-02550\T-32200 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA Dimen (2D) | Dist | meas | m_ped_LA_Dimen_2D | r_Distance | unilateral | 125195\121206\G-03A2\T-32300 | SRT, T-32300, Left Atrium | DCM, 121206, Distance |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LAed Major - A4C | Dist | meas | m_ped_LAED_MAJOR_A4C | r_Distance | unilateral | 125195\F-32011\G-A193\G-A19C\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LAes Major - A4C | Dist | meas | m_ped_LAES_MAJOR_A4C | r_Distance | unilateral | 125195\G-A193\G-A19C\R-FAB5B\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LAed Minor - A4C | Dist | meas | m_ped_LAED_MINOR_A4C | r_Distance | unilateral | 125195\F-32011\G-A194\G-A19C\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LAes Minor - A4C | Dist | meas | m_ped_LAES_MINOR_A4C | r_Distance | unilateral | 125195\G-A194\G-A19C\R-FAB5B\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASD VTI | Vmean | meas | m_p_ASD_VTI | r_CardTAPV | unilateral | 125195\11692-1\D4-31220\R-42E61 | SRT, D4-31220, Atrial Septal Defect | LN, 11692-1, Time Averaged Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Retrograde Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASD VTI | Max PG | meas | m_p_ASD_VTI | r_MaxPG | unilateral | 125195\20247-3\D4-31220 | SRT, D4-31220, Atrial Septal Defect | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASD VTI | Mean PG | meas | m_p_ASD_VTI | r_CardPGmean | unilateral | 125195\20256-4\D4-31220 | SRT, D4-31220, Atrial Septal Defect | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASD VTI | Vmax | meas | m_p_ASD_VTI | r_Vmax | unilateral | 125195\20351-3\D4-31220 | SRT, D4-31220, Atrial Septal Defect | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASD VTI | VTI | meas | m_p_ASD_VTI | r_VTI | unilateral | 125195\20354-7\D4-31220 | SRT, D4-31220, Atrial Septal Defect | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASD Major | Dist | meas | m_ped_ASD_Major | r_Distance | unilateral | 125195\D4-31220\G-03A2\G-A193 | SRT, D4-31220, Atrial Septal Defect | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASD Minor | Dist | meas | m_ped_ASD_Minor | r_Distance | unilateral | 125195\D4-31220\G-03A2\G-A194 | SRT, D4-31220, Atrial Septal Defect | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV P½t | Vmax | meas | m_ped_TV_PHalfT | r_Vmax | unilateral | 125195\C12222-03\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12222-03, Pressure Half-Time Peak velocity |  |  | UCUM, mm/s, mm/s
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR VTI | Vmean | meas | m_ped_TR_VTI | r_CardTAPV | unilateral | 125195\11692-1\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 11692-1, Time Averaged Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV VTI | Vmean | meas | m_ped_TV_VTI | r_CardTAPV | unilateral | 125195\11692-1\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 11692-1, Time Averaged Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Accel | Slope | meas | m_ped_TV_Accel | r_DopSlope | unilateral | 125195\20167-3\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20167-3, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Accel | Time | meas | m_ped_TV_Accel | r_DopTime | unilateral | 125195\20168-1\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Decel Slope | Slope | meas | m_ped_TV_DecelSlope | r_DopSlope | unilateral | 125195\20216-8\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20216-8, Deceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Decel Time | Time | meas | m_ped_TV_DecelTime | r_DopTime | unilateral | 125195\20217-6\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR Vmax | Max PG | meas | m_ped_TR_Vmax | r_MaxPG | unilateral | 125195\20247-3\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Vmax | Max PG | meas | m_ped_TV_Vmax | r_MaxPG | unilateral | 125195\20247-3\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR VTI | Mean PG | meas | m_ped_TR_VTI | r_CardPGmean | unilateral | 125195\20256-4\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Mean PG | Mean PG | meas | m_ped_TV_MeanPG | r_CardPGmean | unilateral | 125195\20256-4\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV P½t | P½t | meas | m_ped_TV_PHalfT | r_PHalftime | unilateral | 125195\20280-4\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20280-4, Pressure Half-Time |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR Vmax | Vmax | meas | m_ped_TR_Vmax | r_Vmax | unilateral | 125195\20351-3\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Vmax | Vmax | meas | m_ped_TV_Vmax | r_Vmax | unilateral | 125195\20351-3\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR VTI | VTI | meas | m_ped_TR_VTI | r_VTI | unilateral | 125195\20354-7\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV VTI | VTI | meas | m_ped_TV_VTI | r_VTI | unilateral | 125195\20354-7\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Peak E Vel | Vel | meas | m_ped_TV_PeakEVel | r_Vel | unilateral | 125195\59080-2\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59080-2, E-Wave Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Peak A Vel | Vel | meas | m_ped_TV_PeakAVel | r_Vel | unilateral | 125195\59081-0\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59081-0, A-Wave Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Closure to Opening | Time | meas | m_ped_TV_Cl_to_Op | r_DopTime | unilateral | 125195\59082-8\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59082-8, Closure to Opening Time |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV D-E Exc | Dist | meas | m_ped_TV_DE_Exc | r_Distance | unilateral | 125195\59091-9\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59091-9, D-E Excursion |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR Radius | Dist | meas | m_ped_TR_Radius | r_Distance | unilateral | 125195\59102-4\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59102-4, Flow Radius |  |  | UCUM, mm, mm |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Radius | Dist | meas | m_ped_TV_Radius | r_Distance | unilateral | 125195\59102-4\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59102-4, Flow Radius |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV A-C Interval | Time | meas | m_ped_TV_AC_Int | r_MMTime | unilateral | 125195\59103-2\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59103-2, A-C Interval |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV A Dur | Time | meas | m_ped_TV_A_Dur | r_DopTime | unilateral | 125195\59105-7\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59105-7, A-Wave Duration |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV D-E Slope | Slope | meas | m_ped_TV_DE_Slope | r_MMSlope | unilateral | 125195\59127-1\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59127-1, D-E Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV E-F Slope | Slope | meas | m_ped_TV_EF_Slope | r_MMSlope | unilateral | 125195\59128-9\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59128-9, E-F Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR Alias Vel | Alias Vel | meas | m_ped_TR_AliasVel | r_2DVel | unilateral | 125195\59130-5\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59130-5, Alias velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Alias Vel | Alias Vel | meas | m_ped_TV_AliasVel | r_2DVel | unilateral | 125195\59130-5\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59130-5, Alias velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | HR - TV | HR | meas | m_ped_HR_TV | r_HeartRate | unilateral | 125195\8867-4\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Diam | Dist | meas | m_ped_TV_Diam | r_Distance | unilateral | 125195\G-038F\G-03A2\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038F, Cardiovascular Orifice Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Annul Diam | Dist | meas | m_ped_TV_Annul_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-3500E\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-3500E, Cardiac valve annulus | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV DFP | Time | meas | m_ped_TV_DFP | r_DopTime | unilateral | 125195\R-003A9\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, R-003A9, Tricuspid Diastolic Filling Period (DFPt) |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV - P½t | Vmax | meas | m_ped_MV_PHalfT | r_Vmax | unilateral | 125195\C12222-03\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12222-03, Pressure Half-Time Peak velocity |  |  | UCUM, mm/s, mm/s
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR VTI | Vmean | meas | m_ped_MR_VTI | r_CardTAPV | unilateral | 125195\11692-1\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 11692-1, Time Averaged Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV VTI | Vmean | meas | m_ped_MV_VTI | r_CardTAPV | unilateral | 125195\11692-1\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 11692-1, Time Averaged Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV EPSS | Dist | meas | m_ped_MV_EPSS | r_Distance | unilateral | 125195\18036-4\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 18036-4, Mitral Valve EPSS, E wave |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Accel | Slope | meas | m_ped_MV_Accel | r_DopSlope | unilateral | 125195\20167-3\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20167-3, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Accel | Time | meas | m_ped_MV_Accel | r_DopTime | unilateral | 125195\20168-1\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Decel Slope | Slope | meas | m_ped_MV_DecelSlope | r_DopSlope | unilateral | 125195\20216-8\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20216-8, Deceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Decel Time | Time | meas | m_ped_MV_DecelTime | r_DopTime | unilateral | 125195\20217-6\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR Vmax | Max PG | meas | m_ped_MR_Vmax | r_MaxPG | unilateral | 125195\20247-3\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Vmax | Max PG | meas | m_ped_MV_Vmax | r_MaxPG | unilateral | 125195\20247-3\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR VTI | Mean PG | meas | m_ped_MR_VTI | r_CardPGmean | unilateral | 125195\20256-4\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Mean PG | Mean PG | meas | m_ped_MV_MeanPG | r_CardPGmean | unilateral | 125195\20256-4\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV - P½t | P½t | meas | m_ped_MV_PHalfT | r_PHalftime | unilateral | 125195\20280-4\T-35300 | SRT, T-35300, Mitral Valve | LN, 20280-4, Pressure Half-Time |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR Vmax | Vmax | meas | m_ped_MR_Vmax | r_Vmax | unilateral | 125195\20351-3\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Vmax | Vmax | meas | m_ped_MV_Vmax | r_Vmax | unilateral | 125195\20351-3\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR VTI | VTI | meas | m_ped_MR_VTI | r_VTI | unilateral | 125195\20354-7\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV VTI | VTI | meas | m_ped_MV_VTI | r_VTI | unilateral | 125195\20354-7\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Peak E Vel | Vel | meas | m_ped_MV_PeakEVel | r_Vel | unilateral | 125195\59080-2\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 59080-2, E-Wave Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Peak A Vel | Vel | meas | m_ped_MV_PeakAVel | r_Vel | unilateral | 125195\59081-0\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 59081-0, A-Wave Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Closure to Opening | Time | meas | m_ped_MV_Cl_to_Op | r_DopTime | unilateral | 125195\59082-8\T-35300 | SRT, T-35300, Mitral Valve | LN, 59082-8, Closure to Opening Time |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV D-E Exc | Dist | meas | m_ped_MV_DE_Exc | r_Distance | unilateral | 125195\59091-9\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 59091-9, D-E Excursion |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV E-E Sep | Dist | meas | m_ped_MV_EE_Separ | r_Distance | unilateral | 125195\59098-4\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 59098-4, Mitral Valve E-septal Separation |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR Radius | Dist | meas | m_ped_MR_Radius | r_Distance | unilateral | 125195\59102-4\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 59102-4, Flow Radius |  |  | UCUM, mm, mm |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Radius | Dist | meas | m_ped_MV_Radius | r_Distance | unilateral | 125195\59102-4\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 59102-4, Flow Radius |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV A-C Interval | Time | meas | m_ped_MV_AC_Int | r_MMTime | unilateral | 125195\59103-2\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 59103-2, A-C Interval |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV A Dur | Time | meas | m_ped_MV_A_Dur | r_DopTime | unilateral | 125195\59105-7\T-35300 | SRT, T-35300, Mitral Valve | LN, 59105-7, A-Wave Duration |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV D-E Slope | Slope | meas | m_ped_MV_DE_Slope | r_MMSlope | unilateral | 125195\59127-1\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 59127-1, D-E Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV E-F Slope | Slope | meas | m_ped_MV_EF_Slope | r_MMSlope | unilateral | 125195\59128-9\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 59128-9, E-F Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR Alias Vel | Alias Vel | meas | m_ped_MR_AliasVel | r_2DVel | unilateral | 125195\59130-5\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 59130-5, Alias velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Alias Vel | Alias Vel | meas | m_ped_MV_AliasVel | r_2DVel | unilateral | 125195\59130-5\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 59130-5, Alias velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | HR - MV | HR | meas | m_ped_HR_MV | r_HeartRate | unilateral | 125195\8867-4\T-35300 | SRT, T-35300, Mitral Valve | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Area (Planim) | Area | meas | m_ped_MV_Area | r_Area | unilateral | 125195\125220\G-038E\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125220, Planimetry
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Diam | Dist | meas | m_ped_MV_Diam | r_Distance | unilateral | 125195\G-038F\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038F, Cardiovascular Orifice Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Major | Dist | meas | m_ped_MV_Major | r_Distance | unilateral | 125195\G-A193\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Minor | Dist | meas | m_ped_MV_Minor | r_Distance | unilateral | 125195\G-A194\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Annul Diam | Dist | meas | m_ped_MV_Annul_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-3500E\T-35300 | SRT, T-35300, Mitral Valve | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-3500E, Cardiac valve annulus | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV DFP | Time | meas | m_ped_MV_DFP | r_DopTime | unilateral | 125195\R-0032C\T-35300 | SRT, T-35300, Mitral Valve | SRT, R-0032C, Mitral Diastolic Filling Period (DFPm) |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT VTI | Vmean | meas | m_ped_RVOT_VTI | r_CardTAPV | unilateral | 125195\11692-1\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 11692-1, Time Averaged Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Accel | Slope | meas | m_p_RVOT_Accel | r_DopSlope | unilateral | 125195\20167-3\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20167-3, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Accel | Time | meas | m_p_RVOT_Accel | r_DopTime | unilateral | 125195\20168-1\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Max PG | Max PG | meas | m_ped_RVOT_MaxPG | r_MaxPG | unilateral | 125195\20247-3\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Vmax | Max PG | meas | m_ped_RVOT_Vmax | r_MaxPG | unilateral | 125195\125214\20247-3\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-32550, Right Ventricle Outflow Tract |  |  |  |  | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Mean PG | Mean PG | meas | m_ped_RVOT_PG | r_CardPGmean | unilateral | 125195\20256-4\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Max PG | Vmax | meas | m_ped_RVOT_MaxPG | r_Vmax | unilateral | 125195\20351-3\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Vmax | Vmax | meas | m_ped_RVOT_Vmax | r_Vmax | unilateral | 125195\125214\20351-3\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, T-32550, Right Ventricle Outflow Tract |  |  |  |  | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT VTI | VTI | meas | m_ped_RVOT_VTI | r_VTI | unilateral | 125195\20354-7\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Diam | Dist | meas | m_ped_RVOT_Diam | r_Distance | unilateral | 125195\G-038F\G-03A2\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, G-038F, Cardiovascular Orifice Diameter |  |  | UCUM, mm, mm | SRT, T-32550, Right Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT VTI | Vmean | meas | m_ped_LVOT_VTI | r_CardTAPV | unilateral | 125195\11692-1\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 11692-1, Time Averaged Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Max PG | Max PG | meas | m_ped_LVOT_MaxPG | r_MaxPG | unilateral | 125195\20247-3\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Vmax | Max PG | meas | m_ped_LVOT_Vmax | r_MaxPG | unilateral | 125195\125214\20247-3\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-32650, Left Ventricle Outflow Tract |  |  |  |  | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Mean PG | Mean PG | meas | m_ped_LVOT_PG | r_CardPGmean | unilateral | 125195\20256-4\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Max PG | Vmax | meas | m_ped_LVOT_MaxPG | r_Vmax | unilateral | 125195\20351-3\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Vmax | Vmax | meas | m_ped_LVOT_Vmax | r_Vmax | unilateral | 125195\125214\20351-3\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, T-32650, Left Ventricle Outflow Tract |  |  |  |  | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT VTI | VTI | meas | m_ped_LVOT_VTI | r_VTI | unilateral | 125195\20354-7\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Diam | Dist | meas | m_ped_LVOT_Diam | r_Distance | unilateral | 125195\G-038F\G-03A2\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, G-038F, Cardiovascular Orifice Diameter |  |  | UCUM, mm, mm | SRT, T-32650, Left Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RPA Max PG | Max PG | meas | m_ped_RPA_MaxPG | r_MaxPG | unilateral | 125195\20247-3\T-44000\T-44200 | SRT, T-44000, Pulmonary artery | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-44200, Right Pulmonary Artery
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LPA Max PG | Max PG | meas | m_ped_LPA_MaxPG | r_MaxPG | unilateral | 125195\20247-3\T-44000\T-44400 | SRT, T-44000, Pulmonary artery | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-44400, Left Pulmonary Artery
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RPA Mean PG | Mean PG | meas | m_ped_RPA_PG | r_CardPGmean | unilateral | 125195\20256-4\T-44000\T-44200 | SRT, T-44000, Pulmonary artery | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-44200, Right Pulmonary Artery
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LPA Mean PG | Mean PG | meas | m_ped_LPA_PG | r_CardPGmean | unilateral | 125195\20256-4\T-44000\T-44400 | SRT, T-44000, Pulmonary artery | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg | SRT, T-44400, Left Pulmonary Artery
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RPA Max PG | Vmax | meas | m_ped_RPA_MaxPG | r_Vmax | unilateral | 125195\20351-3\T-44000\T-44200 | SRT, T-44000, Pulmonary artery | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, T-44200, Right Pulmonary Artery
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LPA Max PG | Vmax | meas | m_ped_LPA_MaxPG | r_Vmax | unilateral | 125195\20351-3\T-44000\T-44400 | SRT, T-44000, Pulmonary artery | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, T-44400, Left Pulmonary Artery
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RPA Diam | Dist | meas | m_ped_RPA_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-44000\T-44200 | SRT, T-44000, Pulmonary artery | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-44200, Right Pulmonary Artery | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LPA Diam | Dist | meas | m_ped_LPA_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-44000\T-44400 | SRT, T-44000, Pulmonary artery | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-44400, Left Pulmonary Artery | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PA Accel Time | Slope | meas | m_ped_PA_AccelTime | r_DopSlope | unilateral | 125195\20167-3\T-44100 | SRT, T-44100, Pulmonary Trunk | LN, 20167-3, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PA Accel Time | Time | meas | m_ped_PA_AccelTime | r_DopTime | unilateral | 125195\20168-1\T-44100 | SRT, T-44100, Pulmonary Trunk | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MPA Max PG | Max PG | meas | m_ped_MPA_MaxPG | r_MaxPG | unilateral | 125195\20247-3\T-44100 | SRT, T-44100, Pulmonary Trunk | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MPA Mean PG | Mean PG | meas | m_ped_MPA_PG | r_CardPGmean | unilateral | 125195\20256-4\T-44100 | SRT, T-44100, Pulmonary Trunk | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MPA Max PG | Vmax | meas | m_ped_MPA_MaxPG | r_Vmax | unilateral | 125195\20351-3\T-44100 | SRT, T-44100, Pulmonary Trunk | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MPA Diam | Dist | meas | m_ped_MPA_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-44100 | SRT, T-44100, Pulmonary Trunk | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RV ET | Time | meas | m_ped_RV_Eject_Tm | r_MMTime | unilateral | 125195\20222-6\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 20222-6, Ejection Time |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PI P½t | Vmax | meas | m_ped_PI_PHalfT | r_Vmax | unilateral | 125195\C12222-03\G-0367\T-35200 | SRT, T-35200, Pulmonic Valve | 99PMSBLUS, C12222-03, Pressure Half-Time Peak velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PI End Dias Vel | Vel | meas | m_ped_PI_EndDiVel | r_Vel | unilateral | 125195\11653-3\G-0367\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV VTI | Vmean | meas | m_ped_PV_VTI | r_CardTAPV | unilateral | 125195\11692-1\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 11692-1, Time Averaged Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV Accel | Slope | meas | m_ped_PV_Accel | r_DopSlope | unilateral | 125195\20167-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20167-3, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV Accel | Time | meas | m_ped_PV_Accel | r_DopTime | unilateral | 125195\20168-1\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PI Decel Slope | Slope | meas | m_ped_PI_DecelSlope | r_DopSlope | unilateral | 125195\20216-8\G-0367\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20216-8, Deceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Late Dias Slope | Slope | meas | m_ped_LateDiasSlope | r_MMSlope | unilateral | 125195\20216-8\F-32011\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20216-8, Deceleration Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PI Decel Time | Time | meas | m_ped_PI_DecelTime | r_DopTime | unilateral | 125195\20217-6\G-0367\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV Max PG | Max PG | meas | m_ped_PV_MaxPG | r_MaxPG | unilateral | 125195\20247-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV Vmax | Max PG | meas | m_ped_PV_Vmax | r_MaxPG | unilateral | 125195\125214\20247-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV Mean PG | Mean PG | meas | m_ped_PV_PG | r_CardPGmean | unilateral | 125195\20256-4\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PI P½t | P½t | meas | m_ped_PI_PHalfT | r_PHalftime | unilateral | 125195\20280-4\G-0367\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20280-4, Pressure Half-Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV Max PG | Vmax | meas | m_ped_PV_MaxPG | r_Vmax | unilateral | 125195\20351-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV Vmax | Vmax | meas | m_ped_PV_Vmax | r_Vmax | unilateral | 125195\125214\20351-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV VTI | VTI | meas | m_ped_PV_VTI | r_VTI | unilateral | 125195\20354-7\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A Wave Amp | Dist | meas | m_ped_AWaveAmp | r_Distance | unilateral | 125195\59100-8\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 59100-8, A-Wave Amplitude |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | B-C Slope | Slope | meas | m_ped_BC_Slope | r_MMSlope | unilateral | 125195\59126-3\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 59126-3, B-C Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | HR - PV | HR | meas | m_ped_HR_PV | r_HeartRate | unilateral | 125195\8867-4\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV Annul Diam | Dist | meas | m_ped_PV_Annul_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-3500E\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-3500E, Cardiac valve annulus | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI End Dias Vel | Vel | meas | m_ped_AI_EndDiVel | r_Vel | unilateral | 125195\11653-3\F-32011\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, F-32011, End Diastole | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI VTI | Vmean | meas | m_ped_AI_VTI | r_CardTAPV | unilateral | 125195\11692-1\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 11692-1, Time Averaged Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV VTI | Vmean | meas | m_ped_AV_VTI | r_CardTAPV | unilateral | 125195\11692-1\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 11692-1, Time Averaged Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Cusp Sep | Dist | meas | m_ped_AV_Cusp_Sep | r_Distance | unilateral | 125195\17996-0\G-0394\T-35400 | SRT, T-35400, Aortic Valve | LN, 17996-0, Aortic Valve Cusp Separation |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Accel | Slope | meas | m_ped_AI_Accel | r_DopSlope | unilateral | 125195\20167-3\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20167-3, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Accel | Slope | meas | m_ped_AV_Accel | r_DopSlope | unilateral | 125195\20167-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20167-3, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Accel | Time | meas | m_ped_AI_Accel | r_DopTime | unilateral | 125195\20168-1\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Accel | Time | meas | m_ped_AV_Accel | r_DopTime | unilateral | 125195\20168-1\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Decel Slope | Slope | meas | m_ped_AI_DecelSlope | r_DopSlope | unilateral | 125195\20216-8\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20216-8, Deceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Decel Time | Time | meas | m_ped_AI_DecelTime | r_DopTime | unilateral | 125195\20217-6\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Vmax | Max PG | meas | m_ped_AI_Vmax | r_MaxPG | unilateral | 125195\20247-3\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Max PG | Max PG | meas | m_ped_AV_MaxPG | r_MaxPG | unilateral | 125195\20247-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Vmax | Max PG | meas | m_ped_AV_Vmax | r_MaxPG | unilateral | 125195\125214\20247-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI VTI | Mean PG | meas | m_ped_AI_VTI | r_CardPGmean | unilateral | 125195\20256-4\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Mean PG | Mean PG | meas | m_ped_AV_PG | r_CardPGmean | unilateral | 125195\20256-4\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI P½t | P½t | meas | m_ped_AI_PHalfT | r_PHalftime | unilateral | 125195\20280-4\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20280-4, Pressure Half-Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Vmax | Vmax | meas | m_ped_AI_Vmax | r_Vmax | unilateral | 125195\20351-3\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Max PG | Vmax | meas | m_ped_AV_MaxPG | r_Vmax | unilateral | 125195\20351-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Vmax | Vmax | meas | m_ped_AV_Vmax | r_Vmax | unilateral | 125195\125214\20351-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI VTI | VTI | meas | m_ped_AI_VTI | r_VTI | unilateral | 125195\20354-7\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV VTI | VTI | meas | m_ped_AV_VTI | r_VTI | unilateral | 125195\20354-7\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Radius | Dist | meas | m_ped_AI_Radius | r_Distance | unilateral | 125195\59102-4\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 59102-4, Flow Radius |  |  | UCUM, mm, mm |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AS Vmax | Max PG | meas | m_ped_AS_Vmax | r_MaxPG | unilateral | 125195\59106-5\T-35400 | SRT, T-35400, Aortic Valve | LN, 59106-5, Stenosis Peak Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AS Vmax | Vmax | meas | m_ped_AS_Vmax | r_Vmax | unilateral | 125195\59107-3\T-35400 | SRT, T-35400, Aortic Valve | LN, 59107-3, Stenosis Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Alias Vel | Alias Vel | meas | m_ped_AI_AliasVel | r_2DVel | unilateral | 125195\59130-5\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 59130-5, Alias velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | HR - AV | HR | meas | m_ped_HR_AV | r_HeartRate | unilateral | 125195\8867-4\T-35400 | SRT, T-35400, Aortic Valve | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Area | Area | meas | m_ped_AV_Area | r_Area | unilateral | 125195\G-03A2\G-A166\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Annul Diam | Dist | meas | m_ped_AV_Annul_Diam | r_Distance | unilateral | 125195\G-03A2\M-02550\T-35400\T-35410 | SRT, T-35400, Aortic Valve | SRT, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-35410, Aortic Valve Ring | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RV ET | Time | meas | m_ped_RV_ET_Dop | r_DopTime | unilateral | 125195\20222-6\T-32500 | SRT, T-32500, Right Ventricle | LN, 20222-6, Ejection Time |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RV PEP | Time | meas | m_ped_RV_PreEject_Tm | r_MMTime | unilateral | 125195\59085-1\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 59085-1, Pre-Ejection Period |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVAWd (MM) | Dist | meas | m_ped_RVAWd_MM | r_Distance | unilateral | 125195\59089-3\F-32011\G-0394\R-4210B\T-32500 | SRT, T-32500, Right Ventricle | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm | SRT, R-4210B, Anterior Wall | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVAWd (2D) | Dist | meas | m_ped_RVAWd_2D | r_Distance | unilateral | 125195\59089-3\F-32011\G-03A2\R-4210B\T-32500 | SRT, T-32500, Right Ventricle | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm | SRT, R-4210B, Anterior Wall | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVIDd (MM) | Dist | meas | m_ped_RVIDd_MM | r_Distance | unilateral | 125195\59090-1\F-32011\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 59090-1, ROI Internal Dimension by US |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVIDd (2D) | Dist | meas | m_ped_RVIDd_2D | r_Distance | unilateral | 125195\59090-1\F-32011\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | LN, 59090-1, ROI Internal Dimension by US |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVed Major - A4C | Dist | meas | m_ped_RVED_MAJOR_A4C | r_Distance | unilateral | 125195\F-32011\G-A193\G-A19C\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVes Major - A4C | Dist | meas | m_ped_RVES_MAJOR_A4C | r_Distance | unilateral | 125195\G-A193\G-A19C\R-FAB5B\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVed Minor - A4C | Dist | meas | m_ped_RVED_MINOR_A4C | r_Distance | unilateral | 125195\F-32011\G-A194\G-A19C\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVes Minor - A4C | Dist | meas | m_ped_RVES_MINOR_A4C | r_Distance | unilateral | 125195\G-A194\G-A19C\R-FAB5B\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVed - SAX CH | Dist | meas | m_ped_LVED_SAX_CH | r_Distance | unilateral | 125195\121206\F-32011\G-0399\T-32600 | SRT, T-32600, Left Ventricle | DCM, 121206, Distance |  |  | UCUM, mm, mm |  |  | SRT, G-0399, Parasternal short axis at the level of the mitral chords | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVes - SAX CH | Dist | meas | m_ped_LVES_SAX_CH | r_Distance | unilateral | 125195\121206\G-0399\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | DCM, 121206, Distance |  |  | UCUM, mm, mm |  |  | SRT, G-0399, Parasternal short axis at the level of the mitral chords | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVed - SAX PM | Dist | meas | m_ped_LVED_SAX_PM | r_Distance | unilateral | 125195\121206\F-32011\G-039B\T-32600 | SRT, T-32600, Left Ventricle | DCM, 121206, Distance |  |  | UCUM, mm, mm |  |  | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVes - SAX PM | Dist | meas | m_ped_LVES_SAX_PM | r_Distance | unilateral | 125195\121206\G-039B\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | DCM, 121206, Distance |  |  | UCUM, mm, mm |  |  | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Accel | Slope | meas | m_p_LVOT_Accel | r_DopSlope | unilateral | 125195\20167-3\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20167-3, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med Accel Time | Time | meas | m_p_TDIMedAccelTime | r_DopTime | unilateral | 125195\20168-1\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat Accel Time | Time | meas | m_p_TDILatAccelTime | r_DopTime | unilateral | 125195\20168-1\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Accel | Time | meas | m_p_LVOT_Accel | r_DopTime | unilateral | 125195\20168-1\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med Decel Time | Time | meas | m_p_TDIMedDecelTime | r_DopTime | unilateral | 125195\20217-6\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat Decel Time | Time | meas | m_p_TDILatDecelTime | r_DopTime | unilateral | 125195\20217-6\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV ET | Time | meas | m_ped_LV_ET_Dop | r_DopTime | unilateral | 125195\20222-6\T-32600 | SRT, T-32600, Left Ventricle | LN, 20222-6, Ejection Time |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV ET | Time | meas | m_ped_LV_Eject_Tm | r_MMTime | unilateral | 125195\20222-6\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 20222-6, Ejection Time |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVRT | Time | meas | m_ped_IVRT | r_DopTime | unilateral | 125195\59083-6\T-32600 | SRT, T-32600, Left Ventricle | LN, 59083-6, Isovolumic Relaxation Time |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med IVRT | Time | meas | m_p_TDI_Med_IVRT | r_DopTime | unilateral | 125195\59083-6\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59083-6, Isovolumic Relaxation Time |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat IVRT | Time | meas | m_p_TDI_Lat_IVRT | r_DopTime | unilateral | 125195\59083-6\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59083-6, Isovolumic Relaxation Time |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVCT | Time | meas | m_ped_IVCT | r_DopTime | unilateral | 125195\59084-4\T-32600 | SRT, T-32600, Left Ventricle | LN, 59084-4, Isovolumic Contraction Time |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med IVCT | Time | meas | m_p_TDI_Med_IVCT | r_DopTime | unilateral | 125195\59084-4\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59084-4, Isovolumic Contraction Time |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat IVCT | Time | meas | m_p_TDI_Lat_IVCT | r_DopTime | unilateral | 125195\59084-4\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59084-4, Isovolumic Contraction Time |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV PEP | Time | meas | m_ped_LV_PreEject_Tm | r_MMTime | unilateral | 125195\59085-1\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59085-1, Pre-Ejection Period |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVPWd (MM) | Dist | meas | m_ped_LVPWd_MM | r_Distance | unilateral | 125195\59089-3\F-32011\G-0394\R-42175\T-32600 | SRT, T-32600, Left Ventricle | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm | SRT, R-42175, Posterior Wall | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVPWs (MM) | Dist | meas | m_ped_LVPWs_MM | r_Distance | unilateral | 125195\59089-3\G-0394\R-42175\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm | SRT, R-42175, Posterior Wall | SRT, G-0394, M mode |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVPWd (2D) | Dist | meas | m_ped_LVPWd_2D | r_Distance | unilateral | 125195\59089-3\F-32011\G-03A2\R-42175\T-32600 | SRT, T-32600, Left Ventricle | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm | SRT, R-42175, Posterior Wall | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVPWs (2D) | Dist | meas | m_ped_LVPWs_2D | r_Distance | unilateral | 125195\59089-3\G-03A2\R-42175\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm | SRT, R-42175, Posterior Wall | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVIDd (MM) | Dist | meas | m_ped_LVIDd_MM | r_Distance | unilateral | 125195\59090-1\F-32011\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59090-1, ROI Internal Dimension by US |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVIDs (MM) | Dist | meas | m_ped_LVIDs_MM | r_Distance | unilateral | 125195\59090-1\G-0394\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59090-1, ROI Internal Dimension by US |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVIDd (2D) | Dist | meas | m_ped_LVIDd_2D | r_Distance | unilateral | 125195\59090-1\F-32011\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 59090-1, ROI Internal Dimension by US |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVIDs (2D) | Dist | meas | m_ped_LVIDs_2D | r_Distance | unilateral | 125195\59090-1\G-03A2\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59090-1, ROI Internal Dimension by US |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVAd Sax Epi | Area | meas | m_ped_LV_AdSaxEpi | r_Area | unilateral | 125195\59093-5\F-32011\G-039B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59093-5, Epicardial Area |  |  | UCUM, mm2, mm2 |  |  | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVAd Sax Endo | Area | meas | m_ped_LV_AdSaxEndo | r_Area | unilateral | 125195\59094-3\F-32011\G-039B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59094-3, Endocardial Area |  |  | UCUM, mm2, mm2 |  |  | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Time to Med S | Time | meas | m_p_TimeMedS | r_DopTime | unilateral | 125195\59095-0\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59095-0, Time to Left Ventricle S Tissue Velocity |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Time to Lat S | Time | meas | m_p_TimeLatS | r_DopTime | unilateral | 125195\59095-0\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59095-0, Time to Left Ventricle S Tissue Velocity |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Time to Med E` | Time | meas | m_p_TimeMedE | r_DopTime | unilateral | 125195\59096-8\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59096-8, Time to Left Ventricle E Tissue Velocity |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Time to Lat E` | Time | meas | m_p_TimeLatE | r_DopTime | unilateral | 125195\59096-8\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59096-8, Time to Left Ventricle E Tissue Velocity |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med E` Area | VTI | meas | m_p_AreaUnderMedE | r_VTI | unilateral | 125195\59124-8\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59124-8, Tissue Velocity Time Integral (VTI) for the area under Left Ventricle E wave |  |  | UCUM, mm, mm | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat E` Area | VTI | meas | m_p_AreaUnderLatE | r_VTI | unilateral | 125195\59124-8\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59124-8, Tissue Velocity Time Integral (VTI) for the area under Left Ventricle E wave |  |  | UCUM, mm, mm | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med A` Area | VTI | meas | m_p_AreaUnderMedA | r_VTI | unilateral | 125195\59125-5\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59125-5, Tissue Velocity Time Integral (VTI) for the area under Left Ventricle A wave |  |  | UCUM, mm, mm | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat A` Area | VTI | meas | m_p_AreaUnderLatA | r_VTI | unilateral | 125195\59125-5\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59125-5, Tissue Velocity Time Integral (VTI) for the area under Left Ventricle A wave |  |  | UCUM, mm, mm | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med E` Vel | Vel | meas | m_p_Med_EVel | r_Vel | unilateral | 125195\59133-9\G-0391\P5-B0128\R-40B1B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59133-9, Peak Tissue Velocity |  |  | UCUM, mm/s, mm/s | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging |  | SRT, R-40B1B, Early Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med S Vel | Vel | meas | m_p_Med_SVel | r_Vel | unilateral | 125195\59133-9\G-0391\P5-B0128\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59133-9, Peak Tissue Velocity |  |  | UCUM, mm/s, mm/s | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat E` Vel | Vel | meas | m_p_Lat_EVel | r_Vel | unilateral | 125195\59133-9\G-0392\P5-B0128\R-40B1B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59133-9, Peak Tissue Velocity |  |  | UCUM, mm/s, mm/s | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging |  | SRT, R-40B1B, Early Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat S Vel | Vel | meas | m_p_Lat_SVel | r_Vel | unilateral | 125195\59133-9\G-0392\P5-B0128\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59133-9, Peak Tissue Velocity |  |  | UCUM, mm/s, mm/s | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | HR - LV | HR | meas | m_ped_HR_LV | r_HeartRate | unilateral | 125195\8867-4\T-32600 | SRT, T-32600, Left Ventricle | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med A` Vel | Vel | meas | m_p_Med_AVel | r_Vel | unilateral | 125195\F-32011\G-037C\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037C, LV Peak Diastolic Tissue Velocity During Atrial Systole |  |  | UCUM, mm/s, mm/s | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat A` Vel | Vel | meas | m_p_Lat_AVel | r_Vel | unilateral | 125195\F-32011\G-037C\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037C, LV Peak Diastolic Tissue Velocity During Atrial Systole |  |  | UCUM, mm/s, mm/s | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | LV Area | meas | m_ped_A2Cd | r_SimpArea | unilateral | 125195\125208\F-32011\G-03A2\G-A166\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | LV Area | meas | m_ped_A4Cd | r_SimpArea | unilateral | 125195\125208\F-32011\G-03A2\G-A166\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | LV Area | meas | m_ped_A2Cs | r_SimpArea | unilateral | 125195\125208\G-03A2\G-A166\G-A19B\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | LV Area | meas | m_ped_A4Cs | r_SimpArea | unilateral | 125195\125208\G-03A2\G-A166\G-A19C\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVd - A2C (A/L) | Area | meas | m_ped_LVd_A2C_SPE | r_Area | unilateral | 125195\125226\F-32011\G-03A2\G-A166\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVs - A2C (A/L) | Area | meas | m_ped_LVs_A2C_SPE | r_Area | unilateral | 125195\125226\G-03A2\G-A166\G-A19B\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVd - A4C (A/L) | Area | meas | m_ped_LVd_A4C_SPE | r_Area | unilateral | 125195\125226\F-32011\G-03A2\G-A166\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVs - A4C (A/L) | Area | meas | m_ped_LVs_A4C_SPE | r_Area | unilateral | 125195\125226\G-03A2\G-A166\G-A19C\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | LV Length | meas | m_ped_A2Cd | r_SimpDistance | unilateral | 125195\125208\F-32011\G-03A2\G-A193\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | LV Length | meas | m_ped_A4Cd | r_SimpDistance | unilateral | 125195\125208\F-32011\G-03A2\G-A193\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | LV Length | meas | m_ped_A2Cs | r_SimpDistance | unilateral | 125195\125208\G-03A2\G-A193\G-A19B\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | LV Length | meas | m_ped_A4Cs | r_SimpDistance | unilateral | 125195\125208\G-03A2\G-A193\G-A19C\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVd - A2C (A/L) | Length | meas | m_ped_LVd_A2C_SPE | r_Length | unilateral | 125195\125226\F-32011\G-03A2\G-A193\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVs - A2C (A/L) | Length | meas | m_ped_LVs_A2C_SPE | r_Length | unilateral | 125195\125226\G-03A2\G-A193\G-A19B\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVd - A4C (A/L) | Length | meas | m_ped_LVd_A4C_SPE | r_Length | unilateral | 125195\125226\F-32011\G-03A2\G-A193\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVs - A4C (A/L) | Length | meas | m_ped_LVs_A4C_SPE | r_Length | unilateral | 125195\125226\G-03A2\G-A193\G-A19C\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVLd Apical | Dist | meas | m_ped_LV_LdApical | r_Distance | unilateral | 125195\F-32011\G-A193\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVed Major - A4C | Dist | meas | m_ped_LVED_MAJOR_A4C | r_Distance | unilateral | 125195\F-32011\G-A193\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVes Major - A4C | Dist | meas | m_ped_LVES_MAJOR_A4C | r_Distance | unilateral | 125195\G-A193\G-A19C\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVed Minor - A4C | Dist | meas | m_ped_LVED_MINOR_A4C | r_Distance | unilateral | 125195\F-32011\G-A194\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVes Minor - A4C | Dist | meas | m_ped_LVES_MINOR_A4C | r_Distance | unilateral | 125195\G-A194\G-A19C\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | LV Vol | meas | m_ped_A2Cd | r_SimpVolume | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | LV Vol | meas | m_ped_A4Cd | r_SimpVolume | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | LV Vol | meas | m_ped_A2Cs | r_SimpVolume | unilateral | 125195\125208\G-03A2\G-A19B\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | LV Vol | meas | m_ped_A4Cs | r_SimpVolume | unilateral | 125195\125208\G-03A2\G-A19C\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 1 | meas | m_ped_A2Cd | r_Seg1 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 10 | meas | m_ped_A2Cd | r_Seg10 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 11 | meas | m_ped_A2Cd | r_Seg11 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 12 | meas | m_ped_A2Cd | r_Seg12 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 13 | meas | m_ped_A2Cd | r_Seg13 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 14 | meas | m_ped_A2Cd | r_Seg14 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 15 | meas | m_ped_A2Cd | r_Seg15 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 16 | meas | m_ped_A2Cd | r_Seg16 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 17 | meas | m_ped_A2Cd | r_Seg17 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 18 | meas | m_ped_A2Cd | r_Seg18 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 19 | meas | m_ped_A2Cd | r_Seg19 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 2 | meas | m_ped_A2Cd | r_Seg2 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 20 | meas | m_ped_A2Cd | r_Seg20 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 3 | meas | m_ped_A2Cd | r_Seg3 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 4 | meas | m_ped_A2Cd | r_Seg4 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 5 | meas | m_ped_A2Cd | r_Seg5 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 6 | meas | m_ped_A2Cd | r_Seg6 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 7 | meas | m_ped_A2Cd | r_Seg7 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 8 | meas | m_ped_A2Cd | r_Seg8 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cd | Seg 9 | meas | m_ped_A2Cd | r_Seg9 | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 1 | meas | m_ped_A2Cs | r_Seg1 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 10 | meas | m_ped_A2Cs | r_Seg10 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 11 | meas | m_ped_A2Cs | r_Seg11 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 12 | meas | m_ped_A2Cs | r_Seg12 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 13 | meas | m_ped_A2Cs | r_Seg13 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 14 | meas | m_ped_A2Cs | r_Seg14 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 15 | meas | m_ped_A2Cs | r_Seg15 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 16 | meas | m_ped_A2Cs | r_Seg16 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 17 | meas | m_ped_A2Cs | r_Seg17 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 18 | meas | m_ped_A2Cs | r_Seg18 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 19 | meas | m_ped_A2Cs | r_Seg19 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 2 | meas | m_ped_A2Cs | r_Seg2 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 20 | meas | m_ped_A2Cs | r_Seg20 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 3 | meas | m_ped_A2Cs | r_Seg3 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 4 | meas | m_ped_A2Cs | r_Seg4 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 5 | meas | m_ped_A2Cs | r_Seg5 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 6 | meas | m_ped_A2Cs | r_Seg6 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 7 | meas | m_ped_A2Cs | r_Seg7 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 8 | meas | m_ped_A2Cs | r_Seg8 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A2Cs | Seg 9 | meas | m_ped_A2Cs | r_Seg9 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 1 | meas | m_ped_A4Cd | r_Seg1 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 10 | meas | m_ped_A4Cd | r_Seg10 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 11 | meas | m_ped_A4Cd | r_Seg11 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 12 | meas | m_ped_A4Cd | r_Seg12 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 13 | meas | m_ped_A4Cd | r_Seg13 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 14 | meas | m_ped_A4Cd | r_Seg14 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 15 | meas | m_ped_A4Cd | r_Seg15 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 16 | meas | m_ped_A4Cd | r_Seg16 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 17 | meas | m_ped_A4Cd | r_Seg17 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 18 | meas | m_ped_A4Cd | r_Seg18 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 19 | meas | m_ped_A4Cd | r_Seg19 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 2 | meas | m_ped_A4Cd | r_Seg2 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 20 | meas | m_ped_A4Cd | r_Seg20 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 3 | meas | m_ped_A4Cd | r_Seg3 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 4 | meas | m_ped_A4Cd | r_Seg4 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 5 | meas | m_ped_A4Cd | r_Seg5 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 6 | meas | m_ped_A4Cd | r_Seg6 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 7 | meas | m_ped_A4Cd | r_Seg7 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 8 | meas | m_ped_A4Cd | r_Seg8 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cd | Seg 9 | meas | m_ped_A4Cd | r_Seg9 | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 1 | meas | m_ped_A4Cs | r_Seg1 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 10 | meas | m_ped_A4Cs | r_Seg10 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 11 | meas | m_ped_A4Cs | r_Seg11 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 12 | meas | m_ped_A4Cs | r_Seg12 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 13 | meas | m_ped_A4Cs | r_Seg13 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 14 | meas | m_ped_A4Cs | r_Seg14 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 15 | meas | m_ped_A4Cs | r_Seg15 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 16 | meas | m_ped_A4Cs | r_Seg16 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 17 | meas | m_ped_A4Cs | r_Seg17 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 18 | meas | m_ped_A4Cs | r_Seg18 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 19 | meas | m_ped_A4Cs | r_Seg19 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 2 | meas | m_ped_A4Cs | r_Seg2 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 20 | meas | m_ped_A4Cs | r_Seg20 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 3 | meas | m_ped_A4Cs | r_Seg3 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 4 | meas | m_ped_A4Cs | r_Seg4 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 5 | meas | m_ped_A4Cs | r_Seg5 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 6 | meas | m_ped_A4Cs | r_Seg6 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 7 | meas | m_ped_A4Cs | r_Seg7 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 8 | meas | m_ped_A4Cs | r_Seg8 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A4Cs | Seg 9 | meas | m_ped_A4Cs | r_Seg9 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV dP/dt | dP/dt | meas | m_ped_LV_dPdt | r_DopDpDt | unilateral | 125195\18035-6\T-35300 | SRT, T-35300, Mitral Valve | LN, 18035-6, Mitral Regurgitation dP/dt derived from Mitral Reg. velocity |  |  | UCUM, mmHg, mmHg
Pediatric Cardiac Ultrasound Report | Pediatric Echo | VSD Vmax | Max PG | meas | m_p_VSD_Vmax | r_MaxPG | unilateral | 125195\20247-3\D4-31150 | SRT, D4-31150, Ventricular Septal Defect | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Cardiac Ultrasound Report | Pediatric Echo | VSD Vmax | Vmax | meas | m_p_VSD_Vmax | r_Vmax | unilateral | 125195\20351-3\D4-31150 | SRT, D4-31150, Ventricular Septal Defect | LN, 20351-3, Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Cardiac Ultrasound Report | Pediatric Echo | VSD Major | Dist | meas | m_ped_VSD_Major | r_Distance | unilateral | 125195\D4-31150\G-03A2\G-A193 | SRT, D4-31150, Ventricular Septal Defect | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | VSD Minor | Dist | meas | m_ped_VSD_Minor | r_Distance | unilateral | 125195\D4-31150\G-03A2\G-A194 | SRT, D4-31150, Ventricular Septal Defect | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVSd (MM) | Dist | meas | m_ped_IVSd_MM | r_Distance | unilateral | 125195\59089-3\F-32011\G-0394\T-32410 | SRT, T-32410, Interventricular septum | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVSs (MM) | Dist | meas | m_ped_IVSs_MM | r_Distance | unilateral | 125195\59089-3\G-0394\R-FAB5B\T-32410 | SRT, T-32410, Interventricular septum | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVSd (2D) | Dist | meas | m_ped_IVSd_2D | r_Distance | unilateral | 125195\59089-3\F-32011\G-03A2\T-32410 | SRT, T-32410, Interventricular septum | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVSs (2D) | Dist | meas | m_ped_IVSs_2D | r_Distance | unilateral | 125195\59089-3\G-03A2\R-FAB5B\T-32410 | SRT, T-32410, Interventricular septum | LN, 59089-3, ROI Thickness by US |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2C-A/L | Area | meas | m_ped_A2Cal_RA | r_Area | unilateral | 125195\125226\G-03A2\G-A166\G-A19B\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Atrial Area | meas | m_ped_A2Cs_RA | r_ASimpArea | unilateral | 125195\125208\G-03A2\G-A166\G-A19B\R-FAB5B\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4C-A/L | Area | meas | m_ped_A4Cal_RA | r_Area | unilateral | 125195\125226\G-03A2\G-A166\G-A19C\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Atrial Area | meas | m_ped_A4Cs_RA | r_ASimpArea | unilateral | 125195\125208\G-03A2\G-A166\G-A19C\R-FAB5B\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2C-A/L | Length | meas | m_ped_A2Cal_RA | r_Length | unilateral | 125195\125226\G-03A2\G-A193\G-A19B\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Atrial Length | meas | m_ped_A2Cs_RA | r_ASimpDistance | unilateral | 125195\125208\G-03A2\G-A193\G-A19B\R-FAB5B\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4C-A/L | Length | meas | m_ped_A4Cal_RA | r_Length | unilateral | 125195\125226\G-03A2\G-A193\G-A19C\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Atrial Length | meas | m_ped_A4Cs_RA | r_ASimpDistance | unilateral | 125195\125208\G-03A2\G-A193\G-A19C\R-FAB5B\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Atrial Volume | meas | m_ped_A2Cs_RA | r_ASimpVolume | unilateral | 125195\125208\G-03A2\G-A19B\G-D705\R-FAB5B\T-32200 | SRT, T-32200, Right Atrium | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Atrial Volume | meas | m_ped_A4Cs_RA | r_ASimpVolume | unilateral | 125195\125208\G-03A2\G-A19C\G-D705\R-FAB5B\T-32200 | SRT, T-32200, Right Atrium | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 1 | meas | m_ped_A2Cs_RA | r_Seg1 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 10 | meas | m_ped_A2Cs_RA | r_Seg10 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 11 | meas | m_ped_A2Cs_RA | r_Seg11 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 12 | meas | m_ped_A2Cs_RA | r_Seg12 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 13 | meas | m_ped_A2Cs_RA | r_Seg13 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 14 | meas | m_ped_A2Cs_RA | r_Seg14 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 15 | meas | m_ped_A2Cs_RA | r_Seg15 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 16 | meas | m_ped_A2Cs_RA | r_Seg16 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 17 | meas | m_ped_A2Cs_RA | r_Seg17 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 18 | meas | m_ped_A2Cs_RA | r_Seg18 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 19 | meas | m_ped_A2Cs_RA | r_Seg19 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 2 | meas | m_ped_A2Cs_RA | r_Seg2 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 20 | meas | m_ped_A2Cs_RA | r_Seg20 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 3 | meas | m_ped_A2Cs_RA | r_Seg3 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 4 | meas | m_ped_A2Cs_RA | r_Seg4 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 5 | meas | m_ped_A2Cs_RA | r_Seg5 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 6 | meas | m_ped_A2Cs_RA | r_Seg6 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 7 | meas | m_ped_A2Cs_RA | r_Seg7 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 8 | meas | m_ped_A2Cs_RA | r_Seg8 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A2Cs | Seg 9 | meas | m_ped_A2Cs_RA | r_Seg9 | unilateral | 125195\125208\G-03A2\G-A19B\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 1 | meas | m_ped_A4Cs_RA | r_Seg1 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 10 | meas | m_ped_A4Cs_RA | r_Seg10 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 11 | meas | m_ped_A4Cs_RA | r_Seg11 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 12 | meas | m_ped_A4Cs_RA | r_Seg12 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 13 | meas | m_ped_A4Cs_RA | r_Seg13 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 14 | meas | m_ped_A4Cs_RA | r_Seg14 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 15 | meas | m_ped_A4Cs_RA | r_Seg15 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 16 | meas | m_ped_A4Cs_RA | r_Seg16 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 17 | meas | m_ped_A4Cs_RA | r_Seg17 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 18 | meas | m_ped_A4Cs_RA | r_Seg18 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 19 | meas | m_ped_A4Cs_RA | r_Seg19 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 2 | meas | m_ped_A4Cs_RA | r_Seg2 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 20 | meas | m_ped_A4Cs_RA | r_Seg20 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 3 | meas | m_ped_A4Cs_RA | r_Seg3 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 4 | meas | m_ped_A4Cs_RA | r_Seg4 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 5 | meas | m_ped_A4Cs_RA | r_Seg5 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 6 | meas | m_ped_A4Cs_RA | r_Seg6 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 7 | meas | m_ped_A4Cs_RA | r_Seg7 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 8 | meas | m_ped_A4Cs_RA | r_Seg8 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA A4Cs | Seg 9 | meas | m_ped_A4Cs_RA | r_Seg9 | unilateral | 125195\125208\G-03A2\G-A19C\M-02550\R-FAB5B\T-32200\T5203-01 | SRT, T-32200, Right Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Atrial Area | meas | m_p_A2Cs_LA | r_ASimpArea | unilateral | 125195\125208\G-03A2\G-A166\G-A19B\R-FAB5B\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2C-A/L | Area | meas | m_ped_A2Cal_LA | r_Area | unilateral | 125195\125226\G-03A2\G-A166\G-A19B\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Atrial Area | meas | m_p_A4Cs_LA | r_ASimpArea | unilateral | 125195\125208\G-03A2\G-A166\G-A19C\R-FAB5B\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4C-A/L | Area | meas | m_ped_A4Cal_LA | r_Area | unilateral | 125195\125226\G-03A2\G-A166\G-A19C\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Atrial Length | meas | m_p_A2Cs_LA | r_ASimpDistance | unilateral | 125195\125208\G-03A2\G-A193\G-A19B\R-FAB5B\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2C-A/L | Length | meas | m_ped_A2Cal_LA | r_Length | unilateral | 125195\125226\G-03A2\G-A193\G-A19B\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Atrial Length | meas | m_p_A4Cs_LA | r_ASimpDistance | unilateral | 125195\125208\G-03A2\G-A193\G-A19C\R-FAB5B\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4C-A/L | Length | meas | m_ped_A4Cal_LA | r_Length | unilateral | 125195\125226\G-03A2\G-A193\G-A19C\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Atrial Volume | meas | m_p_A2Cs_LA | r_ASimpVolume | unilateral | 125195\125208\G-03A2\G-A19B\G-D705\R-FAB5B\T-32300 | SRT, T-32300, Left Atrium | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Atrial Volume | meas | m_p_A4Cs_LA | r_ASimpVolume | unilateral | 125195\125208\G-03A2\G-A19C\G-D705\R-FAB5B\T-32300 | SRT, T-32300, Left Atrium | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 1 | meas | m_p_A2Cs_LA | r_Seg1 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 10 | meas | m_p_A2Cs_LA | r_Seg10 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 11 | meas | m_p_A2Cs_LA | r_Seg11 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 12 | meas | m_p_A2Cs_LA | r_Seg12 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 13 | meas | m_p_A2Cs_LA | r_Seg13 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 14 | meas | m_p_A2Cs_LA | r_Seg14 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 15 | meas | m_p_A2Cs_LA | r_Seg15 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 16 | meas | m_p_A2Cs_LA | r_Seg16 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 17 | meas | m_p_A2Cs_LA | r_Seg17 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 18 | meas | m_p_A2Cs_LA | r_Seg18 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 19 | meas | m_p_A2Cs_LA | r_Seg19 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 2 | meas | m_p_A2Cs_LA | r_Seg2 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 20 | meas | m_p_A2Cs_LA | r_Seg20 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 3 | meas | m_p_A2Cs_LA | r_Seg3 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 4 | meas | m_p_A2Cs_LA | r_Seg4 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 5 | meas | m_p_A2Cs_LA | r_Seg5 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 6 | meas | m_p_A2Cs_LA | r_Seg6 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 7 | meas | m_p_A2Cs_LA | r_Seg7 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 8 | meas | m_p_A2Cs_LA | r_Seg8 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A2Cs | Seg 9 | meas | m_p_A2Cs_LA | r_Seg9 | unilateral | 125195\G-A19B\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 1 | meas | m_p_A4Cs_LA | r_Seg1 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 10 | meas | m_p_A4Cs_LA | r_Seg10 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 11 | meas | m_p_A4Cs_LA | r_Seg11 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 12 | meas | m_p_A4Cs_LA | r_Seg12 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 13 | meas | m_p_A4Cs_LA | r_Seg13 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 14 | meas | m_p_A4Cs_LA | r_Seg14 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 15 | meas | m_p_A4Cs_LA | r_Seg15 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 16 | meas | m_p_A4Cs_LA | r_Seg16 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 17 | meas | m_p_A4Cs_LA | r_Seg17 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 18 | meas | m_p_A4Cs_LA | r_Seg18 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 19 | meas | m_p_A4Cs_LA | r_Seg19 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 2 | meas | m_p_A4Cs_LA | r_Seg2 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 20 | meas | m_p_A4Cs_LA | r_Seg20 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 3 | meas | m_p_A4Cs_LA | r_Seg3 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 4 | meas | m_p_A4Cs_LA | r_Seg4 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 5 | meas | m_p_A4Cs_LA | r_Seg5 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 6 | meas | m_p_A4Cs_LA | r_Seg6 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 7 | meas | m_p_A4Cs_LA | r_Seg7 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 8 | meas | m_p_A4Cs_LA | r_Seg8 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA A4Cs | Seg 9 | meas | m_p_A4Cs_LA | r_Seg9 | unilateral | 125195\G-A19C\M-02550\R-FAB5B\T-32300\T5203-01 | SRT, T-32300, Left Atrium | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SVC S/D | none | calc | c_p_SVC_SD | none | unilateral | 125195\12144-2\M-2460D | SRT, M-2460D, Right Superior vena cava | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVC S/D | none | calc | c_p_IVC_SD | none | unilateral | 125195\12144-2\T-48710 | SRT, T-48710, Inferior vena cava | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Hepatic S/D | none | calc | c_p_HepSD | none | unilateral | 125195\12144-2\T-48720 | SRT, T-48720, Hepatic Vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Pulm S/D | none | calc | c_p_PulmSD | none | unilateral | 125195\12144-2\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RU PulmV S/D | none | calc | c_p_RU_PulmSD | none | unilateral | 125195\12144-2\T-48510\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units | SRT, T-48510, Right Superior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RL PulmV S/D | none | calc | c_p_RL_PulmSD | none | unilateral | 125195\12144-2\T-48520\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units | SRT, T-48520, Right Inferior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LU PulmV S/D | none | calc | c_p_LU_PulmSD | none | unilateral | 125195\12144-2\T-48530\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units | SRT, T-48530, Left Superior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LL PulmV S/D | none | calc | c_p_LL_PulmSD | none | unilateral | 125195\12144-2\T-48540\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units | SRT, T-48540, Left Inferior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Desc Ao Max PG (full) | none | calc | c_p_DscAo_MaxPG_FB | none | unilateral | 125195\125217\20247-3\T-42070 | SRT, T-42070, Thoracic aorta | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Desc Ao Mean PG (full) | none | calc | c_p_DscAo_MeanPG_FB | none | unilateral | 125195\125217\20256-4\T-42070 | SRT, T-42070, Thoracic aorta | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Asc Ao Max PG (full) | none | calc | c_p_AscAo_MaxPG_FB | none | unilateral | 125195\125217\20247-3\T-42100 | SRT, T-42100, Ascending aorta | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Asc Ao Mean PG (full) | none | calc | c_p_AscAo_MeanPG_FB | none | unilateral | 125195\125217\20256-4\T-42100 | SRT, T-42100, Ascending aorta | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AoR Area | none | calc | c_ped_AoR_Area | none | unilateral | 125195\G-038E\G-03A2\T-42110 | SRT, T-42110, Root of Aorta | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA/Ao (MM) | none | calc | c_ped_LA_Ao_MM | none | unilateral | 125195\17985-3\G-0394\T-32300 | SRT, T-32300, Left Atrium | LN, 17985-3, Left Atrium to Aortic Root Ratio |  |  | UCUM, 1, no units |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA/Ao (2D) | none | calc | c_ped_LA_Ao_2D | none | unilateral | 125195\17985-3\G-03A2\T-32300 | SRT, T-32300, Left Atrium | LN, 17985-3, Left Atrium to Aortic Root Ratio |  |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASD Diam | none | calc | c_p_ASD_Diam | none | unilateral | 125195\D4-31220\G-03A2\M-02550\R-41D2D | SRT, D4-31220, Atrial Septal Defect | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | [7 empty] | SRT, R-41D2D, Calculated
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PISA (TR) | none | calc | c_ped_PISA_TR | none | unilateral | 125195\125216\20226-7\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20226-7, Flow Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR Volume | none | calc | c_ped_TR_Vol | none | unilateral | 125195\125216\33878-0\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 33878-0, Volume Flow |  |  | UCUM, mm3, mm3 |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR Flow Rate | none | calc | c_ped_TR_FlowRate | none | unilateral | 125195\34141-2\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 34141-2, Peak Instantaneous Flow Rate |  |  | UCUM, mm3/s, mm3/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV E/A | none | calc | c_ped_TV_EA | none | unilateral | 125195\59104-0\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59104-0, Peak E wave/Peak A wave by US |  |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (TV) | none | calc | c_ped_CO_TV | none | unilateral | 125195\F-32100\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (TV) | none | calc | c_ped_SV_TV | none | unilateral | 125195\F-32120\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TVA (PISA) | none | calc | c_ped_TVA_PISA | none | unilateral | 125195\125216\G-038E\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR ERO | none | calc | c_ped_TR_ERO | none | unilateral | 125195\125216\G-0367\G-038E\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Annul Area | none | calc | c_ped_TV_Annul_Area | none | unilateral | 125195\G-038E\G-03A2\T-3500E\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | SRT, T-3500E, Cardiac valve annulus | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR Fraction | none | calc | c_ped_TR_Frac | none | unilateral | 125195\G-0367\G-0390\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-0390, Regurgitant Fraction |  |  | UCUM, %, Percent |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Area | none | calc | c_ped_TV_Area | none | unilateral | 125195\G-03A2\G-A166\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PISA (MR) | none | calc | c_ped_PISA_MR | none | unilateral | 125195\125216\20226-7\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20226-7, Flow Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR Volume | none | calc | c_ped_MR_Vol | none | unilateral | 125195\125216\33878-0\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 33878-0, Volume Flow |  |  | UCUM, mm3, mm3 |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR Flow Rate | none | calc | c_ped_MR_FlowRate | none | unilateral | 125195\34141-2\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 34141-2, Peak Instantaneous Flow Rate |  |  | UCUM, mm3/s, mm3/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Tei Index | none | calc | c_ped_TeiIndex | none | unilateral | 125195\59099-2\T-35300 | SRT, T-35300, Mitral Valve | LN, 59099-2, Myocardial Performance Index (Tei) |  |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV E/A | none | calc | c_ped_MV_EA | none | unilateral | 125195\59104-0\T-35300 | SRT, T-35300, Mitral Valve | LN, 59104-0, Peak E wave/Peak A wave by US |  |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (MV) | none | calc | c_ped_CO_MV | none | unilateral | 125195\F-32100\T-35300 | SRT, T-35300, Mitral Valve | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (MV) | none | calc | c_ped_SV_MV | none | unilateral | 125195\F-32120\T-35300 | SRT, T-35300, Mitral Valve | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MVA (P½t ) | none | calc | c_ped_MVA_PHalfT | none | unilateral | 125195\125210\G-038E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125210, Area by Pressure Half-Time
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MVA (VTI) | none | calc | c_ped_MVA_VTI | none | unilateral | 125195\125215\G-038E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125215, Continuity Equation by Velocity Time Integral
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MVA (PISA) | none | calc | c_ped_MVA_PISA | none | unilateral | 125195\125216\G-038E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR ERO | none | calc | c_ped_MR_ERO | none | unilateral | 125195\125216\G-0367\G-038E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Area (Ellipse) | none | calc | c_ped_MV_Area_Ellps | none | unilateral | 125195\125211\G-038E\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125211, Biplane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Annul Area | none | calc | c_ped_MV_Annul_Area | none | unilateral | 125195\G-038E\G-03A2\T-3500E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | SRT, T-3500E, Cardiac valve annulus | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR Fraction | none | calc | c_ped_MR_Frac | none | unilateral | 125195\G-0367\G-0390\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-0390, Regurgitant Fraction |  |  | UCUM, %, Percent |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Area | none | calc | c_ped_MV_Area | none | unilateral | 125195\G-03A2\G-A166\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Area | none | calc | c_ped_RVOT_Area | none | unilateral | 125195\G-038E\G-03A2\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | SRT, T-32550, Right Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Area | none | calc | c_ped_LVOT_Area | none | unilateral | 125195\G-038E\G-03A2\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | SRT, T-32650, Left Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MPA Area | none | calc | c_ped_MPA_Area | none | unilateral | 125195\G-038E\G-03A2\T-44100 | SRT, T-44100, Pulmonary Trunk | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PAP (AT) | none | calc | c_ped_PAP_AT | none | unilateral | 125195\59101-6\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 59101-6, Pulmonary Artery Pressure using Accel Time |  |  | UCUM, mmHg, mmHg
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PVA (Vmax) | none | calc | c_ped_PVA_Vmax | none | unilateral | 125195\125214\G-038E\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PVA (VTI) | none | calc | c_ped_PVA_VTI | none | unilateral | 125195\125215\G-038E\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125215, Continuity Equation by Velocity Time Integral
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVSP (AS) | none | calc | c_ped_LVSP_AS | none | unilateral | 125195\C12212-01\R-FAB5B\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12212-01, Left Ventricle Systolic Pressure with Aortic Stenosis |  |  | UCUM, mmHg, mmHg |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVDP (AI) | none | calc | c_ped_LVDP_AI | none | unilateral | 125195\C12212-02\F-32011\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12212-02, Left Ventricle Diastolic Pressure with Aortic Insufficiency |  |  | UCUM, mmHg, mmHg |  |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PISA (AI) | none | calc | c_ped_PISA_AI | none | unilateral | 125195\125216\20226-7\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20226-7, Flow Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Max PG (full) | none | calc | c_p_AV_MaxPG_FB | none | unilateral | 125195\125217\20247-3\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Mean PG (full) | none | calc | c_p_AV_MeanPG_FB | none | unilateral | 125195\125217\20256-4\T-35400 | SRT, T-35400, Aortic Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Volume | none | calc | c_ped_AI_Vol | none | unilateral | 125195\125216\33878-0\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 33878-0, Volume Flow |  |  | UCUM, mm3, mm3 |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Flow Rate | none | calc | c_ped_AI_FlowRate | none | unilateral | 125195\34141-2\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 34141-2, Peak Instantaneous Flow Rate |  |  | UCUM, mm3/s, mm3/s |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AVA (Vmax) | none | calc | c_ped_AVA_Vmax | none | unilateral | 125195\125214\G-038E\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AVA (VTI) | none | calc | c_ped_AVA_VTI | none | unilateral | 125195\125215\G-038E\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125215, Continuity Equation by Velocity Time Integral
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI ERO | none | calc | c_ped_AI_ERO | none | unilateral | 125195\125216\G-0367\G-038E\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Fraction | none | calc | c_ped_AI_Frac | none | unilateral | 125195\G-0367\G-0390\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-0390, Regurgitant Fraction |  |  | UCUM, %, Percent |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVS % (MM) | none | calc | c_ped_IVS_MM | none | unilateral | 125195\59092-7\G-0394\T-32410 | SRT, T-32410, Interventricular septum | LN, 59092-7, % Thickening |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVS % (2D) | none | calc | c_ped_IVS | none | unilateral | 125195\59092-7\G-03A2\T-32410 | SRT, T-32410, Interventricular septum | LN, 59092-7, % Thickening |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVPEPc | none | calc | c_ped_RVPEPc | none | unilateral | 125195\59087-7\T-32500 | SRT, T-32500, Right Ventricle | LN, 59087-7, Heart Rate-Corrected Pre-Ejection Period |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RV PEP/ET | none | calc | c_ped_RV_PEP_ET | none | unilateral | 125195\59088-5\T-32500 | SRT, T-32500, Right Ventricle | LN, 59088-5, Pre-Ejection Period/Ejection Time Ratio |  |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVSP | none | calc | c_ped_RVSP | none | unilateral | 125195\F-31000\R-FAB5B\T-32500 | SRT, T-32500, Right Ventricle | SRT, F-31000, Blood Pressure |  |  | UCUM, mmHg, mmHg |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (RVOT) | none | calc | c_ped_CO_RVOT | none | unilateral | 125195\F-32100\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (RVOT) | none | calc | c_ped_SV_RVOT | none | unilateral | 125195\F-32120\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RV MPI | none | calc | c_ped_RV_MPI | none | unilateral | 125195\G-0381\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-0381, Right Ventricular Index of Myocardial Performance |  |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV Mass Index(Cubed) | none | calc | c_ped_LV_Mass_BSA | none | unilateral | 125195\125206\C12203-01\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-01, Left Ventricle Mass Index |  |  | UCUM, g/m2, g/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV Mass Index (A/L) | none | calc | c_ped_LVMass_A_LBSA | none | unilateral | 125195\125270\C12203-01\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-01, Left Ventricle Mass Index |  |  | UCUM, g/m2, g/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125270, Left Ventricle Mass by Area Length
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV Mass (Cubed) | none | calc | c_ped_LV_Mass | none | unilateral | 125195\125206\18087-7\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18087-7, Left Ventricle Mass |  |  | UCUM, g, grams |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV Mass (A/L) | none | calc | c_ped_LVMass_A_L | none | unilateral | 125195\125270\18087-7\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18087-7, Left Ventricle Mass |  |  | UCUM, g, grams |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125270, Left Ventricle Mass by Area Length
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVS/LVPW (MM) | none | calc | c_ped_IVS_LVPW_MM | none | unilateral | 125195\18155-2\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18155-2, Interventricular Septum to Posterior Wall Thickness Ratio |  |  | UCUM, 1, no units |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVS/LVPW (2D) | none | calc | c_ped_IVS_LVPW | none | unilateral | 125195\18155-2\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18155-2, Interventricular Septum to Posterior Wall Thickness Ratio |  |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVETc | none | calc | c_ped_LVETc | none | unilateral | 125195\59086-9\T-32600 | SRT, T-32600, Left Ventricle | LN, 59086-9, Heart Rate-Corrected Ejection Time |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVPEPc | none | calc | c_ped_LVPEPc | none | unilateral | 125195\59087-7\T-32600 | SRT, T-32600, Left Ventricle | LN, 59087-7, Heart Rate-Corrected Pre-Ejection Period |  |  | UCUM, ms, ms
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV PEP/ET | none | calc | c_ped_LV_PEP_ET | none | unilateral | 125195\59088-5\T-32600 | SRT, T-32600, Left Ventricle | LN, 59088-5, Pre-Ejection Period/Ejection Time Ratio |  |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVPW % (MM) | none | calc | c_ped_LVPW_MM | none | unilateral | 125195\59092-7\G-0394\R-42175\T-32600 | SRT, T-32600, Left Ventricle | LN, 59092-7, % Thickening |  |  | UCUM, %, Percent | SRT, R-42175, Posterior Wall | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVPW % (2D) | none | calc | c_ped_LVPW | none | unilateral | 125195\59092-7\G-03A2\R-42175\T-32600 | SRT, T-32600, Left Ventricle | LN, 59092-7, % Thickening |  |  | UCUM, %, Percent | SRT, R-42175, Posterior Wall | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Wall Stress | none | calc | c_ped_Wall_Stress | none | unilateral | 125195\59097-6\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59097-6, Left Ventricle Meridional Wall Stress |  |  | UCUM, g/m2, g/m2 |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Mean VCF | none | calc | c_ped_Mean_VCF | none | unilateral | 125195\59117-2\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59117-2, Mean Velocity of Circumferential Fiber Shortening (Mean VcFv) |  |  | 99PMSBLUS, circ/s, circ/s |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Mean VCFc | none | calc | c_ped_Mean_VCFc | none | unilateral | 125195\59118-0\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59118-0, HR-Corrected Mean Velocity of Circumferential Fiber Shortening |  |  | 99PMSBLUS, circ/s, circ/s |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | E`/A` Medial | none | calc | c_p_Med_E_A | none | unilateral | 125195\59129-7\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59129-7, Left Ventricle E to A Tissue Velocity Ratio |  |  | UCUM, 1, no units | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | E`/A` Lateral | none | calc | c_p_Lat_E_A | none | unilateral | 125195\59129-7\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59129-7, Left Ventricle E to A Tissue Velocity Ratio |  |  | UCUM, 1, no units | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | FS (MM-Cubed) | none | calc | c_ped_FS_Cubed_MM | none | unilateral | 125195\125206\59132-1\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59132-1, Fractional Shortening |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | FS (MM-Teich) | none | calc | c_ped_FS_Teich_MM | none | unilateral | 125195\125209\59132-1\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59132-1, Fractional Shortening |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | FS (2D-Cubed) | none | calc | c_ped_FS_Cubed | none | unilateral | 125195\125206\59132-1\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 59132-1, Fractional Shortening |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | FS (2D-Teich) | none | calc | c_ped_FS_Teich | none | unilateral | 125195\125209\59132-1\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 59132-1, Fractional Shortening |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (MM-Cubed) | none | calc | c_ped_SI_Cubed_MM | none | unilateral | 125195\125206\F-00078\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (MM-Teich) | none | calc | c_ped_SI_Teich_MM | none | unilateral | 125195\125209\F-00078\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (2D-Cubed) | none | calc | c_ped_SI_Cubed | none | unilateral | 125195\125206\F-00078\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (BP) | none | calc | c_ped_SI_BP | none | unilateral | 125195\125207\F-00078\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (2D-Teich) | none | calc | c_ped_SI_Teich | none | unilateral | 125195\125209\F-00078\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (A2C) | none | calc | c_ped_SI_AP2 | none | unilateral | 125195\125208\F-00078\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (A2C-A/L) | none | calc | c_ped_SI_A2C_SPE | none | unilateral | 125195\125226\F-00078\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (A4C) | none | calc | c_ped_SI_AP4 | none | unilateral | 125195\125208\F-00078\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (A4C-A/L) | none | calc | c_ped_SI_A4C_SPE | none | unilateral | 125195\125226\F-00078\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (MM-Cubed) | none | calc | c_ped_EF_Cubed_MM | none | unilateral | 125195\125206\F-32070\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (MM-Teich) | none | calc | c_ped_EF_Teich_MM | none | unilateral | 125195\125209\F-32070\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (2D-Cubed) | none | calc | c_ped_EF_Cubed | none | unilateral | 125195\125206\F-32070\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (BP) | none | calc | c_ped_EF_BP | none | unilateral | 125195\125207\F-32070\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (2D-Teich) | none | calc | c_ped_EF_Teich | none | unilateral | 125195\125209\F-32070\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (A2C) | none | calc | c_ped_EF_AP2 | none | unilateral | 125195\125208\F-32070\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (A2C-A/L) | none | calc | c_ped_EF_A2C_SPE | none | unilateral | 125195\125226\F-32070\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (A4C) | none | calc | c_ped_EF_AP4 | none | unilateral | 125195\125208\F-32070\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (A4C-A/L) | none | calc | c_ped_EF_A4C_SPE | none | unilateral | 125195\125226\F-32070\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (MM-Cubed) | none | calc | c_ped_CO_Cubed_MM | none | unilateral | 125195\125206\F-32100\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (MM-Teich) | none | calc | c_ped_CO_Teich_MM | none | unilateral | 125195\125209\F-32100\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (2D-Cubed) | none | calc | c_ped_CO_Cubed | none | unilateral | 125195\125206\F-32100\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (BP) | none | calc | c_ped_CO_BP | none | unilateral | 125195\125207\F-32100\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (2D-Teich) | none | calc | c_ped_CO_Teich | none | unilateral | 125195\125209\F-32100\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (A2C) | none | calc | c_ped_CO_AP2 | none | unilateral | 125195\125208\F-32100\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (A2C-A/L) | none | calc | c_ped_CO_A2C_SPE | none | unilateral | 125195\125226\F-32100\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (A4C) | none | calc | c_ped_CO_AP4 | none | unilateral | 125195\125208\F-32100\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (A4C-A/L) | none | calc | c_ped_CO_A4C_SPE | none | unilateral | 125195\125226\F-32100\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (LVOT) | none | calc | c_ped_CO_LVOT | none | unilateral | 125195\F-32100\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (MM-Cubed) | none | calc | c_ped_CI_Cubed_MM | none | unilateral | 125195\125206\F-32110\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (MM-Teich) | none | calc | c_ped_CI_Teich_MM | none | unilateral | 125195\125209\F-32110\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (2D-Cubed) | none | calc | c_ped_CI_Cubed | none | unilateral | 125195\125206\F-32110\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (BP) | none | calc | c_ped_CI_BP | none | unilateral | 125195\125207\F-32110\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (2D-Teich) | none | calc | c_ped_CI_Teich | none | unilateral | 125195\125209\F-32110\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (A2C) | none | calc | c_ped_CI_AP2 | none | unilateral | 125195\125208\F-32110\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (A2C-A/L) | none | calc | c_ped_CI_A2C_SPE | none | unilateral | 125195\125226\F-32110\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (A4C) | none | calc | c_ped_CI_AP4 | none | unilateral | 125195\125208\F-32110\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (A4C-A/L) | none | calc | c_ped_CI_A4C_SPE | none | unilateral | 125195\125226\F-32110\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (MM-Cubed) | none | calc | c_ped_SV_Cubed_MM | none | unilateral | 125195\125206\F-32120\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (MM-Teich) | none | calc | c_ped_SV_Teich_MM | none | unilateral | 125195\125209\F-32120\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (2D-Cubed) | none | calc | c_ped_SV_Cubed | none | unilateral | 125195\125206\F-32120\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (BP) | none | calc | c_ped_SV_BP | none | unilateral | 125195\125207\F-32120\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (2D-Teich) | none | calc | c_ped_SV_Teich | none | unilateral | 125195\125209\F-32120\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (A2C) | none | calc | c_ped_SV_AP2 | none | unilateral | 125195\125208\F-32120\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (A2C-A/L) | none | calc | c_ped_SV_A2C_SPE | none | unilateral | 125195\125226\F-32120\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (A4C) | none | calc | c_ped_SV_AP4 | none | unilateral | 125195\125208\F-32120\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (A4C-A/L) | none | calc | c_ped_SV_A4C_SPE | none | unilateral | 125195\125226\F-32120\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (LVOT) | none | calc | c_ped_SV_LVOT | none | unilateral | 125195\F-32120\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | E/E` Medial | none | calc | c_p_E_MedE | none | unilateral | 125195\G-037B\G-0391\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037B, Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave |  |  | UCUM, 1, no units | SRT, G-0391, Medial Mitral Annulus
Pediatric Cardiac Ultrasound Report | Pediatric Echo | E/E` Lateral | none | calc | c_p_E_LatE | none | unilateral | 125195\G-037B\G-0392\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037B, Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave |  |  | UCUM, 1, no units | SRT, G-0392, Lateral Mitral Annulus
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV MPI | none | calc | c_ped_LV_MPI | none | unilateral | 125195\G-037F\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037F, Left Ventricular Index of Myocardial Performance |  |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (MM-Cubed) | none | calc | c_ped_EDV_Cubed_MM | none | unilateral | 125195\125206\F-32011\G-0394\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (MM-Cubed) | none | calc | c_ped_ESV_Cubed_MM | none | unilateral | 125195\125206\G-0394\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (MM-Teich) | none | calc | c_ped_EDV_Teich_MM | none | unilateral | 125195\125209\F-32011\G-0394\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (MM-Teich) | none | calc | c_ped_ESV_Teich_MM | none | unilateral | 125195\125209\G-0394\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (2D-Cubed) | none | calc | c_ped_EDV_Cubed | none | unilateral | 125195\125206\F-32011\G-03A2\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (2D-Cubed) | none | calc | c_ped_ESV_Cubed | none | unilateral | 125195\125206\G-03A2\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (BP) | none | calc | c_ped_EDV_BP | none | unilateral | 125195\125207\F-32011\G-03A2\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (BP) | none | calc | c_ped_ESV_BP | none | unilateral | 125195\125207\G-03A2\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (2D-Teich) | none | calc | c_ped_EDV_Teich | none | unilateral | 125195\125209\F-32011\G-03A2\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (2D-Teich) | none | calc | c_ped_ESV_Teich | none | unilateral | 125195\125209\G-03A2\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (A2C-A/L) | none | calc | c_ped_EDV_A2C_SPE | none | unilateral | 125195\125226\F-32011\G-03A2\G-A19B\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (A4C-A/L) | none | calc | c_ped_EDV_A4C_SPE | none | unilateral | 125195\125226\F-32011\G-03A2\G-A19C\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (A2C-A/L) | none | calc | c_ped_ESV_A2C_SPE | none | unilateral | 125195\125226\G-03A2\G-A19B\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (A4C-A/L) | none | calc | c_ped_ESV_A4C_SPE | none | unilateral | 125195\125226\G-03A2\G-A19C\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (PV) | none | calc | c_ped_CO_PV | none | unilateral | 125195\F-32100\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (PV) | none | calc | c_ped_SV_PV | none | unilateral | 125195\F-32120\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVSP (VSD) | none | calc | c_p_RVSP_VSD | none | unilateral | 125195\D4-31150\F-31000\R-FAB5B | SRT, D4-31150, Ventricular Septal Defect | SRT, F-31000, Blood Pressure |  |  | UCUM, mmHg, mmHg |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Qp/Qs | none | calc | c_ped_QpQs | none | unilateral | 125195\29462-9\P5-30031 | SRT, P5-30031, Cardiac Shunt Study | LN, 29462-9, Pulmonary-to-Systemic Shunt Flow Ratio |  |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | VSD Diam | none | calc | c_p_VSD_Diam | none | unilateral | 125195\D4-31150\G-03A2\M-02550\R-41D2D | SRT, D4-31150, Ventricular Septal Defect | SRT, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | [7 empty] | SRT, R-41D2D, Calculated
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA ESV Index (BP) | none | calc | c_ped_BP_RA_ESV_BSA | none | unilateral | 125195\125207\C12206-03\G-03A2\T-32200 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-03, Right Atrium Systolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA ESV Index (A2C) | none | calc | c_ped_A2C_RA_ESV_BSA | none | unilateral | 125195\125208\C12206-03\G-03A2\G-A19B\T-32200 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-03, Right Atrium Systolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA ESV Index (A4C) | none | calc | c_ped_A4C_RA_ESV_BSA | none | unilateral | 125195\125208\C12206-03\G-03A2\G-A19C\T-32200 | SRT, T-32200, Right Atrium | 99PMSBLUS, C12206-03, Right Atrium Systolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA ESV(A/L) | none | calc | c_ped_RA_ESV_AL | none | unilateral | 125195\125204\G-03A2\G-D705\T-32200 | SRT, T-32200, Right Atrium | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125204, Area-Length Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA ESV (BP) | none | calc | c_ped_ESV_RA_BP | none | unilateral | 125195\125207\G-03A2\G-D705\T-32200 | SRT, T-32200, Right Atrium | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA ESV Index (BP) | none | calc | c_p_BP_LA_ESV_BSA | none | unilateral | 125195\125207\C12205-03\G-03A2\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-03, Left Atrium Systolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA ESV Index (A2C) | none | calc | c_p_A2C_LA_ESV_BSA | none | unilateral | 125195\125208\C12205-03\G-03A2\G-A19B\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-03, Left Atrium Systolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA ESV Index (A4C) | none | calc | c_p_A4C_LA_ESV_BSA | none | unilateral | 125195\125208\C12205-03\G-03A2\G-A19C\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-03, Left Atrium Systolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA ESV(A/L) | none | calc | c_ped_LA_ESV_AL | none | unilateral | 125195\125204\G-03A2\G-D705\T-32300 | SRT, T-32300, Left Atrium | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125204, Area-Length Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA ESV (BP) | none | calc | c_p_ESV_LA_BP | none | unilateral | 125195\125207\G-03A2\G-D705\T-32300 | SRT, T-32300, Left Atrium | SRT, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Echocardiography Procedure Report | OB | MR Vmax | MaxPG | meas | m_ped_MR_Vmax | r_MaxPG | unilateral | 995300\20247-3\R-42E61\T-35300 | SRT, T-35300, Mitral Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | MV Mean PG | CardPGmean | meas | m_ped_MV_MeanPG | r_CardPGmean | unilateral | 995300\20256-4\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | MR VTI | CardPGmean | meas | m_ped_MR_VTI | r_CardPGmean | unilateral | 995300\20256-4\R-42E61\T-35300 | SRT, T-35300, Mitral Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | MVA - P½t | PHalftime | meas | m_ped_MV_PHalfT | r_PHalftime | unilateral | 995300\20280-4\T-35300 | SRT, T-35300, Mitral Valve | LN, 20280-4, Pressure Half-Time |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | MV VTI | CardTAPV | meas | m_ped_MV_VTI | r_CardTAPV | unilateral | 995300\20352-1\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | MR VTI | CardTAPV | meas | m_ped_MR_VTI | r_CardTAPV | unilateral | 995300\20352-1\R-42E61\T-35300 | SRT, T-35300, Mitral Valve | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | MV VTI | VTI | meas | m_ped_MV_VTI | r_VTI | unilateral | 995300\20354-7\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | MR VTI | VTI | meas | m_ped_MR_VTI | r_VTI | unilateral | 995300\20354-7\R-42E61\T-35300 | SRT, T-35300, Mitral Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | HR - MV | HeartRate | meas | m_ped_HR_MV | r_HeartRate | unilateral | 995300\8867-4\T-35300 | SRT, T-35300, Mitral Valve | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Pediatric Echocardiography Procedure Report | OB | MV Annul Diam | Distance | meas | m_ped_MV_Annul_Diam | r_Distance | unilateral | 995300\M-02550\T-35300\T-35313 | SRT, T-35300, Mitral Valve | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-35313, Mitral Annulus
Pediatric Echocardiography Procedure Report | OB | MV Area (Planim) | Area | meas | m_ped_MV_Area | r_Area | unilateral | 995300\125220\G-038E\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125220, Planimetry
Pediatric Echocardiography Procedure Report | OB | MV Diam | Distance | meas | m_ped_MV_Diam | r_Distance | unilateral | 995300\G-038F\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038F, Cardiovascular Orifice Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | MV Major | Distance | meas | m_ped_MV_Major | r_Distance | unilateral | 995300\G-A193\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | MV Minor | Distance | meas | m_ped_MV_Minor | r_Distance | unilateral | 995300\G-A194\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | MV DFP | DopTime | meas | m_ped_MV_DFP | r_DopTime | unilateral | 995300\R-0032C\T-35300 | SRT, T-35300, Mitral Valve | SRT, R-0032C, Mitral Diastolic Filling Period (DFPm) |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | RVOT Accel | DopSlope | meas | m_p_RVOT_Accel | r_DopSlope | unilateral | 995300\C12222-04\T-32500\T-32550 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12222-04, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 | SNM3, T-32550, Right Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | RVOT Max PG | Vmax | meas | m_ped_RVOT_MaxPG | r_Vmax | unilateral | 995300\11726-7\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s | SNM3, T-32550, Right Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | RVOT Vmax | Vmax | meas | m_ped_RVOT_Vmax | r_Vmax | unilateral | 995300\11726-7\125214\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s | SNM3, T-32550, Right Ventricle Outflow Tract |  |  |  |  | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Echocardiography Procedure Report | OB | RVOT Accel | DopTime | meas | m_p_RVOT_Accel | r_DopTime | unilateral | 995300\20168-1\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms | SNM3, T-32550, Right Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | RVOT Max PG | MaxPG | meas | m_ped_RVOT_MaxPG | r_MaxPG | unilateral | 995300\20247-3\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SNM3, T-32550, Right Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | RVOT Vmax | MaxPG | meas | m_ped_RVOT_Vmax | r_MaxPG | unilateral | 995300\125214\20247-3\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SNM3, T-32550, Right Ventricle Outflow Tract |  |  |  |  | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Echocardiography Procedure Report | OB | RVOT Mean PG | CardPGmean | meas | m_ped_RVOT_PG | r_CardPGmean | unilateral | 995300\20256-4\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg | SNM3, T-32550, Right Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | RVOT VTI | CardTAPV | meas | m_ped_RVOT_VTI | r_CardTAPV | unilateral | 995300\20352-1\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s | SNM3, T-32550, Right Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | RVOT VTI | VTI | meas | m_ped_RVOT_VTI | r_VTI | unilateral | 995300\20354-7\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm | SNM3, T-32550, Right Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | RVOT Diam | Distance | meas | m_ped_RVOT_Diam | r_Distance | unilateral | 995300\G-038F\G-03A2\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, G-038F, Cardiovascular Orifice Diameter |  |  | UCUM, mm, mm | SNM3, T-32550, Right Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | LVOT Max PG | Vmax | meas | m_ped_LVOT_MaxPG | r_Vmax | unilateral | 995300\11726-7\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s | SNM3, T-32650, Left Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | LVOT Vmax | Vmax | meas | m_ped_LVOT_Vmax | r_Vmax | unilateral | 995300\11726-7\125214\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s | SNM3, T-32650, Left Ventricle Outflow Tract |  |  |  |  | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Echocardiography Procedure Report | OB | LVOT Max PG | MaxPG | meas | m_ped_LVOT_MaxPG | r_MaxPG | unilateral | 995300\20247-3\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SNM3, T-32650, Left Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | LVOT Vmax | MaxPG | meas | m_ped_LVOT_Vmax | r_MaxPG | unilateral | 995300\125214\20247-3\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | SNM3, T-32650, Left Ventricle Outflow Tract |  |  |  |  | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Echocardiography Procedure Report | OB | LVOT Mean PG | CardPGmean | meas | m_ped_LVOT_PG | r_CardPGmean | unilateral | 995300\20256-4\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg | SNM3, T-32650, Left Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | LVOT VTI | CardTAPV | meas | m_ped_LVOT_VTI | r_CardTAPV | unilateral | 995300\20352-1\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s | SNM3, T-32650, Left Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | LVOT VTI | VTI | meas | m_ped_LVOT_VTI | r_VTI | unilateral | 995300\20354-7\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm | SNM3, T-32650, Left Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | LVOT Diam | Distance | meas | m_ped_LVOT_Diam | r_Distance | unilateral | 995300\G-038F\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, G-038F, Cardiovascular Orifice Diameter |  |  | UCUM, mm, mm | SNM3, T-32650, Left Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | PA Accel Time | DopSlope | meas | m_ped_PA_AccelTime | r_DopSlope | unilateral | 995300\C12222-04\C3010-01 | 99PMSBLUS, C3010-01, Main pulmonary artery | 99PMSBLUS, C12222-04, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2
Pediatric Echocardiography Procedure Report | OB | MPA Max PG | Vmax | meas | m_ped_MPA_MaxPG | r_Vmax | unilateral | 995300\11726-7\C3010-01 | 99PMSBLUS, C3010-01, Main pulmonary artery | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Echocardiography Procedure Report | OB | PA Accel Time | DopTime | meas | m_ped_PA_AccelTime | r_DopTime | unilateral | 995300\20168-1\C3010-01 | 99PMSBLUS, C3010-01, Main pulmonary artery | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | MPA Max PG | MaxPG | meas | m_ped_MPA_MaxPG | r_MaxPG | unilateral | 995300\20247-3\C3010-01 | 99PMSBLUS, C3010-01, Main pulmonary artery | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Echocardiography Procedure Report | OB | MPA Mean PG | CardPGmean | meas | m_ped_MPA_PG | r_CardPGmean | unilateral | 995300\20256-4\C3010-01 | 99PMSBLUS, C3010-01, Main pulmonary artery | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Echocardiography Procedure Report | OB | MPA Diam | Distance | meas | m_ped_MPA_Diam | r_Distance | unilateral | 995300\C3010-01\M-02550 | 99PMSBLUS, C3010-01, Main pulmonary artery | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | RPA Max PG | Vmax | meas | m_ped_RPA_MaxPG | r_Vmax | unilateral | 995300\11726-7\T-44200 | SNM3, T-44200, Right pulmonary artery | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Echocardiography Procedure Report | OB | RPA Max PG | MaxPG | meas | m_ped_RPA_MaxPG | r_MaxPG | unilateral | 995300\20247-3\T-44200 | SNM3, T-44200, Right pulmonary artery | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Echocardiography Procedure Report | OB | RPA Mean PG | CardPGmean | meas | m_ped_RPA_PG | r_CardPGmean | unilateral | 995300\20256-4\T-44200 | SNM3, T-44200, Right pulmonary artery | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Echocardiography Procedure Report | OB | RPA Diam | Distance | meas | m_ped_RPA_Diam | r_Distance | unilateral | 995300\M-02550\T-44200 | SNM3, T-44200, Right pulmonary artery | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | LPA Max PG | Vmax | meas | m_ped_LPA_MaxPG | r_Vmax | unilateral | 995300\11726-7\T-44400 | SNM3, T-44400, Left pulmonary artery | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Echocardiography Procedure Report | OB | LPA Max PG | MaxPG | meas | m_ped_LPA_MaxPG | r_MaxPG | unilateral | 995300\20247-3\T-44400 | SNM3, T-44400, Left pulmonary artery | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Echocardiography Procedure Report | OB | LPA Mean PG | CardPGmean | meas | m_ped_LPA_PG | r_CardPGmean | unilateral | 995300\20256-4\T-44400 | SNM3, T-44400, Left pulmonary artery | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Echocardiography Procedure Report | OB | LPA Diam | Distance | meas | m_ped_LPA_Diam | r_Distance | unilateral | 995300\M-02550\T-44400 | SNM3, T-44400, Left pulmonary artery | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | RV ET | MMTime | meas | m_ped_RV_Eject_Tm | r_MMTime | unilateral | 995300\C12203-02\G-0394\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12203-02, Eject Time |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | Late Dias Slope | MMSlope | meas | m_ped_LateDiasSlope | r_MMSlope | unilateral | 995300\C12209-01\F-32011\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | 99PMSBLUS, C12209-01, Late Diastolic Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | A Wave Amp | Distance | meas | m_ped_AWaveAmp | r_Distance | unilateral | 995300\C12209-02\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | 99PMSBLUS, C12209-02, A Wave Amp |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | B-C Slope | MMSlope | meas | m_ped_BC_Slope | r_MMSlope | unilateral | 995300\C12209-03\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | 99PMSBLUS, C12209-03, B-C Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | PI P½t | Vmax | meas | m_ped_PI_PHalfT | r_Vmax | unilateral | 995300\C12222-03\R-42E61\T-35200 | SRT, T-35200, Pulmonic Valve | 99PMSBLUS, C12222-03, Pressure Half-Time Peak velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | PV Accel | DopSlope | meas | m_ped_PV_Accel | r_DopSlope | unilateral | 995300\C12222-04\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | 99PMSBLUS, C12222-04, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | PI End Dias Vel | Vel | meas | m_ped_PI_EndDiVel | r_Vel | unilateral | 995300\11653-3\R-42E61\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | PV Max PG | Vmax | meas | m_ped_PV_MaxPG | r_Vmax | unilateral | 995300\11726-7\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | PV Vmax | Vmax | meas | m_ped_PV_Vmax | r_Vmax | unilateral | 995300\11726-7\125214\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Echocardiography Procedure Report | OB | PV Accel | DopTime | meas | m_ped_PV_Accel | r_DopTime | unilateral | 995300\20168-1\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | PI Decel Slope | DopSlope | meas | m_ped_PI_DecelSlope | r_DopSlope | unilateral | 995300\20216-8\R-42E61\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20216-8, Deceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | PI Decel Time | DopTime | meas | m_ped_PI_DecelTime | r_DopTime | unilateral | 995300\20217-6\R-42E61\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | PV Max PG | MaxPG | meas | m_ped_PV_MaxPG | r_MaxPG | unilateral | 995300\20247-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | PV Vmax | MaxPG | meas | m_ped_PV_Vmax | r_MaxPG | unilateral | 995300\125214\20247-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Echocardiography Procedure Report | OB | PV Mean PG | CardPGmean | meas | m_ped_PV_PG | r_CardPGmean | unilateral | 995300\20256-4\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | PI P½t | PHalftime | meas | m_ped_PI_PHalfT | r_PHalftime | unilateral | 995300\20280-4\R-42E61\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20280-4, Pressure Half-Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | PV VTI | CardTAPV | meas | m_ped_PV_VTI | r_CardTAPV | unilateral | 995300\20352-1\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | PV VTI | VTI | meas | m_ped_PV_VTI | r_VTI | unilateral | 995300\20354-7\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | HR - PV | HeartRate | meas | m_ped_HR_PV | r_HeartRate | unilateral | 995300\8867-4\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Pediatric Echocardiography Procedure Report | OB | PV Annul Diam | Distance | meas | m_ped_PV_Annul_Diam | r_Distance | unilateral | 995300\C12209-04\M-02550\T-35200 | SRT, T-35200, Pulmonic Valve | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm | 99PMSBLUS, C12209-04, Pulmonic Annulus
Pediatric Echocardiography Procedure Report | OB | AI Radius | Distance | meas | m_ped_AI_Radius | r_Distance | unilateral | 995300\C12222-01\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12222-01, Flow Radius |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | AI Alias Vel | 2DVel | meas | m_ped_AI_AliasVel | r_2DVel | unilateral | 995300\C12222-02\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12222-02, Alias velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | AV Accel | DopSlope | meas | m_ped_AV_Accel | r_DopSlope | unilateral | 995300\C12222-04\R-42047\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12222-04, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | AI Accel | DopSlope | meas | m_ped_AI_Accel | r_DopSlope | unilateral | 995300\C12222-04\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12222-04, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | AS Vmax | MaxPG | meas | m_ped_AS_Vmax | r_MaxPG | unilateral | 995300\C12222-09\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12222-09, Stenosis Peak Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Echocardiography Procedure Report | OB | AS Vmax | Vmax | meas | m_ped_AS_Vmax | r_Vmax | unilateral | 995300\C12222-10\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12222-10, Stenosis Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Echocardiography Procedure Report | OB | AI End Dias Vel | Vel | meas | m_ped_AI_EndDiVel | r_Vel | unilateral | 995300\11653-3\F-32011\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 11653-3, End Diastolic Velocity |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, F-32011, End Diastole | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | AV Max PG | Vmax | meas | m_ped_AV_MaxPG | r_Vmax | unilateral | 995300\11726-7\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | AV Vmax | Vmax | meas | m_ped_AV_Vmax | r_Vmax | unilateral | 995300\11726-7\125214\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Echocardiography Procedure Report | OB | AI Vmax | Vmax | meas | m_ped_AI_Vmax | r_Vmax | unilateral | 995300\11726-7\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | AV Cusp Sep | Distance | meas | m_ped_AV_Cusp_Sep | r_Distance | unilateral | 995300\17996-0\G-0394\T-35400 | SRT, T-35400, Aortic Valve | LN, 17996-0, Aortic Valve Cusp Separation |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | AV Accel | DopTime | meas | m_ped_AV_Accel | r_DopTime | unilateral | 995300\20168-1\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | AI Accel | DopTime | meas | m_ped_AI_Accel | r_DopTime | unilateral | 995300\20168-1\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | AI Decel Slope | DopSlope | meas | m_ped_AI_DecelSlope | r_DopSlope | unilateral | 995300\20216-8\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20216-8, Deceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | AI Decel Time | DopTime | meas | m_ped_AI_DecelTime | r_DopTime | unilateral | 995300\20217-6\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | AV Vmax | MaxPG | meas | m_ped_AV_Vmax | r_MaxPG | unilateral | 995300\125214\20247-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Echocardiography Procedure Report | OB | AI Vmax | MaxPG | meas | m_ped_AI_Vmax | r_MaxPG | unilateral | 995300\20247-3\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | AV Mean PG | CardPGmean | meas | m_ped_AV_PG | r_CardPGmean | unilateral | 995300\20256-4\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | AI VTI | CardPGmean | meas | m_ped_AI_VTI | r_CardPGmean | unilateral | 995300\20256-4\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | AI P½t | PHalftime | meas | m_ped_AI_PHalfT | r_PHalftime | unilateral | 995300\20280-4\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20280-4, Pressure Half-Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | AV VTI | CardTAPV | meas | m_ped_AV_VTI | r_CardTAPV | unilateral | 995300\20352-1\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | AI VTI | CardTAPV | meas | m_ped_AI_VTI | r_CardTAPV | unilateral | 995300\20352-1\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | AV VTI | VTI | meas | m_ped_AV_VTI | r_VTI | unilateral | 995300\20354-7\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | AI VTI | VTI | meas | m_ped_AI_VTI | r_VTI | unilateral | 995300\20354-7\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | HR - AV | HeartRate | meas | m_ped_HR_AV | r_HeartRate | unilateral | 995300\8867-4\T-35400 | SRT, T-35400, Aortic Valve | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Pediatric Echocardiography Procedure Report | OB | AV Area | Area | meas | m_ped_AV_Area | r_Area | unilateral | 995300\G-A166\T-35400 | SRT, T-35400, Aortic Valve | SNM3, G-A166, Area |  |  | UCUM, mm2, mm2
Pediatric Echocardiography Procedure Report | OB | AV Annul Diam | Distance | meas | m_ped_AV_Annul_Diam | r_Distance | unilateral | 995300\M-02550\T-35400\T-35410 | SRT, T-35400, Aortic Valve | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-35410, Aortic Valve Ring
Pediatric Echocardiography Procedure Report | OB | RV ET | DopTime | meas | m_ped_RV_ET_Dop | r_DopTime | unilateral | 995300\C12203-02\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12203-02, Eject Time |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | RV PEP | MMTime | meas | m_ped_RV_PreEject_Tm | r_MMTime | unilateral | 995300\C12203-03\G-0394\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12203-03, Pre-Eject Time |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | RVAWd (MM) | Distance | meas | m_ped_RVAWd_MM | r_Distance | unilateral | 995300\C12204-03\C7470-25\F-32011\G-0394\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C7470-25, Thickness |  |  | UCUM, mm, mm | 99PMSBLUS, C12204-03, Right Ventricular Anterior Wall | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | RVAWd (2D) | Distance | meas | m_ped_RVAWd_2D | r_Distance | unilateral | 995300\C12204-03\C7470-25\F-32011\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C7470-25, Thickness |  |  | UCUM, mm, mm | 99PMSBLUS, C12204-03, Right Ventricular Anterior Wall | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | RVIDd (MM) | Distance | meas | m_ped_RVIDd_MM | r_Distance | unilateral | 995300\C7470-26\F-32011\G-0394\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C7470-26, Internal Dimension |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | RVIDd (2D) | Distance | meas | m_ped_RVIDd_2D | r_Distance | unilateral | 995300\C7470-26\F-32011\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C7470-26, Internal Dimension |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | RVes Major - A4C | Distance | meas | m_ped_RVES_MAJOR_A4C | r_Distance | unilateral | 995300\109070\G-A193\G-A19C\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | RVed Major - A4C | Distance | meas | m_ped_RVED_MAJOR_A4C | r_Distance | unilateral | 995300\F-32011\G-A193\G-A19C\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | RVes Minor - A4C | Distance | meas | m_ped_RVES_MINOR_A4C | r_Distance | unilateral | 995300\109070\G-A194\G-A19C\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | RVed Minor - A4C | Distance | meas | m_ped_RVED_MINOR_A4C | r_Distance | unilateral | 995300\F-32011\G-A194\G-A19C\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | LV ET | DopTime | meas | m_ped_LV_ET_Dop | r_DopTime | unilateral | 995300\C12203-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-02, Eject Time |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | LV ET | MMTime | meas | m_ped_LV_Eject_Tm | r_MMTime | unilateral | 995300\C12203-02\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-02, Eject Time |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | LV PEP | MMTime | meas | m_ped_LV_PreEject_Tm | r_MMTime | unilateral | 995300\C12203-03\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-03, Pre-Eject Time |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | Time to Med S | DopTime | meas | m_p_TimeMedS | r_DopTime | unilateral | 995300\C12203-05\G-0391\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-05, Time to LV S Tissue Velocity |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | Time to Lat S | DopTime | meas | m_p_TimeLatS | r_DopTime | unilateral | 995300\C12203-05\G-0392\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-05, Time to LV S Tissue Velocity |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | Time to Med E` | DopTime | meas | m_p_TimeMedE | r_DopTime | unilateral | 995300\C12203-06\G-0391\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-06, Time to LV E Tissue Velocity |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | Time to Lat E` | DopTime | meas | m_p_TimeLatE | r_DopTime | unilateral | 995300\C12203-06\G-0392\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-06, Time to LV E Tissue Velocity |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | Med E` Area | VTI | meas | m_p_AreaUnderMedE | r_VTI | unilateral | 995300\C12203-07\G-0391\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-07, Area under LV E Tissue Velocity |  |  | UCUM, mm, mm | SRT, G-0391, Medial Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | Lat E` Area | VTI | meas | m_p_AreaUnderLatE | r_VTI | unilateral | 995300\C12203-07\G-0392\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-07, Area under LV E Tissue Velocity |  |  | UCUM, mm, mm | SRT, G-0392, Lateral Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | Med A` Area | VTI | meas | m_p_AreaUnderMedA | r_VTI | unilateral | 995300\C12203-08\G-0391\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-08, Area under LV A Tissue Velocity |  |  | UCUM, mm, mm | SRT, G-0391, Medial Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | Lat A` Area | VTI | meas | m_p_AreaUnderLatA | r_VTI | unilateral | 995300\C12203-08\G-0392\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-08, Area under LV A Tissue Velocity |  |  | UCUM, mm, mm | SRT, G-0392, Lateral Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | LVOT Accel | DopSlope | meas | m_p_LVOT_Accel | r_DopSlope | unilateral | 995300\C12222-04\T-32600\T-32650 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12222-04, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 | SNM3, T-32650, Left Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | Med E` Vel | Vel | meas | m_p_Med_EVel | r_Vel | unilateral | 995300\C12233-01\C3612-04\G-0391\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3612-04, Tissue Velocity |  |  | UCUM, mm/s, mm/s | SRT, G-0391, Medial Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging |  | 99PMSBLUS, C12233-01, Early Diastole
Pediatric Echocardiography Procedure Report | OB | Med S Vel | Vel | meas | m_p_Med_SVel | r_Vel | unilateral | 995300\109070\C3612-04\G-0391\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3612-04, Tissue Velocity |  |  | UCUM, mm/s, mm/s | SRT, G-0391, Medial Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | Lat E` Vel | Vel | meas | m_p_Lat_EVel | r_Vel | unilateral | 995300\C12233-01\C3612-04\G-0392\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3612-04, Tissue Velocity |  |  | UCUM, mm/s, mm/s | SRT, G-0392, Lateral Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging |  | 99PMSBLUS, C12233-01, Early Diastole
Pediatric Echocardiography Procedure Report | OB | Lat S Vel | Vel | meas | m_p_Lat_SVel | r_Vel | unilateral | 995300\109070\C3612-04\G-0392\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3612-04, Tissue Velocity |  |  | UCUM, mm/s, mm/s | SRT, G-0392, Lateral Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | Med A` Vel | Vel | meas | m_p_Med_AVel | r_Vel | unilateral | 995300\C3612-05\F-32011\G-0391\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3612-05, Tissue Velocity During Atrial Systole |  |  | UCUM, mm/s, mm/s | SRT, G-0391, Medial Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | Lat A` Vel | Vel | meas | m_p_Lat_AVel | r_Vel | unilateral | 995300\C3612-05\F-32011\G-0392\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3612-05, Tissue Velocity During Atrial Systole |  |  | UCUM, mm/s, mm/s | SRT, G-0392, Lateral Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | IVRT | DopTime | meas | m_ped_IVRT | r_DopTime | unilateral | 995300\C3613-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3613-02, Isovolumic Relaxation Time |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | Med IVRT | DopTime | meas | m_p_TDI_Med_IVRT | r_DopTime | unilateral | 995300\C3613-02\G-0391\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3613-02, Isovolumic Relaxation Time |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | Lat IVRT | DopTime | meas | m_p_TDI_Lat_IVRT | r_DopTime | unilateral | 995300\C3613-02\G-0392\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3613-02, Isovolumic Relaxation Time |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | IVCT | DopTime | meas | m_ped_IVCT | r_DopTime | unilateral | 995300\C3613-03\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3613-03, Isovolumic Contraction Time |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | Med IVCT | DopTime | meas | m_p_TDI_Med_IVCT | r_DopTime | unilateral | 995300\C3613-03\G-0391\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3613-03, Isovolumic Contraction Time |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | Lat IVCT | DopTime | meas | m_p_TDI_Lat_IVCT | r_DopTime | unilateral | 995300\C3613-03\G-0392\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3613-03, Isovolumic Contraction Time |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | LVPWs (MM) | Distance | meas | m_ped_LVPWs_MM | r_Distance | unilateral | 995300\109070\C12243-01\C7470-25\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C7470-25, Thickness |  |  | UCUM, mm, mm | 99PMSBLUS, C12243-01, Left Ventricle Posterior Wall | SRT, G-0394, M mode |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LVPWd (MM) | Distance | meas | m_ped_LVPWd_MM | r_Distance | unilateral | 995300\C12243-01\C7470-25\F-32011\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C7470-25, Thickness |  |  | UCUM, mm, mm | 99PMSBLUS, C12243-01, Left Ventricle Posterior Wall | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | LVPWs (2D) | Distance | meas | m_ped_LVPWs_2D | r_Distance | unilateral | 995300\109070\C12243-01\C7470-25\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C7470-25, Thickness |  |  | UCUM, mm, mm | 99PMSBLUS, C12243-01, Left Ventricle Posterior Wall | SRT, G-03A2, 2D mode |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LVPWd (2D) | Distance | meas | m_ped_LVPWd_2D | r_Distance | unilateral | 995300\C12243-01\C7470-25\F-32011\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C7470-25, Thickness |  |  | UCUM, mm, mm | 99PMSBLUS, C12243-01, Left Ventricle Posterior Wall | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | LVIDs (MM) | Distance | meas | m_ped_LVIDs_MM | r_Distance | unilateral | 995300\109070\C7470-26\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C7470-26, Internal Dimension |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LVIDd (MM) | Distance | meas | m_ped_LVIDd_MM | r_Distance | unilateral | 995300\C7470-26\F-32011\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C7470-26, Internal Dimension |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | LVIDs (2D) | Distance | meas | m_ped_LVIDs_2D | r_Distance | unilateral | 995300\109070\C7470-26\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C7470-26, Internal Dimension |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LVIDd (2D) | Distance | meas | m_ped_LVIDd_2D | r_Distance | unilateral | 995300\C7470-26\F-32011\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C7470-26, Internal Dimension |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | LVAd Sax Epi | Area | meas | m_ped_LV_AdSaxEpi | r_Area | unilateral | 995300\C7471-05\F-32011\G-039B\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C7471-05, Epicardial Area |  |  | UCUM, mm2, mm2 |  |  | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | LVAd Sax Endo | Area | meas | m_ped_LV_AdSaxEndo | r_Area | unilateral | 995300\C7471-07\F-32011\G-039B\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C7471-07, Endocardial Area |  |  | UCUM, mm2, mm2 |  |  | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | LVes - SAX CH | Distance | meas | m_ped_LVES_SAX_CH | r_Distance | unilateral | 995300\109070\121206\G-0399\T-32600 | SRT, T-32600, Left Ventricle | DCM, 121206, Distance |  |  | UCUM, mm, mm |  |  | SRT, G-0399, Parasternal short axis at the level of the mitral chords | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LVed - SAX CH | Distance | meas | m_ped_LVED_SAX_CH | r_Distance | unilateral | 995300\121206\F-32011\G-0399\T-32600 | SRT, T-32600, Left Ventricle | DCM, 121206, Distance |  |  | UCUM, mm, mm |  |  | SRT, G-0399, Parasternal short axis at the level of the mitral chords | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | LVes - SAX PM | Distance | meas | m_ped_LVES_SAX_PM | r_Distance | unilateral | 995300\109070\121206\G-039B\T-32600 | SRT, T-32600, Left Ventricle | DCM, 121206, Distance |  |  | UCUM, mm, mm |  |  | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LVed - SAX PM | Distance | meas | m_ped_LVED_SAX_PM | r_Distance | unilateral | 995300\121206\F-32011\G-039B\T-32600 | SRT, T-32600, Left Ventricle | DCM, 121206, Distance |  |  | UCUM, mm, mm |  |  | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | Med Accel Time | DopTime | meas | m_p_TDIMedAccelTime | r_DopTime | unilateral | 995300\20168-1\G-0391\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | Lat Accel Time | DopTime | meas | m_p_TDILatAccelTime | r_DopTime | unilateral | 995300\20168-1\G-0392\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | LVOT Accel | DopTime | meas | m_p_LVOT_Accel | r_DopTime | unilateral | 995300\20168-1\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms | SNM3, T-32650, Left Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | Med Decel Time | DopTime | meas | m_p_TDIMedDecelTime | r_DopTime | unilateral | 995300\20217-6\G-0391\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms | SRT, G-0391, Medial Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | Lat Decel Time | DopTime | meas | m_p_TDILatDecelTime | r_DopTime | unilateral | 995300\20217-6\G-0392\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms | SRT, G-0392, Lateral Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | HR - LV | HeartRate | meas | m_ped_HR_LV | r_HeartRate | unilateral | 995300\8867-4\T-32600 | SRT, T-32600, Left Ventricle | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Pediatric Echocardiography Procedure Report | OB | A2Cs | SimpArea | meas | m_ped_A2Cs | r_SimpArea | unilateral | 995300\109070\125208\G-03A2\G-A166\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | SimpArea | meas | m_ped_A4Cs | r_SimpArea | unilateral | 995300\109070\125208\G-03A2\G-A166\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | SimpArea | meas | m_ped_A2Cd | r_SimpArea | unilateral | 995300\125208\F-32011\G-03A2\G-A166\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | SimpArea | meas | m_ped_A4Cd | r_SimpArea | unilateral | 995300\125208\F-32011\G-03A2\G-A166\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | SimpVolume | meas | m_ped_A2Cs | r_SimpVolume | unilateral | 995300\109070\125208\G-03A2\G-A19B\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | SimpVolume | meas | m_ped_A4Cs | r_SimpVolume | unilateral | 995300\109070\125208\G-03A2\G-A19C\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | SimpVolume | meas | m_ped_A2Cd | r_SimpVolume | unilateral | 995300\125208\F-32011\G-03A2\G-A19B\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | SimpVolume | meas | m_ped_A4Cd | r_SimpVolume | unilateral | 995300\125208\F-32011\G-03A2\G-A19C\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg1 | meas | m_ped_A2Cs | r_Seg1 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg10 | meas | m_ped_A2Cs | r_Seg10 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg11 | meas | m_ped_A2Cs | r_Seg11 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg12 | meas | m_ped_A2Cs | r_Seg12 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg13 | meas | m_ped_A2Cs | r_Seg13 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg14 | meas | m_ped_A2Cs | r_Seg14 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg15 | meas | m_ped_A2Cs | r_Seg15 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg16 | meas | m_ped_A2Cs | r_Seg16 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg17 | meas | m_ped_A2Cs | r_Seg17 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg18 | meas | m_ped_A2Cs | r_Seg18 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg19 | meas | m_ped_A2Cs | r_Seg19 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg2 | meas | m_ped_A2Cs | r_Seg2 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg20 | meas | m_ped_A2Cs | r_Seg20 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg3 | meas | m_ped_A2Cs | r_Seg3 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg4 | meas | m_ped_A2Cs | r_Seg4 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg5 | meas | m_ped_A2Cs | r_Seg5 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg6 | meas | m_ped_A2Cs | r_Seg6 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg7 | meas | m_ped_A2Cs | r_Seg7 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg8 | meas | m_ped_A2Cs | r_Seg8 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | Seg9 | meas | m_ped_A2Cs | r_Seg9 | unilateral | 995300\109070\125208\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg1 | meas | m_ped_A2Cd | r_Seg1 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg10 | meas | m_ped_A2Cd | r_Seg10 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg11 | meas | m_ped_A2Cd | r_Seg11 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg12 | meas | m_ped_A2Cd | r_Seg12 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg13 | meas | m_ped_A2Cd | r_Seg13 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg14 | meas | m_ped_A2Cd | r_Seg14 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg15 | meas | m_ped_A2Cd | r_Seg15 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg16 | meas | m_ped_A2Cd | r_Seg16 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg17 | meas | m_ped_A2Cd | r_Seg17 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg18 | meas | m_ped_A2Cd | r_Seg18 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg19 | meas | m_ped_A2Cd | r_Seg19 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg2 | meas | m_ped_A2Cd | r_Seg2 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg20 | meas | m_ped_A2Cd | r_Seg20 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg3 | meas | m_ped_A2Cd | r_Seg3 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg4 | meas | m_ped_A2Cd | r_Seg4 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg5 | meas | m_ped_A2Cd | r_Seg5 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg6 | meas | m_ped_A2Cd | r_Seg6 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg7 | meas | m_ped_A2Cd | r_Seg7 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg8 | meas | m_ped_A2Cd | r_Seg8 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | Seg9 | meas | m_ped_A2Cd | r_Seg9 | unilateral | 995300\125208\F-32011\G-A19B\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg1 | meas | m_ped_A4Cs | r_Seg1 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg10 | meas | m_ped_A4Cs | r_Seg10 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg11 | meas | m_ped_A4Cs | r_Seg11 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg12 | meas | m_ped_A4Cs | r_Seg12 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg13 | meas | m_ped_A4Cs | r_Seg13 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg14 | meas | m_ped_A4Cs | r_Seg14 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg15 | meas | m_ped_A4Cs | r_Seg15 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg16 | meas | m_ped_A4Cs | r_Seg16 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg17 | meas | m_ped_A4Cs | r_Seg17 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg18 | meas | m_ped_A4Cs | r_Seg18 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg19 | meas | m_ped_A4Cs | r_Seg19 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg2 | meas | m_ped_A4Cs | r_Seg2 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg20 | meas | m_ped_A4Cs | r_Seg20 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg3 | meas | m_ped_A4Cs | r_Seg3 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg4 | meas | m_ped_A4Cs | r_Seg4 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg5 | meas | m_ped_A4Cs | r_Seg5 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg6 | meas | m_ped_A4Cs | r_Seg6 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg7 | meas | m_ped_A4Cs | r_Seg7 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg8 | meas | m_ped_A4Cs | r_Seg8 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | Seg9 | meas | m_ped_A4Cs | r_Seg9 | unilateral | 995300\109070\125208\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg1 | meas | m_ped_A4Cd | r_Seg1 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg10 | meas | m_ped_A4Cd | r_Seg10 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | AV Max PG | MaxPG | meas | m_ped_AV_MaxPG | r_MaxPG | unilateral | 995300\20247-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg12 | meas | m_ped_A4Cd | r_Seg12 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg13 | meas | m_ped_A4Cd | r_Seg13 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg14 | meas | m_ped_A4Cd | r_Seg14 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg15 | meas | m_ped_A4Cd | r_Seg15 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg16 | meas | m_ped_A4Cd | r_Seg16 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg17 | meas | m_ped_A4Cd | r_Seg17 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg18 | meas | m_ped_A4Cd | r_Seg18 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg19 | meas | m_ped_A4Cd | r_Seg19 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg2 | meas | m_ped_A4Cd | r_Seg2 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg20 | meas | m_ped_A4Cd | r_Seg20 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg3 | meas | m_ped_A4Cd | r_Seg3 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg4 | meas | m_ped_A4Cd | r_Seg4 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg5 | meas | m_ped_A4Cd | r_Seg5 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg6 | meas | m_ped_A4Cd | r_Seg6 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg7 | meas | m_ped_A4Cd | r_Seg7 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg8 | meas | m_ped_A4Cd | r_Seg8 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg9 | meas | m_ped_A4Cd | r_Seg9 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cs | SimpDistance | meas | m_ped_A2Cs | r_SimpDistance | unilateral | 995300\109070\125208\G-03A2\G-A193\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cs | SimpDistance | meas | m_ped_A4Cs | r_SimpDistance | unilateral | 995300\109070\125208\G-03A2\G-A193\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A2Cd | SimpDistance | meas | m_ped_A2Cd | r_SimpDistance | unilateral | 995300\125208\F-32011\G-03A2\G-A193\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | A4Cd | SimpDistance | meas | m_ped_A4Cd | r_SimpDistance | unilateral | 995300\125208\F-32011\G-03A2\G-A193\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | LVLd Apical | Distance | meas | m_ped_LV_LdApical | r_Distance | unilateral | 995300\F-32011\G-A193\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | LVes Major - A4C | Distance | meas | m_ped_LVES_MAJOR_A4C | r_Distance | unilateral | 995300\109070\G-A193\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LVed Major - A4C | Distance | meas | m_ped_LVED_MAJOR_A4C | r_Distance | unilateral | 995300\F-32011\G-A193\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | LVes Minor - A4C | Distance | meas | m_ped_LVES_MINOR_A4C | r_Distance | unilateral | 995300\109070\G-A194\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LVed Minor - A4C | Distance | meas | m_ped_LVED_MINOR_A4C | r_Distance | unilateral | 995300\F-32011\G-A194\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | LV dP/dt | DopDpDt | meas | m_ped_LV_dPdt | r_DopDpDt | unilateral | 995300\18035-6\T-35300 | SRT, T-35300, Mitral Valve | LN, 18035-6, Mitral Regurgitation dP/dt derived from Mitral Reg velocity |  |  | UCUM, mmHg, mmHg
Pediatric Echocardiography Procedure Report | OB | VSD Vmax | Vmax | meas | m_p_VSD_Vmax | r_Vmax | unilateral | 995300\11726-7\D4-31150 | SRT, D4-31150, Ventricular Septal Defect | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Echocardiography Procedure Report | OB | VSD Vmax | MaxPG | meas | m_p_VSD_Vmax | r_MaxPG | unilateral | 995300\20247-3\D4-31150 | SRT, D4-31150, Ventricular Septal Defect | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Echocardiography Procedure Report | OB | VSD Major | Distance | meas | m_ped_VSD_Major | r_Distance | unilateral | 995300\D4-31150\G-A193 | SRT, D4-31150, Ventricular Septal Defect | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | VSD Minor | Distance | meas | m_ped_VSD_Minor | r_Distance | unilateral | 995300\D4-31150\G-A194 | SRT, D4-31150, Ventricular Septal Defect | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | IVSs (MM) | Distance | meas | m_ped_IVSs_MM | r_Distance | unilateral | 995300\109070\C7470-25\G-0394\T-32410 | SRT, T-32410, Interventricular septum | 99PMSBLUS, C7470-25, Thickness |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | IVSd (MM) | Distance | meas | m_ped_IVSd_MM | r_Distance | unilateral | 995300\C7470-25\F-32011\G-0394\T-32410 | SRT, T-32410, Interventricular septum | 99PMSBLUS, C7470-25, Thickness |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | IVSs (2D) | Distance | meas | m_ped_IVSs_2D | r_Distance | unilateral | 995300\109070\C7470-25\G-03A2\T-32410 | SRT, T-32410, Interventricular septum | 99PMSBLUS, C7470-25, Thickness |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | IVSd (2D) | Distance | meas | m_ped_IVSd_2D | r_Distance | unilateral | 995300\C7470-25\F-32011\G-03A2\T-32410 | SRT, T-32410, Interventricular septum | 99PMSBLUS, C7470-25, Thickness |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | ASimpDistance | meas | m_p_A2Cs_LA | r_ASimpDistance | unilateral | 995300\125208\C12205-01\G-03A2\G-A19B\T-32300 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-01, Left Atrium systolic major axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | ASimpDistance | meas | m_p_A4Cs_LA | r_ASimpDistance | unilateral | 995300\125208\C12205-01\G-03A2\G-A19C\T-32300 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-01, Left Atrium systolic major axis |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg1 | meas | m_p_A2Cs_LA | r_Seg1 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg10 | meas | m_p_A2Cs_LA | r_Seg10 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg11 | meas | m_p_A2Cs_LA | r_Seg11 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg12 | meas | m_p_A2Cs_LA | r_Seg12 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg13 | meas | m_p_A2Cs_LA | r_Seg13 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg14 | meas | m_p_A2Cs_LA | r_Seg14 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg15 | meas | m_p_A2Cs_LA | r_Seg15 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg16 | meas | m_p_A2Cs_LA | r_Seg16 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg17 | meas | m_p_A2Cs_LA | r_Seg17 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg18 | meas | m_p_A2Cs_LA | r_Seg18 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg19 | meas | m_p_A2Cs_LA | r_Seg19 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg2 | meas | m_p_A2Cs_LA | r_Seg2 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg20 | meas | m_p_A2Cs_LA | r_Seg20 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg3 | meas | m_p_A2Cs_LA | r_Seg3 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg4 | meas | m_p_A2Cs_LA | r_Seg4 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg5 | meas | m_p_A2Cs_LA | r_Seg5 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg6 | meas | m_p_A2Cs_LA | r_Seg6 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg7 | meas | m_p_A2Cs_LA | r_Seg7 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg8 | meas | m_p_A2Cs_LA | r_Seg8 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | Seg9 | meas | m_p_A2Cs_LA | r_Seg9 | unilateral | 995300\109070\C12205-02\G-A19B\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg1 | meas | m_p_A4Cs_LA | r_Seg1 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg10 | meas | m_p_A4Cs_LA | r_Seg10 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg11 | meas | m_p_A4Cs_LA | r_Seg11 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg12 | meas | m_p_A4Cs_LA | r_Seg12 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg13 | meas | m_p_A4Cs_LA | r_Seg13 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg14 | meas | m_p_A4Cs_LA | r_Seg14 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg15 | meas | m_p_A4Cs_LA | r_Seg15 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg16 | meas | m_p_A4Cs_LA | r_Seg16 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg17 | meas | m_p_A4Cs_LA | r_Seg17 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg18 | meas | m_p_A4Cs_LA | r_Seg18 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg19 | meas | m_p_A4Cs_LA | r_Seg19 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg2 | meas | m_p_A4Cs_LA | r_Seg2 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg20 | meas | m_p_A4Cs_LA | r_Seg20 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg3 | meas | m_p_A4Cs_LA | r_Seg3 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg4 | meas | m_p_A4Cs_LA | r_Seg4 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg5 | meas | m_p_A4Cs_LA | r_Seg5 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg6 | meas | m_p_A4Cs_LA | r_Seg6 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg7 | meas | m_p_A4Cs_LA | r_Seg7 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg8 | meas | m_p_A4Cs_LA | r_Seg8 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | Seg9 | meas | m_p_A4Cs_LA | r_Seg9 | unilateral | 995300\109070\C12205-02\G-A19C\T-32300\T5203-01 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-02, Left Atrium MOD Diam |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | ASimpArea | meas | m_p_A2Cs_LA | r_ASimpArea | unilateral | 995300\125208\17977-0\G-03A2\G-A19B\T-32300 | SNM3, T-32300, Left Atrium | LN, 17977-0, Left Atrium Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | ASimpArea | meas | m_p_A4Cs_LA | r_ASimpArea | unilateral | 995300\125208\17977-0\G-03A2\G-A19C\T-32300 | SNM3, T-32300, Left Atrium | LN, 17977-0, Left Atrium Systolic Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | LA A2Cs | ASimpVolume | meas | m_p_A2Cs_LA | r_ASimpVolume | unilateral | 995300\125208\G-0383\G-03A2\G-A19B\T-32300 | SNM3, T-32300, Left Atrium | SRT, G-0383, Left Atrium Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | LA A4Cs | ASimpVolume | meas | m_p_A4Cs_LA | r_ASimpVolume | unilateral | 995300\125208\G-0383\G-03A2\G-A19C\T-32300 | SNM3, T-32300, Left Atrium | SRT, G-0383, Left Atrium Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | SVC A Vel | Vel | meas | m_p_SVC_ARevsVel | r_Vel | unilateral | 995300\C3010-03\C3612-01 | 99PMSBLUS, C3010-03, Right Superior vena cava | 99PMSBLUS, C3612-01, Atrial Contraction Reversal Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Echocardiography Procedure Report | OB | SVC A Dur | DopTime | meas | m_p_SVC_ARevsDur | r_DopTime | unilateral | 995300\C3010-03\C3613-08 | 99PMSBLUS, C3010-03, Right Superior vena cava | 99PMSBLUS, C3613-08, A Wave Duration |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | SVC Dias Vel | Vel | meas | m_p_SVC_DiasVel | r_Vel | unilateral | 995300\122202\C3010-03\F-32011 | 99PMSBLUS, C3010-03, Right Superior vena cava | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | SVC Sys Vel | Vel | meas | m_p_SVC_SysVel | r_Vel | unilateral | 995300\109070\122204\C3010-03 | 99PMSBLUS, C3010-03, Right Superior vena cava | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | SVC Diam | Distance | meas | m_ped_SVC_Diam | r_Distance | unilateral | 995300\C3010-03\M-02550 | 99PMSBLUS, C3010-03, Right Superior vena cava | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | Hepatic A Vel | Vel | meas | m_p_HepARevsVel | r_Vel | unilateral | 995300\C3612-01\T5200-01 | 99PMSBLUS, T5200-01, Hepatic Veins | 99PMSBLUS, C3612-01, Atrial Contraction Reversal Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Echocardiography Procedure Report | OB | Hepatic A Dur | DopTime | meas | m_p_HepARevsDur | r_DopTime | unilateral | 995300\C3613-08\T5200-01 | 99PMSBLUS, T5200-01, Hepatic Veins | 99PMSBLUS, C3613-08, A Wave Duration |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | Hepatic Dias Vel | Vel | meas | m_p_HepDiasVel | r_Vel | unilateral | 995300\122202\F-32011\T5200-01 | 99PMSBLUS, T5200-01, Hepatic Veins | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | Hepatic Sys Vel | Vel | meas | m_p_HepSysVel | r_Vel | unilateral | 995300\109070\122204\T5200-01 | 99PMSBLUS, T5200-01, Hepatic Veins | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | IVC A Vel | Vel | meas | m_p_IVC_ARevsVel | r_Vel | unilateral | 995300\C3612-01\T-48710 | SNM3, T-48710, Inferior vena cava | 99PMSBLUS, C3612-01, Atrial Contraction Reversal Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Echocardiography Procedure Report | OB | IVC A Dur | DopTime | meas | m_p_IVC_ARevsDur | r_DopTime | unilateral | 995300\C3613-08\T-48710 | SNM3, T-48710, Inferior vena cava | 99PMSBLUS, C3613-08, A Wave Duration |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | IVC Dias Vel | Vel | meas | m_p_IVC_DiasVel | r_Vel | unilateral | 995300\122202\F-32011\T-48710 | SNM3, T-48710, Inferior vena cava | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | IVC Sys Vel | Vel | meas | m_p_IVC_SysVel | r_Vel | unilateral | 995300\109070\122204\T-48710 | SNM3, T-48710, Inferior vena cava | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | IVC Diam | Distance | meas | m_ped_IVC_Diam | r_Distance | unilateral | 995300\M-02550\T-48710 | SNM3, T-48710, Inferior vena cava | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | Pulm A Vel | Vel | meas | m_p_PulmARevsVel | r_Vel | unilateral | 995300\C3612-01\T-48500 | SRT, T-48500, Pulmonary vein | 99PMSBLUS, C3612-01, Atrial Contraction Reversal Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Echocardiography Procedure Report | OB | LU PulmV A Vel | Vel | meas | m_p_LU_PulmARevsVel | r_Vel | unilateral | 995300\C3612-01\R-40491\T-48500 | SRT, T-48500, Pulmonary vein | 99PMSBLUS, C3612-01, Atrial Contraction Reversal Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, R-40491, Left Upper Segment
Pediatric Echocardiography Procedure Report | OB | RL PulmV A Vel | Vel | meas | m_p_RL_PulmARevsVel | r_Vel | unilateral | 995300\C3612-01\R-4049E\T-48500 | SRT, T-48500, Pulmonary vein | 99PMSBLUS, C3612-01, Atrial Contraction Reversal Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, R-4049E, Right Lower Segment
Pediatric Echocardiography Procedure Report | OB | RU PulmV A Vel | Vel | meas | m_p_RU_PulmARevsVel | r_Vel | unilateral | 995300\C3612-01\R-404A0\T-48500 | SRT, T-48500, Pulmonary vein | 99PMSBLUS, C3612-01, Atrial Contraction Reversal Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, R-404A0, Right Upper Segment
Pediatric Echocardiography Procedure Report | OB | LL PulmV A Vel | Vel | meas | m_p_LL_PulmARevsVel | r_Vel | unilateral | 995300\C3612-01\R-4214B\T-48500 | SRT, T-48500, Pulmonary vein | 99PMSBLUS, C3612-01, Atrial Contraction Reversal Peak Velocity |  |  | UCUM, mm/s, mm/s | SRT, R-4214B, Left Lower Segment
Pediatric Echocardiography Procedure Report | OB | Pulm A Dur | DopTime | meas | m_p_PulmARevsDur | r_DopTime | unilateral | 995300\C3613-08\T-48500 | SRT, T-48500, Pulmonary vein | 99PMSBLUS, C3613-08, A Wave Duration |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | LU PulmV A Dur | DopTime | meas | m_p_LU_PulmARevsDur | r_DopTime | unilateral | 995300\C3613-08\R-40491\T-48500 | SRT, T-48500, Pulmonary vein | 99PMSBLUS, C3613-08, A Wave Duration |  |  | UCUM, ms, ms | SRT, R-40491, Left Upper Segment
Pediatric Echocardiography Procedure Report | OB | RL PulmV A Dur | DopTime | meas | m_p_RL_PulmARevsDur | r_DopTime | unilateral | 995300\C3613-08\R-4049E\T-48500 | SRT, T-48500, Pulmonary vein | 99PMSBLUS, C3613-08, A Wave Duration |  |  | UCUM, ms, ms | SRT, R-4049E, Right Lower Segment
Pediatric Echocardiography Procedure Report | OB | RU PulmV A Dur | DopTime | meas | m_p_RU_PulmARevsDur | r_DopTime | unilateral | 995300\C3613-08\R-404A0\T-48500 | SRT, T-48500, Pulmonary vein | 99PMSBLUS, C3613-08, A Wave Duration |  |  | UCUM, ms, ms | SRT, R-404A0, Right Upper Segment
Pediatric Echocardiography Procedure Report | OB | LL PulmV A Dur | DopTime | meas | m_p_LL_PulmARevsDur | r_DopTime | unilateral | 995300\C3613-08\R-4214B\T-48500 | SRT, T-48500, Pulmonary vein | 99PMSBLUS, C3613-08, A Wave Duration |  |  | UCUM, ms, ms | SRT, R-4214B, Left Lower Segment
Pediatric Echocardiography Procedure Report | OB | Pulm Dias Vel | Vel | meas | m_p_PulmDiasVel | r_Vel | unilateral | 995300\122202\F-32011\T-48500 | SRT, T-48500, Pulmonary vein | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | LU PulmV Dias Vel | Vel | meas | m_p_LU_PulmDiasVel | r_Vel | unilateral | 995300\122202\F-32011\R-40491\T-48500 | SRT, T-48500, Pulmonary vein | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s | SRT, R-40491, Left Upper Segment |  |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | RL PulmV Dias Vel | Vel | meas | m_p_RL_PulmDiasVel | r_Vel | unilateral | 995300\122202\F-32011\R-4049E\T-48500 | SRT, T-48500, Pulmonary vein | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s | SRT, R-4049E, Right Lower Segment |  |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | RU PulmV Dias Vel | Vel | meas | m_p_RU_PulmDiasVel | r_Vel | unilateral | 995300\122202\F-32011\R-404A0\T-48500 | SRT, T-48500, Pulmonary vein | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s | SRT, R-404A0, Right Upper Segment |  |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | LL PulmV Dias Vel | Vel | meas | m_p_LL_PulmDiasVel | r_Vel | unilateral | 995300\122202\F-32011\R-4214B\T-48500 | SRT, T-48500, Pulmonary vein | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s | SRT, R-4214B, Left Lower Segment |  |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | Pulm Sys Vel | Vel | meas | m_p_PulmSysVel | r_Vel | unilateral | 995300\109070\122204\T-48500 | SRT, T-48500, Pulmonary vein | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LU PulmV Sys Vel | Vel | meas | m_p_LU_PulmSysVel | r_Vel | unilateral | 995300\109070\122204\R-40491\T-48500 | SRT, T-48500, Pulmonary vein | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s | SRT, R-40491, Left Upper Segment |  |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | RL PulmV Sys Vel | Vel | meas | m_p_RL_PulmSysVel | r_Vel | unilateral | 995300\109070\122204\R-4049E\T-48500 | SRT, T-48500, Pulmonary vein | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s | SRT, R-4049E, Right Lower Segment |  |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | RU PulmV Sys Vel | Vel | meas | m_p_RU_PulmSysVel | r_Vel | unilateral | 995300\109070\122204\R-404A0\T-48500 | SRT, T-48500, Pulmonary vein | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s | SRT, R-404A0, Right Upper Segment |  |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LL PulmV Sys Vel | Vel | meas | m_p_LL_PulmSysVel | r_Vel | unilateral | 995300\109070\122204\R-4214B\T-48500 | SRT, T-48500, Pulmonary vein | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s | SRT, R-4214B, Left Lower Segment |  |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LU PulmV Diam | Distance | meas | m_ped_LU_PulmV_Diam | r_Distance | unilateral | 995300\M-02550\R-40491\T-48500 | SRT, T-48500, Pulmonary vein | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, R-40491, Left Upper Segment
Pediatric Echocardiography Procedure Report | OB | RL PulmV Diam | Distance | meas | m_ped_RL_PulmV_Diam | r_Distance | unilateral | 995300\M-02550\R-4049E\T-48500 | SRT, T-48500, Pulmonary vein | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, R-4049E, Right Lower Segment
Pediatric Echocardiography Procedure Report | OB | RU PulmV Diam | Distance | meas | m_ped_RU_PulmV_Diam | r_Distance | unilateral | 995300\M-02550\R-404A0\T-48500 | SRT, T-48500, Pulmonary vein | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, R-404A0, Right Upper Segment
Pediatric Echocardiography Procedure Report | OB | LL PulmV Diam | Distance | meas | m_ped_LL_PulmV_Diam | r_Distance | unilateral | 995300\M-02550\R-4214B\T-48500 | SRT, T-48500, Pulmonary vein | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, R-4214B, Left Lower Segment
Pediatric Echocardiography Procedure Report | OB | Ao Arch Dist Diam | Distance | meas | m_p_AoArch_Dist_Diam | r_Distance | unilateral | 995300\C12242-01\M-02550 | 99PMSBLUS, C12242-01, Distal Aorta | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | Ao ST Jx Diam | Distance | meas | m_p_ST_J_Diam | r_Distance | unilateral | 995300\C12245-01\M-02550 | 99SUP78, C12245-01, Aortic sinotubular junction | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | Ao Sinus Diam | Distance | meas | m_ped_Ao_Sinus_Diam | r_Distance | unilateral | 995300\C12245-02\M-02550 | 99SUP78, C12245-02, Aortic sinus | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | LA Dimen (MM) | Distance | meas | m_ped_LA_Dimen_MM | r_Distance | unilateral | 995300\109070\29469-4\G-0394\T-32300 | SNM3, T-32300, Left Atrium | LN, 29469-4, Left Atrium Antero-posterior Systolic Dimension |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | Desc Ao Max PG | Vmax | meas | m_ped_DSC_Ao_MaxPG | r_Vmax | unilateral | 995300\11726-7\T-42070 | SNM3, T-42070, Thoracic aorta | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Echocardiography Procedure Report | OB | Desc Ao Max PG | MaxPG | meas | m_ped_DSC_Ao_MaxPG | r_MaxPG | unilateral | 995300\20247-3\T-42070 | SNM3, T-42070, Thoracic aorta | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Echocardiography Procedure Report | OB | Desc Ao Mean PG | CardPGmean | meas | m_ped_DSC_Ao_PG | r_CardPGmean | unilateral | 995300\20256-4\T-42070 | SNM3, T-42070, Thoracic aorta | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Echocardiography Procedure Report | OB | Desc Ao Diam | Distance | meas | m_ped_DescAorta_Diam | r_Distance | unilateral | 995300\M-02550\T-42070 | SNM3, T-42070, Thoracic aorta | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | ASC Ao Max PG | Vmax | meas | m_ped_ASC_Ao_MaxPG | r_Vmax | unilateral | 995300\11726-7\T-42100 | SNM3, T-42100, Ascending aorta | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Echocardiography Procedure Report | OB | ASC Ao Max PG | MaxPG | meas | m_ped_ASC_Ao_MaxPG | r_MaxPG | unilateral | 995300\20247-3\T-42100 | SNM3, T-42100, Ascending aorta | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Echocardiography Procedure Report | OB | Asc Ao Mean PG | CardPGmean | meas | m_ped_ASC_Ao_PG | r_CardPGmean | unilateral | 995300\20256-4\T-42100 | SNM3, T-42100, Ascending aorta | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Echocardiography Procedure Report | OB | Asc Ao Diam | Distance | meas | m_ped_AscAorta_Diam | r_Distance | unilateral | 995300\M-02550\T-42100 | SNM3, T-42100, Ascending aorta | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | Ao Arch Diam | Distance | meas | m_ped_AoArch_Diam | r_Distance | unilateral | 995300\M-02550\T-42300 | SNM3, T-42300, Aortic arch | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | LAD | Distance | meas | m_ped_LAD_Diam | r_Distance | unilateral | 995300\M-02550\T-4311A | SNM3, T-4311A, Left Anterior Descending Coronary Artery | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | Cx | Distance | meas | m_ped_LCx_Diam | r_Distance | unilateral | 995300\M-02550\T-43120 | SNM3, T-43120, Circumflex Coronary Artery | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | RCA | Distance | meas | m_ped_RCA_Diam | r_Distance | unilateral | 995300\M-02550\T-43203 | SNM3, T-43203, Right Coronary Artery | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | Coarctation Diam | Distance | meas | m_ped_Coarc_Diam | r_Distance | unilateral | 995300\D4-32014\M-02550 | SRT, D4-32014, Coarctation of aorta | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | AoR Diam (MM) | Distance | meas | m_ped_AoR_Diam_MM | r_Distance | unilateral | 995300\F-04403\G-0394\M-02550 | SRT, F-04403, Aortic root | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | AoR Diam (2D) | Distance | meas | m_ped_AoR_Diam_2D | r_Distance | unilateral | 995300\F-04403\G-03A2\M-02550 | SRT, F-04403, Aortic root | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | Ao Isthmus Diam | Distance | meas | m_ped_AoIsth_Diam | r_Distance | unilateral | 995300\M-02550\T-42310 | SRT, T-42310, Aortic isthmus | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | Left Main | Distance | meas | m_ped_L_Main_Diam | r_Distance | unilateral | 995300\M-02550\T-43107 | SRT, T-43107, Left Main Coronary Artery | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | PD | Distance | meas | m_ped_PDC_Diam | r_Distance | unilateral | 995300\M-02550\T-43210 | SRT, T-43210, Posterior Descending Right Coronary Artery | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | PDA Dias Vel | Vel | meas | m_p_PDA_DiasVel | r_Vel | unilateral | 995300\122202\D4-32012\F-32011 | SNM3, D4-32012, Patent ductus arteriosus | DCM, 122202, Diastolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | PDA Sys Vel | Vel | meas | m_p_PDA_SysVel | r_Vel | unilateral | 995300\109070\122204\D4-32012 | SNM3, D4-32012, Patent ductus arteriosus | DCM, 122204, Systolic blood velocity, peak |  |  | UCUM, mm/s, mm/s |  |  |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | PDA Diam | Distance | meas | m_ped_PDA_Diam | r_Distance | unilateral | 995300\D4-32012\M-02550 | SNM3, D4-32012, Patent ductus arteriosus | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | RA Pressure | UserPressure | meas | m_ped_RA_Press | r_UserPressure | unilateral | 995300\109070\C12220-09\T-32200 | SNM3, T-32200, Right Atrium | 99PMSBLUS, C12220-09, Pressure |  |  | UCUM, mmHg, mmHg |  |  |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | RA Dimen (2D) | Distance | meas | m_ped_RA_Dimen_2D | r_Distance | unilateral | 995300\121206\G-03A2\T-32200 | SNM3, T-32200, Right Atrium | DCM, 121206, Distance |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | RA Dimen (MM) | Distance | meas | m_ped_RA_Dimen_MM | r_Distance | unilateral | 995300\G-0394\M-02550\T-32200 | SNM3, T-32200, Right Atrium | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | RAes Major - A4C | Distance | meas | m_ped_RAES_MAJOR_A4C | r_Distance | unilateral | 995300\109070\G-A193\G-A19C\T-32200 | SNM3, T-32200, Right Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | RAed Major - A4C | Distance | meas | m_ped_RAED_MAJOR_A4C | r_Distance | unilateral | 995300\F-32011\G-A193\G-A19C\T-32200 | SNM3, T-32200, Right Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | RAes Minor - A4C | Distance | meas | m_ped_RAES_MINOR_A4C | r_Distance | unilateral | 995300\109070\G-A194\G-A19C\T-32200 | SNM3, T-32200, Right Atrium | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | RAed Minor - A4C | Distance | meas | m_ped_RAED_MINOR_A4C | r_Distance | unilateral | 995300\F-32011\G-A194\G-A19C\T-32200 | SNM3, T-32200, Right Atrium | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | LA Dimen (2D) | Distance | meas | m_ped_LA_Dimen_2D | r_Distance | unilateral | 995300\121206\G-03A2\T-32300 | SNM3, T-32300, Left Atrium | DCM, 121206, Distance |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | LAes Major - A4C | Distance | meas | m_ped_LAES_MAJOR_A4C | r_Distance | unilateral | 995300\109070\G-A193\G-A19C\T-32300 | SNM3, T-32300, Left Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LAed Major - A4C | Distance | meas | m_ped_LAED_MAJOR_A4C | r_Distance | unilateral | 995300\F-32011\G-A193\G-A19C\T-32300 | SNM3, T-32300, Left Atrium | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | LAes Minor - A4C | Distance | meas | m_ped_LAES_MINOR_A4C | r_Distance | unilateral | 995300\109070\G-A194\G-A19C\T-32300 | SNM3, T-32300, Left Atrium | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LAed Minor - A4C | Distance | meas | m_ped_LAED_MINOR_A4C | r_Distance | unilateral | 995300\F-32011\G-A194\G-A19C\T-32300 | SNM3, T-32300, Left Atrium | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | ASD VTI | Vmax | meas | m_p_ASD_VTI | r_Vmax | unilateral | 995300\11726-7\D4-31220 | SRT, D4-31220, Atrial Septal Defect | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s
Pediatric Echocardiography Procedure Report | OB | ASD VTI | MaxPG | meas | m_p_ASD_VTI | r_MaxPG | unilateral | 995300\20247-3\D4-31220 | SRT, D4-31220, Atrial Septal Defect | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg
Pediatric Echocardiography Procedure Report | OB | ASD VTI | CardPGmean | meas | m_p_ASD_VTI | r_CardPGmean | unilateral | 995300\20256-4\D4-31220\R-42047 | SRT, D4-31220, Atrial Septal Defect | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | ASD VTI | CardTAPV | meas | m_p_ASD_VTI | r_CardTAPV | unilateral | 995300\20352-1\D4-31220\R-42047 | SRT, D4-31220, Atrial Septal Defect | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | ASD VTI | VTI | meas | m_p_ASD_VTI | r_VTI | unilateral | 995300\20354-7\D4-31220\R-42047 | SRT, D4-31220, Atrial Septal Defect | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | ASD Major | Distance | meas | m_ped_ASD_Major | r_Distance | unilateral | 995300\D4-31220\G-A193 | SRT, D4-31220, Atrial Septal Defect | SRT, G-A193, Major Axis |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | ASD Minor | Distance | meas | m_ped_ASD_Minor | r_Distance | unilateral | 995300\D4-31220\G-A194 | SRT, D4-31220, Atrial Septal Defect | SRT, G-A194, Minor Axis |  |  | UCUM, mm, mm
Pediatric Echocardiography Procedure Report | OB | TV Radius | Distance | meas | m_ped_TV_Radius | r_Distance | unilateral | 995300\C12222-01\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12222-01, Flow Radius |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | TR Radius | Distance | meas | m_ped_TR_Radius | r_Distance | unilateral | 995300\C12222-01\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12222-01, Flow Radius |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | TV Alias Vel | 2DVel | meas | m_ped_TV_AliasVel | r_2DVel | unilateral | 995300\C12222-02\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12222-02, Alias velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | TR Alias Vel | 2DVel | meas | m_ped_TR_AliasVel | r_2DVel | unilateral | 995300\C12222-02\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12222-02, Alias velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | TV P½t | Vmax | meas | m_ped_TV_PHalfT | r_Vmax | unilateral | 995300\C12222-03\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12222-03, Pressure Half-Time Peak velocity |  |  | UCUM, mm/s, mm/s
Pediatric Echocardiography Procedure Report | OB | TV Accel | DopSlope | meas | m_ped_TV_Accel | r_DopSlope | unilateral | 995300\C12222-04\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12222-04, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | TV D-E Slope | MMSlope | meas | m_ped_TV_DE_Slope | r_MMSlope | unilateral | 995300\C12222-05\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12222-05, D-E Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | TV E-F Slope | MMSlope | meas | m_ped_TV_EF_Slope | r_MMSlope | unilateral | 995300\C12222-06\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12222-06, E-F Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | TV A-C Interval | MMTime | meas | m_ped_TV_AC_Int | r_MMTime | unilateral | 995300\C12222-07\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12222-07, A-C Interval |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | TV Peak E Vel | Vel | meas | m_ped_TV_PeakEVel | r_Vel | unilateral | 995300\C3612-02\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C3612-02, E Wave Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | TV Peak A Vel | Vel | meas | m_ped_TV_PeakAVel | r_Vel | unilateral | 995300\C3612-03\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C3612-03, A Wave Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | TV Closure to Opening | DopTime | meas | m_ped_TV_Cl_to_Op | r_DopTime | unilateral | 995300\C3613-01\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C3613-01, Closure to Opening Time |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | TV A Dur | DopTime | meas | m_ped_TV_A_Dur | r_DopTime | unilateral | 995300\C3613-08\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C3613-08, A Wave Duration |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | TV D-E Exc | Distance | meas | m_ped_TV_DE_Exc | r_Distance | unilateral | 995300\C7470-27\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C7470-27, D-E Excursion |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | TV Vmax | Vmax | meas | m_ped_TV_Vmax | r_Vmax | unilateral | 995300\11726-7\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | TR Vmax | Vmax | meas | m_ped_TR_Vmax | r_Vmax | unilateral | 995300\11726-7\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | TV Accel | DopTime | meas | m_ped_TV_Accel | r_DopTime | unilateral | 995300\20168-1\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | TV Decel Slope | DopSlope | meas | m_ped_TV_DecelSlope | r_DopSlope | unilateral | 995300\20216-8\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20216-8, Deceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | TV Decel Time | DopTime | meas | m_ped_TV_DecelTime | r_DopTime | unilateral | 995300\20217-6\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | TV Vmax | MaxPG | meas | m_ped_TV_Vmax | r_MaxPG | unilateral | 995300\20247-3\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | TR Vmax | MaxPG | meas | m_ped_TR_Vmax | r_MaxPG | unilateral | 995300\20247-3\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | TV Mean PG | CardPGmean | meas | m_ped_TV_MeanPG | r_CardPGmean | unilateral | 995300\20256-4\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | TR VTI | CardPGmean | meas | m_ped_TR_VTI | r_CardPGmean | unilateral | 995300\20256-4\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | TV P½t | PHalftime | meas | m_ped_TV_PHalfT | r_PHalftime | unilateral | 995300\20280-4\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20280-4, Pressure Half-Time |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | TV VTI | CardTAPV | meas | m_ped_TV_VTI | r_CardTAPV | unilateral | 995300\20352-1\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | A4Cd | Seg11 | meas | m_ped_A4Cd | r_Seg11 | unilateral | 995300\125208\F-32011\G-A19C\M-02550\T-32600\T5203-01 | SRT, T-32600, Left Ventricle | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  |  | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | TR VTI | CardTAPV | meas | m_ped_TR_VTI | r_CardTAPV | unilateral | 995300\20352-1\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20352-1, Mean Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | TV VTI | VTI | meas | m_ped_TV_VTI | r_VTI | unilateral | 995300\20354-7\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | TR VTI | VTI | meas | m_ped_TR_VTI | r_VTI | unilateral | 995300\20354-7\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20354-7, Velocity Time Integral |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | HR - TV | HeartRate | meas | m_ped_HR_TV | r_HeartRate | unilateral | 995300\8867-4\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 8867-4, Heart Rate |  |  | UCUM, {H.B.}/min, bpm
Pediatric Echocardiography Procedure Report | OB | TV Annul Diam | Distance | meas | m_ped_TV_Annul_Diam | r_Distance | unilateral | 995300\M-02550\T-35100\T-35111 | SRT, T-35100, Tricuspid Valve | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm | SRT, T-35111, Tricuspid Annulus
Pediatric Echocardiography Procedure Report | OB | TV Diam | Distance | meas | m_ped_TV_Diam | r_Distance | unilateral | 995300\G-038F\G-03A2\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038F, Cardiovascular Orifice Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | TV DFP | DopTime | meas | m_ped_TV_DFP | r_DopTime | unilateral | 995300\R-003A9\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, R-003A9, Tricuspid Diastolic Filling Period (DFPt) |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | MV E-E Sep | Distance | meas | m_ped_MV_EE_Separ | r_Distance | unilateral | 995300\C12207-03\G-0394\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-03, Mitral Valve E-E Separation |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | MV Radius | Distance | meas | m_ped_MV_Radius | r_Distance | unilateral | 995300\C12222-01\R-42047\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12222-01, Flow Radius |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | MR Radius | Distance | meas | m_ped_MR_Radius | r_Distance | unilateral | 995300\C12222-01\R-42E61\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12222-01, Flow Radius |  |  | UCUM, mm, mm |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | MV Alias Vel | 2DVel | meas | m_ped_MV_AliasVel | r_2DVel | unilateral | 995300\C12222-02\R-42047\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12222-02, Alias velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | MR Alias Vel | 2DVel | meas | m_ped_MR_AliasVel | r_2DVel | unilateral | 995300\C12222-02\R-42E61\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12222-02, Alias velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | MVA - P½t | Vmax | meas | m_ped_MV_PHalfT | r_Vmax | unilateral | 995300\C12222-03\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12222-03, Pressure Half-Time Peak velocity |  |  | UCUM, mm/s, mm/s
Pediatric Echocardiography Procedure Report | OB | MV Accel | DopSlope | meas | m_ped_MV_Accel | r_DopSlope | unilateral | 995300\C12222-04\R-42047\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12222-04, Acceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | MV D-E Slope | MMSlope | meas | m_ped_MV_DE_Slope | r_MMSlope | unilateral | 995300\C12222-05\G-0394\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12222-05, D-E Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | MV E-F Slope | MMSlope | meas | m_ped_MV_EF_Slope | r_MMSlope | unilateral | 995300\C12222-06\G-0394\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12222-06, E-F Slope |  |  | UCUM, mm/s, mm/s |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | MV A-C Interval | MMTime | meas | m_ped_MV_AC_Int | r_MMTime | unilateral | 995300\C12222-07\G-0394\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12222-07, A-C Interval |  |  | UCUM, ms, ms |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | MV Peak E Vel | Vel | meas | m_ped_MV_PeakEVel | r_Vel | unilateral | 995300\C3612-02\R-42047\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C3612-02, E Wave Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | MV Peak A Vel | Vel | meas | m_ped_MV_PeakAVel | r_Vel | unilateral | 995300\C3612-03\R-42047\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C3612-03, A Wave Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | MV Closure to Opening | DopTime | meas | m_ped_MV_Cl_to_Op | r_DopTime | unilateral | 995300\C3613-01\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C3613-01, Closure to Opening Time |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | MV A Dur | DopTime | meas | m_ped_MV_A_Dur | r_DopTime | unilateral | 995300\C3613-08\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C3613-08, A Wave Duration |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | MV D-E Exc | Distance | meas | m_ped_MV_DE_Exc | r_Distance | unilateral | 995300\C7470-27\G-0394\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C7470-27, D-E Excursion |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | MV Vmax | Vmax | meas | m_ped_MV_Vmax | r_Vmax | unilateral | 995300\11726-7\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | MR Vmax | Vmax | meas | m_ped_MR_Vmax | r_Vmax | unilateral | 995300\11726-7\R-42E61\T-35300 | SRT, T-35300, Mitral Valve | LN, 11726-7, Peak Velocity |  |  | UCUM, mm/s, mm/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | MV EPSS | Distance | meas | m_ped_MV_EPSS | r_Distance | unilateral | 995300\18036-4\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 18036-4, Mitral Valve EPSS, E wave |  |  | UCUM, mm, mm |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | MV Accel | DopTime | meas | m_ped_MV_Accel | r_DopTime | unilateral | 995300\20168-1\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20168-1, Acceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | MV Decel Slope | DopSlope | meas | m_ped_MV_DecelSlope | r_DopSlope | unilateral | 995300\20216-8\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20216-8, Deceleration Slope |  |  | UCUM, mm/s2, mm/s2 |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | MV Decel Time | DopTime | meas | m_ped_MV_DecelTime | r_DopTime | unilateral | 995300\20217-6\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20217-6, Deceleration Time |  |  | UCUM, ms, ms |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | MV Vmax | MaxPG | meas | m_ped_MV_Vmax | r_MaxPG | unilateral | 995300\20247-3\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Echocardiography Procedure Report | OB | SVC S/D | none | calc | c_p_SVC_SD | none | unilateral | 995300\12144-2\C3010-03 | 99PMSBLUS, C3010-03, Right Superior vena cava | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Pediatric Echocardiography Procedure Report | OB | Hepatic S/D | none | calc | c_p_HepSD | none | unilateral | 995300\12144-2\T5200-01 | 99PMSBLUS, T5200-01, Hepatic Veins | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Pediatric Echocardiography Procedure Report | OB | IVC S/D | none | calc | c_p_IVC_SD | none | unilateral | 995300\12144-2\T-48710 | SNM3, T-48710, Inferior vena cava | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Pediatric Echocardiography Procedure Report | OB | Pulm S/D | none | calc | c_p_PulmSD | none | unilateral | 995300\12144-2\T-48500 | SRT, T-48500, Pulmonary vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units
Pediatric Echocardiography Procedure Report | OB | LU PulmV S/D | none | calc | c_p_LU_PulmSD | none | unilateral | 995300\12144-2\R-40491\T-48500 | SRT, T-48500, Pulmonary vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units | SRT, R-40491, Left Upper Segment
Pediatric Echocardiography Procedure Report | OB | RL PulmV S/D | none | calc | c_p_RL_PulmSD | none | unilateral | 995300\12144-2\R-4049E\T-48500 | SRT, T-48500, Pulmonary vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units | SRT, R-4049E, Right Lower Segment
Pediatric Echocardiography Procedure Report | OB | RU PulmV S/D | none | calc | c_p_RU_PulmSD | none | unilateral | 995300\12144-2\R-404A0\T-48500 | SRT, T-48500, Pulmonary vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units | SRT, R-404A0, Right Upper Segment
Pediatric Echocardiography Procedure Report | OB | LL PulmV S/D | none | calc | c_p_LL_PulmSD | none | unilateral | 995300\12144-2\R-4214B\T-48500 | SRT, T-48500, Pulmonary vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  |  | UCUM, 1, no units | SRT, R-4214B, Left Lower Segment
Pediatric Echocardiography Procedure Report | OB | Desc Ao Max PG (full) | none | calc | c_p_DscAo_MaxPG_FB | none | unilateral | 995300\125217\20247-3\T-42070 | SNM3, T-42070, Thoracic aorta | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Echocardiography Procedure Report | OB | Desc Ao Mean PG (full) | none | calc | c_p_DscAo_MeanPG_FB | none | unilateral | 995300\125217\20256-4\T-42070 | SNM3, T-42070, Thoracic aorta | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Echocardiography Procedure Report | OB | Asc Ao Max PG (full) | none | calc | c_p_AscAo_MaxPG_FB | none | unilateral | 995300\125217\20247-3\T-42100 | SNM3, T-42100, Ascending aorta | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Echocardiography Procedure Report | OB | Asc Ao Mean PG (full) | none | calc | c_p_AscAo_MeanPG_FB | none | unilateral | 995300\125217\20256-4\T-42100 | SNM3, T-42100, Ascending aorta | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Echocardiography Procedure Report | OB | AoR Area | none | calc | c_ped_AoR_Area | none | unilateral | 995300\F-04403\G-038E\G-03A2 | SRT, F-04403, Aortic root | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | LA/Ao (MM) | none | calc | c_ped_LA_Ao_MM | none | unilateral | 995300\17985-3\G-0394\T-32300 | SNM3, T-32300, Left Atrium | LN, 17985-3, Left Atrium to Aortic Root Ratio |  |  | UCUM, 1, no units |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | LA/Ao (2D) | none | calc | c_ped_LA_Ao_2D | none | unilateral | 995300\17985-3\G-03A2\T-32300 | SNM3, T-32300, Left Atrium | LN, 17985-3, Left Atrium to Aortic Root Ratio |  |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | ASD Diam | none | calc | c_p_ASD_Diam | none | unilateral | 995300\121428\D4-31220\G-03A2\M-02550 | SRT, D4-31220, Atrial Septal Defect | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | [7 empty] | DCM, 121428, Calculated
Pediatric Echocardiography Procedure Report | OB | TV E/A | none | calc | c_ped_TV_EA | none | unilateral | 995300\C12222-08\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12222-08, E to A Ratio |  |  | UCUM, 1, no units
Pediatric Echocardiography Procedure Report | OB | PISA (TR) | none | calc | c_ped_PISA_TR | none | unilateral | 995300\125216\C7471-06\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C7471-06, Flow Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Echocardiography Procedure Report | OB | TR Volume | none | calc | c_ped_TR_Vol | none | unilateral | 995300\125216\33878-0\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 33878-0, Volume Flow |  |  | UCUM, mm3, mm3 |  |  |  |  | SRT, R-42E61, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Echocardiography Procedure Report | OB | TR Flow Rate | none | calc | c_ped_TR_FlowRate | none | unilateral | 995300\34141-2\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 34141-2, Peak Instantaneous Flow Rate |  |  | UCUM, mm3/s, mm3/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | TV Area | none | calc | c_ped_TV_Area | none | unilateral | 995300\G-03A2\G-A166\T-35100 | SRT, T-35100, Tricuspid Valve | SNM3, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | CO (TV) | none | calc | c_ped_CO_TV | none | unilateral | 995300\F-32100\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s
Pediatric Echocardiography Procedure Report | OB | SV (TV) | none | calc | c_ped_SV_TV | none | unilateral | 995300\F-32120\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3
Pediatric Echocardiography Procedure Report | OB | TVA (PISA) | none | calc | c_ped_TVA_PISA | none | unilateral | 995300\125216\G-038E\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Echocardiography Procedure Report | OB | TR ERO | none | calc | c_ped_TR_ERO | none | unilateral | 995300\125216\G-038E\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Echocardiography Procedure Report | OB | TV Annul Area | none | calc | c_ped_TV_Annul_Area | none | unilateral | 995300\G-038E\G-03A2\T-35100\T-35111 | SRT, T-35100, Tricuspid Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | SRT, T-35111, Tricuspid Annulus | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | TR Fraction | none | calc | c_ped_TR_Frac | none | unilateral | 995300\G-0390\R-42E61\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-0390, Regurgitant Fraction |  |  | UCUM, %, Percent |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | Tei Index | none | calc | c_ped_TeiIndex | none | unilateral | 995300\C12207-05\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-05, Tei Index |  |  | UCUM, 1, no units
Pediatric Echocardiography Procedure Report | OB | MV E/A | none | calc | c_ped_MV_EA | none | unilateral | 995300\C12222-08\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12222-08, E to A Ratio |  |  | UCUM, 1, no units
Pediatric Echocardiography Procedure Report | OB | PISA (MR) | none | calc | c_ped_PISA_MR | none | unilateral | 995300\125216\C7471-06\R-42E61\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C7471-06, Flow Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Echocardiography Procedure Report | OB | MR Volume | none | calc | c_ped_MR_Vol | none | unilateral | 995300\125216\33878-0\R-42E61\T-35300 | SRT, T-35300, Mitral Valve | LN, 33878-0, Volume Flow |  |  | UCUM, mm3, mm3 |  |  |  |  | SRT, R-42E61, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Echocardiography Procedure Report | OB | MR Flow Rate | none | calc | c_ped_MR_FlowRate | none | unilateral | 995300\34141-2\R-42E61\T-35300 | SRT, T-35300, Mitral Valve | LN, 34141-2, Peak Instantaneous Flow Rate |  |  | UCUM, mm3/s, mm3/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | MV Area | none | calc | c_ped_MV_Area | none | unilateral | 995300\G-03A2\G-A166\T-35300 | SRT, T-35300, Mitral Valve | SNM3, G-A166, Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | CO (MV) | none | calc | c_ped_CO_MV | none | unilateral | 995300\F-32100\T-35300 | SRT, T-35300, Mitral Valve | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s
Pediatric Echocardiography Procedure Report | OB | SV (MV) | none | calc | c_ped_SV_MV | none | unilateral | 995300\F-32120\T-35300 | SRT, T-35300, Mitral Valve | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3
Pediatric Echocardiography Procedure Report | OB | MVA (P½t ) | none | calc | c_ped_MVA_PHalfT | none | unilateral | 995300\125210\G-038E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125210, Area by Pressure Half-Time
Pediatric Echocardiography Procedure Report | OB | MVA (VTI) | none | calc | c_ped_MVA_VTI | none | unilateral | 995300\125215\G-038E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125215, Continuity Equation by Velocity Time Integral
Pediatric Echocardiography Procedure Report | OB | MVA (PISA) | none | calc | c_ped_MVA_PISA | none | unilateral | 995300\125216\G-038E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Echocardiography Procedure Report | OB | MR ERO | none | calc | c_ped_MR_ERO | none | unilateral | 995300\125216\G-038E\R-42E61\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Echocardiography Procedure Report | OB | MV Area (Ellipse) | none | calc | c_ped_MV_Area_Ellps | none | unilateral | 995300\125211\G-038E\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125211, Biplane Ellipse
Pediatric Echocardiography Procedure Report | OB | MV Annul Area | none | calc | c_ped_MV_Annul_Area | none | unilateral | 995300\G-038E\G-03A2\T-35300\T-35313 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | SRT, T-35313, Mitral Annulus | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | MR Fraction | none | calc | c_ped_MR_Frac | none | unilateral | 995300\G-0390\R-42E61\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-0390, Regurgitant Fraction |  |  | UCUM, %, Percent |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | RVOT Area | none | calc | c_ped_RVOT_Area | none | unilateral | 995300\G-038E\G-03A2\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | SNM3, T-32550, Right Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | LVOT Area | none | calc | c_ped_LVOT_Area | none | unilateral | 995300\G-038E\G-03A2\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | SNM3, T-32650, Left Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | MPA Area | none | calc | c_ped_MPA_Area | none | unilateral | 995300\C3010-01\G-038E\G-03A2 | 99PMSBLUS, C3010-01, Main pulmonary artery | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | PAP (AT) | none | calc | c_ped_PAP_AT | none | unilateral | 995300\C12210-01\T-35200 | SRT, T-35200, Pulmonic Valve | 99PMSBLUS, C12210-01, Pulmonary Artery Pressure using Accel Time |  |  | UCUM, mmHg, mmHg
Pediatric Echocardiography Procedure Report | OB | PVA (Vmax) | none | calc | c_ped_PVA_Vmax | none | unilateral | 995300\125214\G-038E\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Echocardiography Procedure Report | OB | PVA (VTI) | none | calc | c_ped_PVA_VTI | none | unilateral | 995300\125215\G-038E\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125215, Continuity Equation by Velocity Time Integral
Pediatric Echocardiography Procedure Report | OB | LVSP (AS) | none | calc | c_ped_LVSP_AS | none | unilateral | 995300\109070\C12212-01\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12212-01, Left Ventricle Systolic Pressure with Aortic Stenosis |  |  | UCUM, mmHg, mmHg |  |  |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | LVDP (AI) | none | calc | c_ped_LVDP_AI | none | unilateral | 995300\C12212-02\F-32011\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12212-02, Left Ventricle Diastolic Pressure with Aortic Insufficiency |  |  | UCUM, mmHg, mmHg |  |  |  | SRT, F-32011, End Diastole
Pediatric Echocardiography Procedure Report | OB | PISA (AI) | none | calc | c_ped_PISA_AI | none | unilateral | 995300\125216\C7471-06\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C7471-06, Flow Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Echocardiography Procedure Report | OB | AV Max PG (full) | none | calc | c_p_AV_MaxPG_FB | none | unilateral | 995300\125217\20247-3\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  |  | UCUM, mmHg, mmHg | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Echocardiography Procedure Report | OB | AV Mean PG (full) | none | calc | c_p_AV_MeanPG_FB | none | unilateral | 995300\125217\20256-4\T-35400 | SRT, T-35400, Aortic Valve | LN, 20256-4, Mean Gradient |  |  | UCUM, mmHg, mmHg | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Echocardiography Procedure Report | OB | AI Volume | none | calc | c_ped_AI_Vol | none | unilateral | 995300\125216\33878-0\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 33878-0, Volume Flow |  |  | UCUM, mm3, mm3 |  |  |  |  | SRT, R-42E61, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Echocardiography Procedure Report | OB | AI Flow Rate | none | calc | c_ped_AI_FlowRate | none | unilateral | 995300\34141-2\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | LN, 34141-2, Peak Instantaneous Flow Rate |  |  | UCUM, mm3/s, mm3/s |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | AVA (Vmax) | none | calc | c_ped_AVA_Vmax | none | unilateral | 995300\125214\G-038E\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Echocardiography Procedure Report | OB | AVA (VTI) | none | calc | c_ped_AVA_VTI | none | unilateral | 995300\125215\G-038E\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 | [5 empty] | DCM, 125215, Continuity Equation by Velocity Time Integral
Pediatric Echocardiography Procedure Report | OB | AI ERO | none | calc | c_ped_AI_ERO | none | unilateral | 995300\125216\G-038E\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  |  | UCUM, mm2, mm2 |  |  |  |  | SRT, R-42E61, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Echocardiography Procedure Report | OB | AI Fraction | none | calc | c_ped_AI_Frac | none | unilateral | 995300\G-0390\R-42E61\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-0390, Regurgitant Fraction |  |  | UCUM, %, Percent |  |  |  |  | SRT, R-42E61, Regurgitant Flow
Pediatric Echocardiography Procedure Report | OB | IVS % (MM) | none | calc | c_ped_IVS_MM | none | unilateral | 995300\C7470-28\G-0394\T-32410 | SRT, T-32410, Interventricular septum | 99PMSBLUS, C7470-28, % Thickening |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | IVS % (2D) | none | calc | c_ped_IVS | none | unilateral | 995300\C7470-28\G-03A2\T-32410 | SRT, T-32410, Interventricular septum | 99PMSBLUS, C7470-28, % Thickening |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | RVSP | none | calc | c_ped_RVSP | none | unilateral | 995300\109070\C12220-09\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12220-09, Pressure |  |  | UCUM, mmHg, mmHg |  |  |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | RVPEPc | none | calc | c_ped_RVPEPc | none | unilateral | 995300\C3613-07\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C3613-07, Heart Rate-Corrected Pre-Ejection Time |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | RV PEP/ET | none | calc | c_ped_RV_PEP_ET | none | unilateral | 995300\C3613-09\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C3613-09, Ratio of Pre-Ejection Time to Ejection Time |  |  | UCUM, 1, no units
Pediatric Echocardiography Procedure Report | OB | CO (RVOT) | none | calc | c_ped_CO_RVOT | none | unilateral | 995300\F-32100\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s | SNM3, T-32550, Right Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | SV (RVOT) | none | calc | c_ped_SV_RVOT | none | unilateral | 995300\F-32120\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 | SNM3, T-32550, Right Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | RV MPI | none | calc | c_ped_RV_MPI | none | unilateral | 995300\G-0381\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-0381, Right Ventricular Index of Myocardial Performance |  |  | UCUM, 1, no units
Pediatric Echocardiography Procedure Report | OB | LV Mass Index(Cubed) | none | calc | c_ped_LV_Mass_BSA | none | unilateral | 995300\125206\C12203-01\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-01, Left Ventricle Mass Index |  |  | UCUM, g/m2, g/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | LV Mass Index (A/L) | none | calc | c_ped_LVMass_A_LBSA | none | unilateral | 995300\C12203-01\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-01, Left Ventricle Mass Index |  |  | UCUM, g/m2, g/m2 |  | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | E`/A` Medial | none | calc | c_p_Med_E_A | none | unilateral | 995300\C12203-09\G-0391\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-09, Ratio of LV E to A Tissue Velocity |  |  | UCUM, 1, no units | SRT, G-0391, Medial Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | E`/A` Lateral | none | calc | c_p_Lat_E_A | none | unilateral | 995300\C12203-09\G-0392\T12224-02\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-09, Ratio of LV E to A Tissue Velocity |  |  | UCUM, 1, no units | SRT, G-0392, Lateral Mitral Annulus | 99PMSBLUS, T12224-02, Tissue Doppler Imaging
Pediatric Echocardiography Procedure Report | OB | Wall Stress | none | calc | c_ped_Wall_Stress | none | unilateral | 995300\C12203-37\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-37, Left Ventricle Meridional Wall Stress |  |  | UCUM, g/m2, g/m2 |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | Mean VCF | none | calc | c_ped_Mean_VCF | none | unilateral | 995300\C12203-38\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-38, Mean Velocity of Circumferential Fiber Shortening |  |  | 99PMSBLUS, circ/s, circ/s |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | Mean VCFc | none | calc | c_ped_Mean_VCFc | none | unilateral | 995300\C12203-39\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-39, HR-Corrected Mean Velocity of Circumferential Fiber Shortening |  |  | 99PMSBLUS, circ/s, circ/s |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | FS (MM-Cubed) | none | calc | c_ped_FS_Cubed_MM | none | unilateral | 995300\125206\C12239-02\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12239-02, Fractional Shortening |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | FS (MM-Teich) | none | calc | c_ped_FS_Teich_MM | none | unilateral | 995300\125209\C12239-02\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12239-02, Fractional Shortening |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Echocardiography Procedure Report | OB | FS (2D-Cubed) | none | calc | c_ped_FS_Cubed | none | unilateral | 995300\125206\C12239-02\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12239-02, Fractional Shortening |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | FS (2D-Teich) | none | calc | c_ped_FS_Teich | none | unilateral | 995300\125209\C12239-02\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12239-02, Fractional Shortening |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Echocardiography Procedure Report | OB | EF (MM-Cubed) | none | calc | c_ped_EF_Cubed_MM | none | unilateral | 995300\125206\C3467-01\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3467-01, Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | EF (MM-Teich) | none | calc | c_ped_EF_Teich_MM | none | unilateral | 995300\125209\C3467-01\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3467-01, Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Echocardiography Procedure Report | OB | EF (2D-Cubed) | none | calc | c_ped_EF_Cubed | none | unilateral | 995300\125206\C3467-01\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3467-01, Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | EF (BP) | none | calc | c_ped_EF_BP | none | unilateral | 995300\125207\C3467-01\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3467-01, Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Echocardiography Procedure Report | OB | EF (2D-Teich) | none | calc | c_ped_EF_Teich | none | unilateral | 995300\125209\C3467-01\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3467-01, Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Echocardiography Procedure Report | OB | EF (A2C) | none | calc | c_ped_EF_AP2 | none | unilateral | 995300\125208\C3467-01\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3467-01, Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | EF (A2C-A/L) | none | calc | c_ped_EF_A2C_SPE | none | unilateral | 995300\125226\C3467-01\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3467-01, Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Echocardiography Procedure Report | OB | EF (A4C) | none | calc | c_ped_EF_AP4 | none | unilateral | 995300\125208\C3467-01\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3467-01, Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | EF (A4C-A/L) | none | calc | c_ped_EF_A4C_SPE | none | unilateral | 995300\125226\C3467-01\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3467-01, Ejection Fraction |  |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Echocardiography Procedure Report | OB | LVETc | none | calc | c_ped_LVETc | none | unilateral | 995300\C3613-06\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3613-06, Heart Rate-Corrected Ejection Time |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | LVPEPc | none | calc | c_ped_LVPEPc | none | unilateral | 995300\C3613-07\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3613-07, Heart Rate-Corrected Pre-Ejection Time |  |  | UCUM, ms, ms
Pediatric Echocardiography Procedure Report | OB | LV PEP/ET | none | calc | c_ped_LV_PEP_ET | none | unilateral | 995300\C3613-09\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3613-09, Ratio of Pre-Ejection Time to Ejection Time |  |  | UCUM, 1, no units
Pediatric Echocardiography Procedure Report | OB | LVPW % (MM) | none | calc | c_ped_LVPW_MM | none | unilateral | 995300\C12243-01\C7470-28\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C7470-28, % Thickening |  |  | UCUM, %, Percent | 99PMSBLUS, C12243-01, Left Ventricle Posterior Wall | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | LVPW % (2D) | none | calc | c_ped_LVPW | none | unilateral | 995300\C12243-01\C7470-28\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C7470-28, % Thickening |  |  | UCUM, %, Percent | 99PMSBLUS, C12243-01, Left Ventricle Posterior Wall | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | LV Mass (Cubed) | none | calc | c_ped_LV_Mass | none | unilateral | 995300\125206\18087-7\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18087-7, Left Ventricle Mass |  |  | UCUM, g, grams |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | LV Mass (A/L) | none | calc | c_ped_LVMass_A_L | none | unilateral | 995300\18087-7\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18087-7, Left Ventricle Mass |  |  | UCUM, g, grams |  | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | IVS/LVPW (MM) | none | calc | c_ped_IVS_LVPW_MM | none | unilateral | 995300\18155-2\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18155-2, Interventricular Septum to Posterior Wall Thickness Ratio |  |  | UCUM, 1, no units |  | SRT, G-0394, M mode
Pediatric Echocardiography Procedure Report | OB | IVS/LVPW (2D) | none | calc | c_ped_IVS_LVPW | none | unilateral | 995300\18155-2\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18155-2, Interventricular Septum to Posterior Wall Thickness Ratio |  |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Pediatric Echocardiography Procedure Report | OB | ESV (MM-Cubed) | none | calc | c_ped_ESV_Cubed_MM | none | unilateral | 995300\109070\125206\G-0394\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  | DCM, 109070, End Systole |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | EDV (MM-Cubed) | none | calc | c_ped_EDV_Cubed_MM | none | unilateral | 995300\125206\F-32011\G-0394\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | ESV (MM-Teich) | none | calc | c_ped_ESV_Teich_MM | none | unilateral | 995300\109070\125209\G-0394\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  | DCM, 109070, End Systole |  | DCM, 125209, Teichholz
Pediatric Echocardiography Procedure Report | OB | EDV (MM-Teich) | none | calc | c_ped_EDV_Teich_MM | none | unilateral | 995300\125209\F-32011\G-0394\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole |  | DCM, 125209, Teichholz
Pediatric Echocardiography Procedure Report | OB | ESV (2D-Cubed) | none | calc | c_ped_ESV_Cubed | none | unilateral | 995300\109070\125206\G-03A2\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | DCM, 109070, End Systole |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | EDV (2D-Cubed) | none | calc | c_ped_EDV_Cubed | none | unilateral | 995300\125206\F-32011\G-03A2\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | ESV (BP) | none | calc | c_ped_ESV_BP | none | unilateral | 995300\109070\125207\G-03A2\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | DCM, 109070, End Systole |  | DCM, 125207, Method of Disks, Biplane
Pediatric Echocardiography Procedure Report | OB | EDV (BP) | none | calc | c_ped_EDV_BP | none | unilateral | 995300\125207\F-32011\G-03A2\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | DCM, 125207, Method of Disks, Biplane
Pediatric Echocardiography Procedure Report | OB | ESV (2D-Teich) | none | calc | c_ped_ESV_Teich | none | unilateral | 995300\109070\125209\G-03A2\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | DCM, 109070, End Systole |  | DCM, 125209, Teichholz
Pediatric Echocardiography Procedure Report | OB | EDV (2D-Teich) | none | calc | c_ped_EDV_Teich | none | unilateral | 995300\125209\F-32011\G-03A2\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | DCM, 125209, Teichholz
Pediatric Echocardiography Procedure Report | OB | ESV (A2C-A/L) | none | calc | c_ped_ESV_A2C_SPE | none | unilateral | 995300\109070\125226\G-03A2\G-A19B\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | DCM, 109070, End Systole |  | DCM, 125226, Single Plane Ellipse
Pediatric Echocardiography Procedure Report | OB | ESV (A4C-A/L) | none | calc | c_ped_ESV_A4C_SPE | none | unilateral | 995300\109070\125226\G-03A2\G-A19C\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | DCM, 109070, End Systole |  | DCM, 125226, Single Plane Ellipse
Pediatric Echocardiography Procedure Report | OB | EDV (A2C-A/L) | none | calc | c_ped_EDV_A2C_SPE | none | unilateral | 995300\125226\F-32011\G-03A2\G-A19B\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125226, Single Plane Ellipse
Pediatric Echocardiography Procedure Report | OB | EDV (A4C-A/L) | none | calc | c_ped_EDV_A4C_SPE | none | unilateral | 995300\125226\F-32011\G-03A2\G-A19C\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SNM3, G-D705, Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125226, Single Plane Ellipse
Pediatric Echocardiography Procedure Report | OB | SI (MM-Cubed) | none | calc | c_ped_SI_Cubed_MM | none | unilateral | 995300\125206\F-00078\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | SI (MM-Teich) | none | calc | c_ped_SI_Teich_MM | none | unilateral | 995300\125209\F-00078\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Echocardiography Procedure Report | OB | SI (2D-Cubed) | none | calc | c_ped_SI_Cubed | none | unilateral | 995300\125206\F-00078\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | SI (BP) | none | calc | c_ped_SI_BP | none | unilateral | 995300\125207\F-00078\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Echocardiography Procedure Report | OB | SI (2D-Teich) | none | calc | c_ped_SI_Teich | none | unilateral | 995300\125209\F-00078\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Echocardiography Procedure Report | OB | SI (A2C) | none | calc | c_ped_SI_AP2 | none | unilateral | 995300\125208\F-00078\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | SI (A2C-A/L) | none | calc | c_ped_SI_A2C_SPE | none | unilateral | 995300\125226\F-00078\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Echocardiography Procedure Report | OB | SI (A4C) | none | calc | c_ped_SI_AP4 | none | unilateral | 995300\125208\F-00078\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | SI (A4C-A/L) | none | calc | c_ped_SI_A4C_SPE | none | unilateral | 995300\125226\F-00078\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Echocardiography Procedure Report | OB | CO (MM-Cubed) | none | calc | c_ped_CO_Cubed_MM | none | unilateral | 995300\125206\F-32100\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | CO (MM-Teich) | none | calc | c_ped_CO_Teich_MM | none | unilateral | 995300\125209\F-32100\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Echocardiography Procedure Report | OB | CO (2D-Cubed) | none | calc | c_ped_CO_Cubed | none | unilateral | 995300\125206\F-32100\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | CO (BP) | none | calc | c_ped_CO_BP | none | unilateral | 995300\125207\F-32100\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Echocardiography Procedure Report | OB | CO (2D-Teich) | none | calc | c_ped_CO_Teich | none | unilateral | 995300\125209\F-32100\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Echocardiography Procedure Report | OB | CO (A2C) | none | calc | c_ped_CO_AP2 | none | unilateral | 995300\125208\F-32100\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | CO (A2C-A/L) | none | calc | c_ped_CO_A2C_SPE | none | unilateral | 995300\125226\F-32100\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Echocardiography Procedure Report | OB | CO (A4C) | none | calc | c_ped_CO_AP4 | none | unilateral | 995300\125208\F-32100\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | CO (A4C-A/L) | none | calc | c_ped_CO_A4C_SPE | none | unilateral | 995300\125226\F-32100\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Echocardiography Procedure Report | OB | CO (LVOT) | none | calc | c_ped_CO_LVOT | none | unilateral | 995300\F-32100\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s | SNM3, T-32650, Left Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | CI (MM-Cubed) | none | calc | c_ped_CI_Cubed_MM | none | unilateral | 995300\125206\F-32110\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | CI (MM-Teich) | none | calc | c_ped_CI_Teich_MM | none | unilateral | 995300\125209\F-32110\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Echocardiography Procedure Report | OB | CI (2D-Cubed) | none | calc | c_ped_CI_Cubed | none | unilateral | 995300\125206\F-32110\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | CI (BP) | none | calc | c_ped_CI_BP | none | unilateral | 995300\125207\F-32110\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Echocardiography Procedure Report | OB | CI (2D-Teich) | none | calc | c_ped_CI_Teich | none | unilateral | 995300\125209\F-32110\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Echocardiography Procedure Report | OB | CI (A2C) | none | calc | c_ped_CI_AP2 | none | unilateral | 995300\125208\F-32110\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | CI (A2C-A/L) | none | calc | c_ped_CI_A2C_SPE | none | unilateral | 995300\125226\F-32110\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Echocardiography Procedure Report | OB | CI (A4C) | none | calc | c_ped_CI_AP4 | none | unilateral | 995300\125208\F-32110\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | CI (A4C-A/L) | none | calc | c_ped_CI_A4C_SPE | none | unilateral | 995300\125226\F-32110\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Echocardiography Procedure Report | OB | SV (MM-Cubed) | none | calc | c_ped_SV_Cubed_MM | none | unilateral | 995300\125206\F-32120\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | SV (MM-Teich) | none | calc | c_ped_SV_Teich_MM | none | unilateral | 995300\125209\F-32120\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Echocardiography Procedure Report | OB | SV (2D-Cubed) | none | calc | c_ped_SV_Cubed | none | unilateral | 995300\125206\F-32120\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Echocardiography Procedure Report | OB | SV (BP) | none | calc | c_ped_SV_BP | none | unilateral | 995300\125207\F-32120\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Echocardiography Procedure Report | OB | SV (2D-Teich) | none | calc | c_ped_SV_Teich | none | unilateral | 995300\125209\F-32120\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Echocardiography Procedure Report | OB | SV (A2C) | none | calc | c_ped_SV_AP2 | none | unilateral | 995300\125208\F-32120\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | SV (A2C-A/L) | none | calc | c_ped_SV_A2C_SPE | none | unilateral | 995300\125226\F-32120\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Echocardiography Procedure Report | OB | SV (A4C) | none | calc | c_ped_SV_AP4 | none | unilateral | 995300\125208\F-32120\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | SV (A4C-A/L) | none | calc | c_ped_SV_A4C_SPE | none | unilateral | 995300\125226\F-32120\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Echocardiography Procedure Report | OB | SV (LVOT) | none | calc | c_ped_SV_LVOT | none | unilateral | 995300\F-32120\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3 | SNM3, T-32650, Left Ventricle Outflow Tract
Pediatric Echocardiography Procedure Report | OB | E/E` Medial | none | calc | c_p_E_MedE | none | unilateral | 995300\G-037B\G-0391\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037B, Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave |  |  | UCUM, 1, no units | SRT, G-0391, Medial Mitral Annulus
Pediatric Echocardiography Procedure Report | OB | E/E` Lateral | none | calc | c_p_E_LatE | none | unilateral | 995300\G-037B\G-0392\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037B, Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave |  |  | UCUM, 1, no units | SRT, G-0392, Lateral Mitral Annulus
Pediatric Echocardiography Procedure Report | OB | LV MPI | none | calc | c_ped_LV_MPI | none | unilateral | 995300\G-037F\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037F, Left Ventricular Index of Myocardial Performance |  |  | UCUM, 1, no units
Pediatric Echocardiography Procedure Report | OB | CO (PV) | none | calc | c_ped_CO_PV | none | unilateral | 995300\F-32100\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, F-32100, Cardiac Output |  |  | UCUM, mm3/s, mm3/s
Pediatric Echocardiography Procedure Report | OB | SV (PV) | none | calc | c_ped_SV_PV | none | unilateral | 995300\F-32120\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, F-32120, Stroke Volume |  |  | UCUM, mm3, mm3
Pediatric Echocardiography Procedure Report | OB | RVSP (VSD) | none | calc | c_p_RVSP_VSD | none | unilateral | 995300\109070\C12220-09\D4-31150 | SRT, D4-31150, Ventricular Septal Defect | 99PMSBLUS, C12220-09, Pressure |  |  | UCUM, mmHg, mmHg |  |  |  | DCM, 109070, End Systole
Pediatric Echocardiography Procedure Report | OB | Qp/Qs | none | calc | c_ped_QpQs | none | unilateral | 995300\29462-9\P5-30031 | SRT, P5-30031, Cardiac Shunt Study | LN, 29462-9, Pulmonary-to-Systemic Shunt Flow Ratio |  |  | UCUM, 1, no units
Pediatric Echocardiography Procedure Report | OB | VSD Diam | none | calc | c_p_VSD_Diam | none | unilateral | 995300\121428\D4-31150\G-03A2\M-02550 | SRT, D4-31150, Ventricular Septal Defect | SNM3, M-02550, Diameter |  |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode | [7 empty] | DCM, 121428, Calculated
Pediatric Echocardiography Procedure Report | OB | LA ESV Index (BP) | none | calc | c_p_BP_LA_ESV_BSA | none | unilateral | 995300\125207\C12205-03\G-03A2\T-32300 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-03, Left Atrium Systolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Echocardiography Procedure Report | OB | LA ESV Index (A2C) | none | calc | c_p_A2C_LA_ESV_BSA | none | unilateral | 995300\125208\C12205-03\G-03A2\G-A19B\T-32300 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-03, Left Atrium Systolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | LA ESV Index (A4C) | none | calc | c_p_A4C_LA_ESV_BSA | none | unilateral | 995300\125208\C12205-03\G-03A2\G-A19C\T-32300 | SNM3, T-32300, Left Atrium | 99PMSBLUS, C12205-03, Left Atrium Systolic Volume Index |  |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Echocardiography Procedure Report | OB | LA ESV (BP) | none | calc | c_p_ESV_LA_BP | none | unilateral | 995300\125207\G-0383\G-03A2\T-32300 | SNM3, T-32300, Left Atrium | SRT, G-0383, Left Atrium Systolic Volume |  |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Small Parts Ultrasound Procedure Report | Small Parts | EPI Body H | Dist | meas | m_epibodyht | r_Distance | bilateral | 995400\121207\C4031-01\G-03A2 | 99PMSBLUS, C4031-01, Epididymis Body |  | DCM, 121207, Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | EPI Body W | Dist | meas | m_epibodywidth | r_Distance | bilateral | 995400\C4031-01\G-03A2\G-A220 | 99PMSBLUS, C4031-01, Epididymis Body |  | SRT, G-A220, Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | EPI Body L | Dist | meas | m_epibodyLen | r_Distance | bilateral | 995400\C4031-01\G-03A2\G-A22A | 99PMSBLUS, C4031-01, Epididymis Body |  | SRT, G-A22A, Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | EPI Head H | Dist | meas | m_epiheadht | r_Distance | bilateral | 995400\121207\C4031-02\G-03A2 | 99PMSBLUS, C4031-02, Epididymis Head |  | DCM, 121207, Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | EPI Head W | Dist | meas | m_epiheadwidth | r_Distance | bilateral | 995400\C4031-02\G-03A2\G-A220 | 99PMSBLUS, C4031-02, Epididymis Head |  | SRT, G-A220, Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | EPI Head L | Dist | meas | m_epiheadLen | r_Distance | bilateral | 995400\C4031-02\G-03A2\G-A22A | 99PMSBLUS, C4031-02, Epididymis Head |  | SRT, G-A22A, Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | EPI Tail H | Dist | meas | m_epitailht | r_Distance | bilateral | 995400\121207\C4031-03\G-03A2 | 99PMSBLUS, C4031-03, Epididymis Tail |  | DCM, 121207, Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | EPI Tail W | Dist | meas | m_epitailwidth | r_Distance | bilateral | 995400\C4031-03\G-03A2\G-A220 | 99PMSBLUS, C4031-03, Epididymis Tail |  | SRT, G-A220, Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | EPI Tail L | Dist | meas | m_epitailLen | r_Distance | bilateral | 995400\C4031-03\G-03A2\G-A22A | 99PMSBLUS, C4031-03, Epididymis Tail |  | SRT, G-A22A, Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | EQI Avg | meas | m_brles_EQI_4 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-004\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-004, ShearWave Elastography Stiffness sample 4 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | EQI Avg | meas | m_brles_EQI_5 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-005\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-005, ShearWave Elastography Stiffness sample 5 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | EQI Avg | meas | m_brles_EQI_6 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-006\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-006, ShearWave Elastography Stiffness sample 6 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | EQI Avg | meas | m_brles_EQI_7 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-007\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-007, ShearWave Elastography Stiffness sample 7 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | EQI Avg | meas | m_brles_EQI_8 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-008\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-008, ShearWave Elastography Stiffness sample 8 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | EQI Avg | meas | m_brles_EQI_9 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-009\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-009, ShearWave Elastography Stiffness sample 9 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | EQI Avg | meas | m_brles_EQI_10 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-010\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-010, ShearWave Elastography Stiffness sample 10 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | EQI Avg Vel | meas | m_brles_EQI_1 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-011\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-011, ShearWave Elastography Velocity sample 1 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | EQI Avg Vel | meas | m_brles_EQI_2 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-012\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-012, ShearWave Elastography Velocity sample 2 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | EQI Avg Vel | meas | m_brles_EQI_3 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-013\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-013, ShearWave Elastography Velocity sample 3 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | EQI Avg Vel | meas | m_brles_EQI_4 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-014\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-014, ShearWave Elastography Velocity sample 4 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | EQI Avg Vel | meas | m_brles_EQI_5 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-015\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-015, ShearWave Elastography Velocity sample 5 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | EQI Avg Vel | meas | m_brles_EQI_6 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-016\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-016, ShearWave Elastography Velocity sample 6 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | EQI Avg Vel | meas | m_brles_EQI_7 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-017\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-017, ShearWave Elastography Velocity sample 7 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | EQI Avg Vel | meas | m_brles_EQI_8 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-018 \C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-018 , ShearWave Elastography Velocity sample 8 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | EQI Avg Vel | meas | m_brles_EQI_9 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-019\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-019, ShearWave Elastography Velocity sample 9 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | EQI Avg Vel | meas | m_brles_EQI_10 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-020\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-020, ShearWave Elastography Velocity sample 10 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | EQI Std | meas | m_brles_EQI_4 | r_StdStiffness_p | bilateral | 995400\ C99-5100-047\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-047, ShearWave Elastography Stiffness sample 4 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | EQI Std | meas | m_brles_EQI_5 | r_StdStiffness_p | bilateral | 995400\ C99-5100-048\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-048, ShearWave Elastography Stiffness sample 5 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | EQI Std | meas | m_brles_EQI_6 | r_StdStiffness_p | bilateral | 995400\ C99-5100-049\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-049, ShearWave Elastography Stiffness sample 6 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | EQI Std | meas | m_brles_EQI_7 | r_StdStiffness_p | bilateral | 995400\ C99-5100-050\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-050, ShearWave Elastography Stiffness sample 7 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | EQI Std | meas | m_brles_EQI_8 | r_StdStiffness_p | bilateral | 995400\ C99-5100-051\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-051, ShearWave Elastography Stiffness sample 8 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | EQI Std | meas | m_brles_EQI_9 | r_StdStiffness_p | bilateral | 995400\ C99-5100-052\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-052, ShearWave Elastography Stiffness sample 9 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | EQI Std | meas | m_brles_EQI_10 | r_StdStiffness_p | bilateral | 995400\ C99-5100-053\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-053, ShearWave Elastography Stiffness sample 10 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | EQI Std Vel | meas | m_brles_EQI_1 | r_StdStiffness_v | bilateral | 995400\ C99-5100-054\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-054, ShearWave Elastography Velocity sample 1 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | EQI Std Vel | meas | m_brles_EQI_2 | r_StdStiffness_v | bilateral | 995400\ C99-5100-055\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-055, ShearWave Elastography Velocity sample 2 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | EQI Std Vel | meas | m_brles_EQI_3 | r_StdStiffness_v | bilateral | 995400\ C99-5100-056\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-056, ShearWave Elastography Velocity sample 3 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | EQI Std Vel | meas | m_brles_EQI_4 | r_StdStiffness_v | bilateral | 995400\ C99-5100-057\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-057, ShearWave Elastography Velocity sample 4 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | EQI Std Vel | meas | m_brles_EQI_5 | r_StdStiffness_v | bilateral | 995400\ C99-5100-058\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-058, ShearWave Elastography Velocity sample 5 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | EQI Std Vel | meas | m_brles_EQI_6 | r_StdStiffness_v | bilateral | 995400\ C99-5100-059\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-059, ShearWave Elastography Velocity sample 6 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | EQI Std Vel | meas | m_brles_EQI_7 | r_StdStiffness_v | bilateral | 995400\ C99-5100-060\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-060, ShearWave Elastography Velocity sample 7 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | EQI Std Vel | meas | m_brles_EQI_8 | r_StdStiffness_v | bilateral | 995400\ C99-5100-061\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-061, ShearWave Elastography Velocity sample 8 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | EQI Std Vel | meas | m_brles_EQI_9 | r_StdStiffness_v | bilateral | 995400\ C99-5100-062\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-062, ShearWave Elastography Velocity sample 9 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | EQI Std Vel | meas | m_brles_EQI_10 | r_StdStiffness_v | bilateral | 995400\ C99-5100-063\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-063, ShearWave Elastography Velocity sample 10 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | EQI Med | meas | m_brles_EQI_4 | r_MedStiffness_p | bilateral | 995400\ C99-5100-067\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-067, ShearWave Elastography Stiffness sample 4 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | EQI Med | meas | m_brles_EQI_5 | r_MedStiffness_p | bilateral | 995400\ C99-5100-068\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-068, ShearWave Elastography Stiffness sample 5 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | EQI Med | meas | m_brles_EQI_6 | r_MedStiffness_p | bilateral | 995400\ C99-5100-069\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-069, ShearWave Elastography Stiffness sample 6 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | EQI Med | meas | m_brles_EQI_7 | r_MedStiffness_p | bilateral | 995400\ C99-5100-070\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-070, ShearWave Elastography Stiffness sample 7 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | EQI Med | meas | m_brles_EQI_8 | r_MedStiffness_p | bilateral | 995400\ C99-5100-071\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-071, ShearWave Elastography Stiffness sample 8 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | EQI Med | meas | m_brles_EQI_9 | r_MedStiffness_p | bilateral | 995400\ C99-5100-072\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-072, ShearWave Elastography Stiffness sample 9 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | EQI Med | meas | m_brles_EQI_10 | r_MedStiffness_p | bilateral | 995400\ C99-5100-073\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-073, ShearWave Elastography Stiffness sample 10 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | EQI Med Vel | meas | m_brles_EQI_1 | r_MedStiffness_v | bilateral | 995400\ C99-5100-074\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-074, ShearWave Elastography Velocity sample 1 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | EQI Med Vel | meas | m_brles_EQI_2 | r_MedStiffness_v | bilateral | 995400\ C99-5100-075\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-075, ShearWave Elastography Velocity sample 2 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | EQI Med Vel | meas | m_brles_EQI_3 | r_MedStiffness_v | bilateral | 995400\ C99-5100-076\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-076, ShearWave Elastography Velocity sample 3 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | EQI Med Vel | meas | m_brles_EQI_4 | r_MedStiffness_v | bilateral | 995400\ C99-5100-077\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-077, ShearWave Elastography Velocity sample 4 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | EQI Med Vel | meas | m_brles_EQI_5 | r_MedStiffness_v | bilateral | 995400\ C99-5100-078\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-078, ShearWave Elastography Velocity sample 5 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | EQI Med Vel | meas | m_brles_EQI_6 | r_MedStiffness_v | bilateral | 995400\ C99-5100-079\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-079, ShearWave Elastography Velocity sample 6 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | EQI Med Vel | meas | m_brles_EQI_7 | r_MedStiffness_v | bilateral | 995400\ C99-5100-080\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-080, ShearWave Elastography Velocity sample 7 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | EQI Med Vel | meas | m_brles_EQI_8 | r_MedStiffness_v | bilateral | 995400\ C99-5100-081\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-081, ShearWave Elastography Velocity sample 8 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | EQI Med Vel | meas | m_brles_EQI_9 | r_MedStiffness_v | bilateral | 995400\ C99-5100-082\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-082, ShearWave Elastography Velocity sample 9 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | EQI Med Vel | meas | m_brles_EQI_10 | r_MedStiffness_v | bilateral | 995400\ C99-5100-083\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-083, ShearWave Elastography Velocity sample 10 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | EQI IQR | meas | m_brles_EQI_4 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-087\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-087, ShearWave Elastography Stiffness sample 4 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | EQI IQR | meas | m_brles_EQI_5 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-088\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-088, ShearWave Elastography Stiffness sample 5 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | EQI IQR | meas | m_brles_EQI_6 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-089\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-089, ShearWave Elastography Stiffness sample 6 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | EQI IQR | meas | m_brles_EQI_7 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-090\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-090, ShearWave Elastography Stiffness sample 7 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | EQI IQR | meas | m_brles_EQI_8 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-091\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-091, ShearWave Elastography Stiffness sample 8 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | EQI IQR | meas | m_brles_EQI_9 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-092\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-092, ShearWave Elastography Stiffness sample 9 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | EQI IQR | meas | m_brles_EQI_10 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-093\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-093, ShearWave Elastography Stiffness sample 10 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | EQI IQR Vel | meas | m_brles_EQI_1 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-094\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-094, ShearWave Elastography Velocity sample 1 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | EQI IQR Vel | meas | m_brles_EQI_2 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-095\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-095, ShearWave Elastography Velocity sample 2 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | EQI IQR Vel | meas | m_brles_EQI_3 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-096\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-096, ShearWave Elastography Velocity sample 3 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | EQI IQR Vel | meas | m_brles_EQI_4 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-097\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-097, ShearWave Elastography Velocity sample 4 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | EQI IQR Vel | meas | m_brles_EQI_5 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-098\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-098, ShearWave Elastography Velocity sample 5 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | EQI IQR Vel | meas | m_brles_EQI_6 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-099\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-099, ShearWave Elastography Velocity sample 6 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | EQI IQR Vel | meas | m_brles_EQI_7 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-100\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-100, ShearWave Elastography Velocity sample 7 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | EQI IQR Vel | meas | m_brles_EQI_8 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-101\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-101, ShearWave Elastography Velocity sample 8 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | EQI IQR Vel | meas | m_brles_EQI_9 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-102\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-102, ShearWave Elastography Velocity sample 9 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | EQI IQR Vel | meas | m_brles_EQI_10 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-103\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-103, ShearWave Elastography Velocity sample 10 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | EQI IQR/Med | meas | m_brles_EQI_4 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-107\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-107, ShearWave Elasto Stiffness sample 4 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | EQI IQR/Med | meas | m_brles_EQI_5 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-108\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-108, ShearWave Elasto Stiffness sample 5 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | EQI IQR/Med | meas | m_brles_EQI_6 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-109\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-109, ShearWave Elasto Stiffness sample 6 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | EQI IQR/Med | meas | m_brles_EQI_7 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-110\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-110, ShearWave Elasto Stiffness sample 7 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | EQI IQR/Med | meas | m_brles_EQI_8 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-111\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-111, ShearWave Elasto Stiffness sample 8 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | EQI IQR/Med | meas | m_brles_EQI_9 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-112\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-112, ShearWave Elasto Stiffness sample 9 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | EQI IQR/Med | meas | m_brles_EQI_10 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-113\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-113, ShearWave Elasto Stiffness sample 10 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | EQI IQR/Med Vel | meas | m_brles_EQI_1 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-114\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-114, ShearWave Elasto Velocity sample 1 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | EQI IQR/Med Vel | meas | m_brles_EQI_2 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-115\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-115, ShearWave Elasto Velocity sample 2 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | EQI IQR/Med Vel | meas | m_brles_EQI_3 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-116\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-116, ShearWave Elasto Velocity sample 3 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | EQI IQR/Med Vel | meas | m_brles_EQI_4 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-117\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-117, ShearWave Elasto Velocity sample 4 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | EQI IQR/Med Vel | meas | m_brles_EQI_5 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-118\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-118, ShearWave Elasto Velocity sample 5 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | EQI IQR/Med Vel | meas | m_brles_EQI_6 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-119\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-119, ShearWave Elasto Velocity sample 6 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | EQI IQR/Med Vel | meas | m_brles_EQI_7 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-120\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-120, ShearWave Elasto Velocity sample 7 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | EQI IQR/Med Vel | meas | m_brles_EQI_8 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-121\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-121, ShearWave Elasto Velocity sample 8 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | EQI IQR/Med Vel | meas | m_brles_EQI_9 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-122\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-122, ShearWave Elasto Velocity sample 9 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | EQI IQR/Med Vel | meas | m_brles_EQI_10 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-123\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-123, ShearWave Elasto Velocity sample 10 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | Conf. Threshold | meas | m_brles_EQI_4 | r_CFStiffness | bilateral | 995400\ C99-5100-127\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-127, ShearWave Elastography Stiffness sample 4 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | Conf. Threshold | meas | m_brles_EQI_5 | r_CFStiffness | bilateral | 995400\ C99-5100-128\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-128, ShearWave Elastography Stiffness sample 5 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | Conf. Threshold | meas | m_brles_EQI_6 | r_CFStiffness | bilateral | 995400\ C99-5100-129\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-129, ShearWave Elastography Stiffness sample 6 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | Conf. Threshold | meas | m_brles_EQI_7 | r_CFStiffness | bilateral | 995400\ C99-5100-130\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-130, ShearWave Elastography Stiffness sample 7 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | Conf. Threshold | meas | m_brles_EQI_8 | r_CFStiffness | bilateral | 995400\ C99-5100-131\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-131, ShearWave Elastography Stiffness sample 8 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | Conf. Threshold | meas | m_brles_EQI_9 | r_CFStiffness | bilateral | 995400\ C99-5100-132\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-132, ShearWave Elastography Stiffness sample 9 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | Conf. Threshold | meas | m_brles_EQI_10 | r_CFStiffness | bilateral | 995400\ C99-5100-133\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-133, ShearWave Elastography Stiffness sample 10 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | Circ | meas | m_brles_EQI_4 | r_Circumference | bilateral | 995400\ C99-5100-137\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-137, ShearWave Elastography sample 4 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | Circ | meas | m_brles_EQI_5 | r_Circumference | bilateral | 995400\ C99-5100-138\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-138, ShearWave Elastography sample 5 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | Circ | meas | m_brles_EQI_6 | r_Circumference | bilateral | 995400\ C99-5100-139\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-139, ShearWave Elastography sample 6 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | Circ | meas | m_brles_EQI_7 | r_Circumference | bilateral | 995400\ C99-5100-140\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-140, ShearWave Elastography sample 7 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | Circ | meas | m_brles_EQI_8 | r_Circumference | bilateral | 995400\ C99-5100-141\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-141, ShearWave Elastography sample 8 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | Circ | meas | m_brles_EQI_9 | r_Circumference | bilateral | 995400\ C99-5100-142\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-142, ShearWave Elastography sample 9 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | Circ | meas | m_brles_EQI_10 | r_Circumference | bilateral | 995400\ C99-5100-143\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-143, ShearWave Elastography sample 10 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | Area | meas | m_brles_EQI_4 | r_Area | bilateral | 995400\ C99-5100-147\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-147, ShearWave Elastography sample 4 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | Area | meas | m_brles_EQI_5 | r_Area | bilateral | 995400\ C99-5100-148\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-148, ShearWave Elastography sample 5 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | Area | meas | m_brles_EQI_6 | r_Area | bilateral | 995400\ C99-5100-149\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-149, ShearWave Elastography sample 6 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | Area | meas | m_brles_EQI_7 | r_Area | bilateral | 995400\ C99-5100-150\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-150, ShearWave Elastography sample 7 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | Area | meas | m_brles_EQI_8 | r_Area | bilateral | 995400\ C99-5100-151\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-151, ShearWave Elastography sample 8 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | Area | meas | m_brles_EQI_9 | r_Area | bilateral | 995400\ C99-5100-152\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-152, ShearWave Elastography sample 9 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | Area | meas | m_brles_EQI_10 | r_Area | bilateral | 995400\ C99-5100-153\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-153, ShearWave Elastography sample 10 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | EQI Avg | meas | m_brles_EQI_1 | r_AvgStiffness_p | bilateral | 995400\C4031-04\C99-5100-001\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-001, ShearWave Elastography Stiffness sample 1 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | EQI Avg | meas | m_brles_EQI_2 | r_AvgStiffness_p | bilateral | 995400\C4031-04\C99-5100-002\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-002, ShearWave Elastography Stiffness sample 2 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | EQI Avg | meas | m_brles_EQI_3 | r_AvgStiffness_p | bilateral | 995400\C4031-04\C99-5100-003\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-003, ShearWave Elastography Stiffness sample 3 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | EQI Std | meas | m_brles_EQI_1 | r_StdStiffness_p | bilateral | 995400\C4031-04\C99-5100-044\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-044, ShearWave Elastography Stiffness sample 1 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | EQI Std | meas | m_brles_EQI_2 | r_StdStiffness_p | bilateral | 995400\C4031-04\C99-5100-045\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-045, ShearWave Elastography Stiffness sample 2 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | EQI Std | meas | m_brles_EQI_3 | r_StdStiffness_p | bilateral | 995400\C4031-04\C99-5100-046\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-046, ShearWave Elastography Stiffness sample 3 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | EQI Med | meas | m_brles_EQI_1 | r_MedStiffness_p | bilateral | 995400\C4031-04\C99-5100-064\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-064, ShearWave Elastography Stiffness sample 1 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | EQI Med | meas | m_brles_EQI_2 | r_MedStiffness_p | bilateral | 995400\C4031-04\C99-5100-065\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-065, ShearWave Elastography Stiffness sample 2 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | EQI Med | meas | m_brles_EQI_3 | r_MedStiffness_p | bilateral | 995400\C4031-04\C99-5100-066\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-066, ShearWave Elastography Stiffness sample 3 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | EQI IQR | meas | m_brles_EQI_1 | r_IqrStiffness_p | bilateral | 995400\C4031-04\C99-5100-084\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-084, ShearWave Elastography Stiffness sample 1 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | EQI IQR | meas | m_brles_EQI_2 | r_IqrStiffness_p | bilateral | 995400\C4031-04\C99-5100-085\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-085, ShearWave Elastography Stiffness sample 2 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | EQI IQR | meas | m_brles_EQI_3 | r_IqrStiffness_p | bilateral | 995400\C4031-04\C99-5100-086\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-086, ShearWave Elastography Stiffness sample 3 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | EQI IQR/Med | meas | m_brles_EQI_1 | r_IqrMedStiffness_p | bilateral | 995400\C4031-04\C99-5100-104\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-104, ShearWave Elasto Stiffness sample 1 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | EQI IQR/Med | meas | m_brles_EQI_2 | r_IqrMedStiffness_p | bilateral | 995400\C4031-04\C99-5100-105\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-105, ShearWave Elasto Stiffness sample 2 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | EQI IQR/Med | meas | m_brles_EQI_3 | r_IqrMedStiffness_p | bilateral | 995400\C4031-04\C99-5100-106\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-106, ShearWave Elasto Stiffness sample 3 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | Conf. Threshold | meas | m_brles_EQI_1 | r_CFStiffness | bilateral | 995400\C4031-04\C99-5100-124\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-124, ShearWave Elastography Stiffness sample 1 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | Conf. Threshold | meas | m_brles_EQI_2 | r_CFStiffness | bilateral | 995400\C4031-04\C99-5100-125\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-125, ShearWave Elastography Stiffness sample 2 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | Conf. Threshold | meas | m_brles_EQI_3 | r_CFStiffness | bilateral | 995400\C4031-04\C99-5100-126\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-126, ShearWave Elastography Stiffness sample 3 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | Circ | meas | m_brles_EQI_1 | r_Circumference | bilateral | 995400\C4031-04\C99-5100-134\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-134, ShearWave Elastography sample 1 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | Circ | meas | m_brles_EQI_2 | r_Circumference | bilateral | 995400\C4031-04\C99-5100-135\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-135, ShearWave Elastography sample 2 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | Circ | meas | m_brles_EQI_3 | r_Circumference | bilateral | 995400\C4031-04\C99-5100-136\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-136, ShearWave Elastography sample 3 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | Area | meas | m_brles_EQI_1 | r_Area | bilateral | 995400\C4031-04\C99-5100-144\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-144, ShearWave Elastography sample 1 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | Area | meas | m_brles_EQI_2 | r_Area | bilateral | 995400\C4031-04\C99-5100-145\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-145, ShearWave Elastography sample 2 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | Area | meas | m_brles_EQI_3 | r_Area | bilateral | 995400\C4031-04\C99-5100-146\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-146, ShearWave Elastography sample 3 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | EQI Min | meas | m_brles_EQI_1 | r_MinStiffness_p | bilateral | 995400\C4031-04\C99-5100-154\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-154, ShearWave Elastography Stiffness sample 1 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | EQI Min | meas | m_brles_EQI_2 | r_MinStiffness_p | bilateral | 995400\C4031-04\C99-5100-155\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-155, ShearWave Elastography Stiffness sample 2 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | EQI Min | meas | m_brles_EQI_3 | r_MinStiffness_p | bilateral | 995400\C4031-04\C99-5100-156\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-156, ShearWave Elastography Stiffness sample 3 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | EQI Min | meas | m_brles_EQI_4 | r_MinStiffness_p | bilateral | 995400\C4031-04\C99-5100-157\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-157, ShearWave Elastography Stiffness sample 4 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | EQI Min | meas | m_brles_EQI_5 | r_MinStiffness_p | bilateral | 995400\C4031-04\C99-5100-158\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-158, ShearWave Elastography Stiffness sample 5 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | EQI Min | meas | m_brles_EQI_6 | r_MinStiffness_p | bilateral | 995400\C4031-04\C99-5100-159\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-159, ShearWave Elastography Stiffness sample 6 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | EQI Min | meas | m_brles_EQI_7 | r_MinStiffness_p | bilateral | 995400\C4031-04\C99-5100-160\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-160, ShearWave Elastography Stiffness sample 7 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | EQI Min | meas | m_brles_EQI_8 | r_MinStiffness_p | bilateral | 995400\C4031-04\C99-5100-161\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-161, ShearWave Elastography Stiffness sample 8 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | EQI Min | meas | m_brles_EQI_9 | r_MinStiffness_p | bilateral | 995400\C4031-04\C99-5100-162\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-162, ShearWave Elastography Stiffness sample 9 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | EQI Min | meas | m_brles_EQI_10 | r_MinStiffness_p | bilateral | 995400\C4031-04\C99-5100-163\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-163, ShearWave Elastography Stiffness sample 10 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | EQI Min Vel | meas | m_brles_EQI_1 | r_MinStiffness_v | bilateral | 995400\C4031-04\C99-5100-164\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-164, ShearWave Elastography Velocity sample 1 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | EQI Min Vel | meas | m_brles_EQI_2 | r_MinStiffness_v | bilateral | 995400\C4031-04\C99-5100-165\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-165, ShearWave Elastography Velocity sample 2 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | EQI Min Vel | meas | m_brles_EQI_3 | r_MinStiffness_v | bilateral | 995400\C4031-04\C99-5100-166\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-166, ShearWave Elastography Velocity sample 3 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | EQI Min Vel | meas | m_brles_EQI_4 | r_MinStiffness_v | bilateral | 995400\C4031-04\C99-5100-167\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-167, ShearWave Elastography Velocity sample 4 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | EQI Min Vel | meas | m_brles_EQI_5 | r_MinStiffness_v | bilateral | 995400\C4031-04\C99-5100-168\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-168, ShearWave Elastography Velocity sample 5 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | EQI Min Vel | meas | m_brles_EQI_6 | r_MinStiffness_v | bilateral | 995400\C4031-04\C99-5100-169\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-169, ShearWave Elastography Velocity sample 6 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | EQI Min Vel | meas | m_brles_EQI_7 | r_MinStiffness_v | bilateral | 995400\C4031-04\C99-5100-170\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-170, ShearWave Elastography Velocity sample 7 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | EQI Min Vel | meas | m_brles_EQI_8 | r_MinStiffness_v | bilateral | 995400\C4031-04\C99-5100-171 \T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-171 , ShearWave Elastography Velocity sample 8 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | EQI Min Vel | meas | m_brles_EQI_9 | r_MinStiffness_v | bilateral | 995400\C4031-04\C99-5100-172\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-172, ShearWave Elastography Velocity sample 9 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | EQI Min Vel | meas | m_brles_EQI_10 | r_MinStiffness_v | bilateral | 995400\C4031-04\C99-5100-173\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-173, ShearWave Elastography Velocity sample 10 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | EQI Max | meas | m_brles_EQI_1 | r_MaxStiffness_p | bilateral | 995400\C4031-04\C99-5100-174\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-174, ShearWave Elastography Stiffness sample 1 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | EQI Max | meas | m_brles_EQI_2 | r_MaxStiffness_p | bilateral | 995400\C4031-04\C99-5100-175\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-175, ShearWave Elastography Stiffness sample 2 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | EQI Max | meas | m_brles_EQI_3 | r_MaxStiffness_p | bilateral | 995400\C4031-04\C99-5100-176\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-176, ShearWave Elastography Stiffness sample 3 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | EQI Max | meas | m_brles_EQI_4 | r_MaxStiffness_p | bilateral | 995400\C4031-04\C99-5100-177\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-177, ShearWave Elastography Stiffness sample 4 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | EQI Max | meas | m_brles_EQI_5 | r_MaxStiffness_p | bilateral | 995400\C4031-04\C99-5100-178\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-178, ShearWave Elastography Stiffness sample 5 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | EQI Max | meas | m_brles_EQI_6 | r_MaxStiffness_p | bilateral | 995400\C4031-04\C99-5100-179\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-179, ShearWave Elastography Stiffness sample 6 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | EQI Max | meas | m_brles_EQI_7 | r_MaxStiffness_p | bilateral | 995400\C4031-04\C99-5100-180\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-180, ShearWave Elastography Stiffness sample 7 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | EQI Max | meas | m_brles_EQI_8 | r_MaxStiffness_p | bilateral | 995400\C4031-04\C99-5100-181\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-181, ShearWave Elastography Stiffness sample 8 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | EQI Max | meas | m_brles_EQI_9 | r_MaxStiffness_p | bilateral | 995400\C4031-04\C99-5100-182\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-182, ShearWave Elastography Stiffness sample 9 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | EQI Max | meas | m_brles_EQI_10 | r_MaxStiffness_p | bilateral | 995400\C4031-04\C99-5100-183\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-183, ShearWave Elastography Stiffness sample 10 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 1 | EQI Max Vel | meas | m_brles_EQI_1 | r_MaxStiffness_v | bilateral | 995400\C4031-04\C99-5100-184\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-184, ShearWave Elastography Velocity sample 1 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 2 | EQI Max Vel | meas | m_brles_EQI_2 | r_MaxStiffness_v | bilateral | 995400\C4031-04\C99-5100-185\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-185, ShearWave Elastography Velocity sample 2 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 3 | EQI Max Vel | meas | m_brles_EQI_3 | r_MaxStiffness_v | bilateral | 995400\C4031-04\C99-5100-186\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-186, ShearWave Elastography Velocity sample 3 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 4 | EQI Max Vel | meas | m_brles_EQI_4 | r_MaxStiffness_v | bilateral | 995400\C4031-04\C99-5100-187\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-187, ShearWave Elastography Velocity sample 4 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 5 | EQI Max Vel | meas | m_brles_EQI_5 | r_MaxStiffness_v | bilateral | 995400\C4031-04\C99-5100-188\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-188, ShearWave Elastography Velocity sample 5 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 6 | EQI Max Vel | meas | m_brles_EQI_6 | r_MaxStiffness_v | bilateral | 995400\C4031-04\C99-5100-189\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-189, ShearWave Elastography Velocity sample 6 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 7 | EQI Max Vel | meas | m_brles_EQI_7 | r_MaxStiffness_v | bilateral | 995400\C4031-04\C99-5100-190\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-190, ShearWave Elastography Velocity sample 7 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 8 | EQI Max Vel | meas | m_brles_EQI_8 | r_MaxStiffness_v | bilateral | 995400\C4031-04\C99-5100-191 \T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-191 , ShearWave Elastography Velocity sample 8 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 9 | EQI Max Vel | meas | m_brles_EQI_9 | r_MaxStiffness_v | bilateral | 995400\C4031-04\C99-5100-192\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-192, ShearWave Elastography Velocity sample 9 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI 10 | EQI Max Vel | meas | m_brles_EQI_10 | r_MaxStiffness_v | bilateral | 995400\C4031-04\C99-5100-193\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-193, ShearWave Elastography Velocity sample 10 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | EQI Avg | meas | m_thyles_EQI_4 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-004\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-004, ShearWave Elastography Stiffness sample 4 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | EQI Avg | meas | m_thyles_EQI_5 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-005\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-005, ShearWave Elastography Stiffness sample 5 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | EQI Avg | meas | m_thyles_EQI_6 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-006\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-006, ShearWave Elastography Stiffness sample 6 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | EQI Avg | meas | m_thyles_EQI_7 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-007\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-007, ShearWave Elastography Stiffness sample 7 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | EQI Avg | meas | m_thyles_EQI_8 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-008\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-008, ShearWave Elastography Stiffness sample 8 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | EQI Avg | meas | m_thyles_EQI_9 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-009\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-009, ShearWave Elastography Stiffness sample 9 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | EQI Avg | meas | m_thyles_EQI_10 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-010\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-010, ShearWave Elastography Stiffness sample 10 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | EQI Avg Vel | meas | m_thyles_EQI_1 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-011\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-011, ShearWave Elastography Velocity sample 1 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | EQI Avg Vel | meas | m_thyles_EQI_2 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-012\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-012, ShearWave Elastography Velocity sample 2 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | EQI Avg Vel | meas | m_thyles_EQI_3 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-013\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-013, ShearWave Elastography Velocity sample 3 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | EQI Avg Vel | meas | m_thyles_EQI_4 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-014\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-014, ShearWave Elastography Velocity sample 4 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | EQI Avg Vel | meas | m_thyles_EQI_5 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-015\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-015, ShearWave Elastography Velocity sample 5 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | EQI Avg Vel | meas | m_thyles_EQI_6 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-016\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-016, ShearWave Elastography Velocity sample 6 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | EQI Avg Vel | meas | m_thyles_EQI_7 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-017\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-017, ShearWave Elastography Velocity sample 7 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | EQI Avg Vel | meas | m_thyles_EQI_8 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-018 \C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-018 , ShearWave Elastography Velocity sample 8 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | EQI Avg Vel | meas | m_thyles_EQI_9 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-019\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-019, ShearWave Elastography Velocity sample 9 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | EQI Avg Vel | meas | m_thyles_EQI_10 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-020\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-020, ShearWave Elastography Velocity sample 10 Average |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | EQI Std | meas | m_thyles_EQI_4 | r_StdStiffness_p | bilateral | 995400\ C99-5100-047\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-047, ShearWave Elastography Stiffness sample 4 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | EQI Std | meas | m_thyles_EQI_5 | r_StdStiffness_p | bilateral | 995400\ C99-5100-048\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-048, ShearWave Elastography Stiffness sample 5 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | EQI Std | meas | m_thyles_EQI_6 | r_StdStiffness_p | bilateral | 995400\ C99-5100-049\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-049, ShearWave Elastography Stiffness sample 6 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | EQI Std | meas | m_thyles_EQI_7 | r_StdStiffness_p | bilateral | 995400\ C99-5100-050\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-050, ShearWave Elastography Stiffness sample 7 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | EQI Std | meas | m_thyles_EQI_8 | r_StdStiffness_p | bilateral | 995400\ C99-5100-051\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-051, ShearWave Elastography Stiffness sample 8 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | EQI Std | meas | m_thyles_EQI_9 | r_StdStiffness_p | bilateral | 995400\ C99-5100-052\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-052, ShearWave Elastography Stiffness sample 9 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | EQI Std | meas | m_thyles_EQI_10 | r_StdStiffness_p | bilateral | 995400\ C99-5100-053\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-053, ShearWave Elastography Stiffness sample 10 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | EQI Std Vel | meas | m_thyles_EQI_1 | r_StdStiffness_v | bilateral | 995400\ C99-5100-054\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-054, ShearWave Elastography Velocity sample 1 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | EQI Std Vel | meas | m_thyles_EQI_2 | r_StdStiffness_v | bilateral | 995400\ C99-5100-055\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-055, ShearWave Elastography Velocity sample 2 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | EQI Std Vel | meas | m_thyles_EQI_3 | r_StdStiffness_v | bilateral | 995400\ C99-5100-056\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-056, ShearWave Elastography Velocity sample 3 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | EQI Std Vel | meas | m_thyles_EQI_4 | r_StdStiffness_v | bilateral | 995400\ C99-5100-057\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-057, ShearWave Elastography Velocity sample 4 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | EQI Std Vel | meas | m_thyles_EQI_5 | r_StdStiffness_v | bilateral | 995400\ C99-5100-058\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-058, ShearWave Elastography Velocity sample 5 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | EQI Std Vel | meas | m_thyles_EQI_6 | r_StdStiffness_v | bilateral | 995400\ C99-5100-059\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-059, ShearWave Elastography Velocity sample 6 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | EQI Std Vel | meas | m_thyles_EQI_7 | r_StdStiffness_v | bilateral | 995400\ C99-5100-060\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-060, ShearWave Elastography Velocity sample 7 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | EQI Std Vel | meas | m_thyles_EQI_8 | r_StdStiffness_v | bilateral | 995400\ C99-5100-061\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-061, ShearWave Elastography Velocity sample 8 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | EQI Std Vel | meas | m_thyles_EQI_9 | r_StdStiffness_v | bilateral | 995400\ C99-5100-062\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-062, ShearWave Elastography Velocity sample 9 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | EQI Std Vel | meas | m_thyles_EQI_10 | r_StdStiffness_v | bilateral | 995400\ C99-5100-063\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-063, ShearWave Elastography Velocity sample 10 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | EQI Med | meas | m_thyles_EQI_4 | r_MedStiffness_p | bilateral | 995400\ C99-5100-067\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-067, ShearWave Elastography Stiffness sample 4 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | EQI Med | meas | m_thyles_EQI_5 | r_MedStiffness_p | bilateral | 995400\ C99-5100-068\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-068, ShearWave Elastography Stiffness sample 5 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | EQI Med | meas | m_thyles_EQI_6 | r_MedStiffness_p | bilateral | 995400\ C99-5100-069\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-069, ShearWave Elastography Stiffness sample 6 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | EQI Med | meas | m_thyles_EQI_7 | r_MedStiffness_p | bilateral | 995400\ C99-5100-070\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-070, ShearWave Elastography Stiffness sample 7 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | EQI Med | meas | m_thyles_EQI_8 | r_MedStiffness_p | bilateral | 995400\ C99-5100-071\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-071, ShearWave Elastography Stiffness sample 8 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | EQI Med | meas | m_thyles_EQI_9 | r_MedStiffness_p | bilateral | 995400\ C99-5100-072\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-072, ShearWave Elastography Stiffness sample 9 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | EQI Med | meas | m_thyles_EQI_10 | r_MedStiffness_p | bilateral | 995400\ C99-5100-073\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-073, ShearWave Elastography Stiffness sample 10 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | EQI Med Vel | meas | m_thyles_EQI_1 | r_MedStiffness_v | bilateral | 995400\ C99-5100-074\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-074, ShearWave Elastography Velocity sample 1 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | EQI Med Vel | meas | m_thyles_EQI_2 | r_MedStiffness_v | bilateral | 995400\ C99-5100-075\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-075, ShearWave Elastography Velocity sample 2 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | EQI Med Vel | meas | m_thyles_EQI_3 | r_MedStiffness_v | bilateral | 995400\ C99-5100-076\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-076, ShearWave Elastography Velocity sample 3 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | EQI Med Vel | meas | m_thyles_EQI_4 | r_MedStiffness_v | bilateral | 995400\ C99-5100-077\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-077, ShearWave Elastography Velocity sample 4 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | EQI Med Vel | meas | m_thyles_EQI_5 | r_MedStiffness_v | bilateral | 995400\ C99-5100-078\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-078, ShearWave Elastography Velocity sample 5 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | EQI Med Vel | meas | m_thyles_EQI_6 | r_MedStiffness_v | bilateral | 995400\ C99-5100-079\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-079, ShearWave Elastography Velocity sample 6 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | EQI Med Vel | meas | m_thyles_EQI_7 | r_MedStiffness_v | bilateral | 995400\ C99-5100-080\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-080, ShearWave Elastography Velocity sample 7 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | EQI Med Vel | meas | m_thyles_EQI_8 | r_MedStiffness_v | bilateral | 995400\ C99-5100-081\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-081, ShearWave Elastography Velocity sample 8 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | EQI Med Vel | meas | m_thyles_EQI_9 | r_MedStiffness_v | bilateral | 995400\ C99-5100-082\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-082, ShearWave Elastography Velocity sample 9 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | EQI Med Vel | meas | m_thyles_EQI_10 | r_MedStiffness_v | bilateral | 995400\ C99-5100-083\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-083, ShearWave Elastography Velocity sample 10 Median |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | EQI IQR | meas | m_thyles_EQI_4 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-087\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-087, ShearWave Elastography Stiffness sample 4 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | EQI IQR | meas | m_thyles_EQI_5 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-088\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-088, ShearWave Elastography Stiffness sample 5 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | EQI IQR | meas | m_thyles_EQI_6 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-089\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-089, ShearWave Elastography Stiffness sample 6 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | EQI IQR | meas | m_thyles_EQI_7 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-090\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-090, ShearWave Elastography Stiffness sample 7 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | EQI IQR | meas | m_thyles_EQI_8 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-091\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-091, ShearWave Elastography Stiffness sample 8 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | EQI IQR | meas | m_thyles_EQI_9 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-092\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-092, ShearWave Elastography Stiffness sample 9 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | EQI IQR | meas | m_thyles_EQI_10 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-093\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-093, ShearWave Elastography Stiffness sample 10 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | EQI IQR Vel | meas | m_thyles_EQI_1 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-094\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-094, ShearWave Elastography Velocity sample 1 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | EQI IQR Vel | meas | m_thyles_EQI_2 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-095\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-095, ShearWave Elastography Velocity sample 2 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | EQI IQR Vel | meas | m_thyles_EQI_3 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-096\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-096, ShearWave Elastography Velocity sample 3 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | EQI IQR Vel | meas | m_thyles_EQI_4 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-097\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-097, ShearWave Elastography Velocity sample 4 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | EQI IQR Vel | meas | m_thyles_EQI_5 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-098\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-098, ShearWave Elastography Velocity sample 5 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | EQI IQR Vel | meas | m_thyles_EQI_6 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-099\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-099, ShearWave Elastography Velocity sample 6 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | EQI IQR Vel | meas | m_thyles_EQI_7 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-100\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-100, ShearWave Elastography Velocity sample 7 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | EQI IQR Vel | meas | m_thyles_EQI_8 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-101\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-101, ShearWave Elastography Velocity sample 8 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | EQI IQR Vel | meas | m_thyles_EQI_9 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-102\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-102, ShearWave Elastography Velocity sample 9 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | EQI IQR Vel | meas | m_thyles_EQI_10 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-103\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-103, ShearWave Elastography Velocity sample 10 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | EQI IQR/Med | meas | m_thyles_EQI_4 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-107\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-107, ShearWave Elasto Stiffness sample 4 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | EQI IQR/Med | meas | m_thyles_EQI_5 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-108\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-108, ShearWave Elasto Stiffness sample 5 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | EQI IQR/Med | meas | m_thyles_EQI_6 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-109\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-109, ShearWave Elasto Stiffness sample 6 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | EQI IQR/Med | meas | m_thyles_EQI_7 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-110\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-110, ShearWave Elasto Stiffness sample 7 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | EQI IQR/Med | meas | m_thyles_EQI_8 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-111\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-111, ShearWave Elasto Stiffness sample 8 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | EQI IQR/Med | meas | m_thyles_EQI_9 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-112\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-112, ShearWave Elasto Stiffness sample 9 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | EQI IQR/Med | meas | m_thyles_EQI_10 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-113\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-113, ShearWave Elasto Stiffness sample 10 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | EQI IQR/Med Vel | meas | m_thyles_EQI_1 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-114\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-114, ShearWave Elasto Velocity sample 1 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | EQI IQR/Med Vel | meas | m_thyles_EQI_2 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-115\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-115, ShearWave Elasto Velocity sample 2 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | EQI IQR/Med Vel | meas | m_thyles_EQI_3 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-116\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-116, ShearWave Elasto Velocity sample 3 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | EQI IQR/Med Vel | meas | m_thyles_EQI_4 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-117\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-117, ShearWave Elasto Velocity sample 4 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | EQI IQR/Med Vel | meas | m_thyles_EQI_5 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-118\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-118, ShearWave Elasto Velocity sample 5 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | EQI IQR/Med Vel | meas | m_thyles_EQI_6 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-119\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-119, ShearWave Elasto Velocity sample 6 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | EQI IQR/Med Vel | meas | m_thyles_EQI_7 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-120\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-120, ShearWave Elasto Velocity sample 7 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | EQI IQR/Med Vel | meas | m_thyles_EQI_8 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-121\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-121, ShearWave Elasto Velocity sample 8 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | EQI IQR/Med Vel | meas | m_thyles_EQI_9 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-122\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-122, ShearWave Elasto Velocity sample 9 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | EQI IQR/Med Vel | meas | m_thyles_EQI_10 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-123\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-123, ShearWave Elasto Velocity sample 10 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | Conf. Threshold | meas | m_thyles_EQI_4 | r_CFStiffness | bilateral | 995400\ C99-5100-127\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-127, ShearWave Elastography Stiffness sample 4 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | Conf. Threshold | meas | m_thyles_EQI_5 | r_CFStiffness | bilateral | 995400\ C99-5100-128\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-128, ShearWave Elastography Stiffness sample 5 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | Conf. Threshold | meas | m_thyles_EQI_6 | r_CFStiffness | bilateral | 995400\ C99-5100-129\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-129, ShearWave Elastography Stiffness sample 6 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | Conf. Threshold | meas | m_thyles_EQI_7 | r_CFStiffness | bilateral | 995400\ C99-5100-130\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-130, ShearWave Elastography Stiffness sample 7 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | Conf. Threshold | meas | m_thyles_EQI_8 | r_CFStiffness | bilateral | 995400\ C99-5100-131\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-131, ShearWave Elastography Stiffness sample 8 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | Conf. Threshold | meas | m_thyles_EQI_9 | r_CFStiffness | bilateral | 995400\ C99-5100-132\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-132, ShearWave Elastography Stiffness sample 9 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | Conf. Threshold | meas | m_thyles_EQI_10 | r_CFStiffness | bilateral | 995400\ C99-5100-133\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-133, ShearWave Elastography Stiffness sample 10 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | Circ | meas | m_thyles_EQI_4 | r_Circumference | bilateral | 995400\ C99-5100-137\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-137, ShearWave Elastography sample 4 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | Circ | meas | m_thyles_EQI_5 | r_Circumference | bilateral | 995400\ C99-5100-138\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-138, ShearWave Elastography sample 5 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | Circ | meas | m_thyles_EQI_6 | r_Circumference | bilateral | 995400\ C99-5100-139\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-139, ShearWave Elastography sample 6 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | Circ | meas | m_thyles_EQI_7 | r_Circumference | bilateral | 995400\ C99-5100-140\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-140, ShearWave Elastography sample 7 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | Circ | meas | m_thyles_EQI_8 | r_Circumference | bilateral | 995400\ C99-5100-141\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-141, ShearWave Elastography sample 8 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | Circ | meas | m_thyles_EQI_9 | r_Circumference | bilateral | 995400\ C99-5100-142\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-142, ShearWave Elastography sample 9 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | Circ | meas | m_thyles_EQI_10 | r_Circumference | bilateral | 995400\ C99-5100-143\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-143, ShearWave Elastography sample 10 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | Area | meas | m_thyles_EQI_4 | r_Area | bilateral | 995400\ C99-5100-147\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-147, ShearWave Elastography sample 4 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | Area | meas | m_thyles_EQI_5 | r_Area | bilateral | 995400\ C99-5100-148\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-148, ShearWave Elastography sample 5 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | Area | meas | m_thyles_EQI_6 | r_Area | bilateral | 995400\ C99-5100-149\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-149, ShearWave Elastography sample 6 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | Area | meas | m_thyles_EQI_7 | r_Area | bilateral | 995400\ C99-5100-150\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-150, ShearWave Elastography sample 7 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | Area | meas | m_thyles_EQI_8 | r_Area | bilateral | 995400\ C99-5100-151\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-151, ShearWave Elastography sample 8 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | Area | meas | m_thyles_EQI_9 | r_Area | bilateral | 995400\ C99-5100-152\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-152, ShearWave Elastography sample 9 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | Area | meas | m_thyles_EQI_10 | r_Area | bilateral | 995400\ C99-5100-153\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-153, ShearWave Elastography sample 10 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | EQI Avg | meas | m_thyles_EQI_1 | r_AvgStiffness_p | bilateral | 995400\C4031-05\C99-5100-001\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-001, ShearWave Elastography Stiffness sample 1 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | EQI Avg | meas | m_thyles_EQI_2 | r_AvgStiffness_p | bilateral | 995400\C4031-05\C99-5100-002\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-002, ShearWave Elastography Stiffness sample 2 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | EQI Avg | meas | m_thyles_EQI_3 | r_AvgStiffness_p | bilateral | 995400\C4031-05\C99-5100-003\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-003, ShearWave Elastography Stiffness sample 3 Average |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | EQI Std | meas | m_thyles_EQI_1 | r_StdStiffness_p | bilateral | 995400\C4031-05\C99-5100-044\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-044, ShearWave Elastography Stiffness sample 1 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | EQI Std | meas | m_thyles_EQI_2 | r_StdStiffness_p | bilateral | 995400\C4031-05\C99-5100-045\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-045, ShearWave Elastography Stiffness sample 2 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | EQI Std | meas | m_thyles_EQI_3 | r_StdStiffness_p | bilateral | 995400\C4031-05\C99-5100-046\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-046, ShearWave Elastography Stiffness sample 3 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | EQI Med | meas | m_thyles_EQI_1 | r_MedStiffness_p | bilateral | 995400\C4031-05\C99-5100-064\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-064, ShearWave Elastography Stiffness sample 1 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | EQI Med | meas | m_thyles_EQI_2 | r_MedStiffness_p | bilateral | 995400\C4031-05\C99-5100-065\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-065, ShearWave Elastography Stiffness sample 2 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | EQI Med | meas | m_thyles_EQI_3 | r_MedStiffness_p | bilateral | 995400\C4031-05\C99-5100-066\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-066, ShearWave Elastography Stiffness sample 3 Median |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | EQI IQR | meas | m_thyles_EQI_1 | r_IqrStiffness_p | bilateral | 995400\C4031-05\C99-5100-084\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-084, ShearWave Elastography Stiffness sample 1 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | EQI IQR | meas | m_thyles_EQI_2 | r_IqrStiffness_p | bilateral | 995400\C4031-05\C99-5100-085\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-085, ShearWave Elastography Stiffness sample 2 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | EQI IQR | meas | m_thyles_EQI_3 | r_IqrStiffness_p | bilateral | 995400\C4031-05\C99-5100-086\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-086, ShearWave Elastography Stiffness sample 3 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | EQI IQR/Med | meas | m_thyles_EQI_1 | r_IqrMedStiffness_p | bilateral | 995400\C4031-05\C99-5100-104\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-104, ShearWave Elasto Stiffness sample 1 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | EQI IQR/Med | meas | m_thyles_EQI_2 | r_IqrMedStiffness_p | bilateral | 995400\C4031-05\C99-5100-105\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-105, ShearWave Elasto Stiffness sample 2 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | EQI IQR/Med | meas | m_thyles_EQI_3 | r_IqrMedStiffness_p | bilateral | 995400\C4031-05\C99-5100-106\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-106, ShearWave Elasto Stiffness sample 3 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | Conf. Threshold | meas | m_thyles_EQI_1 | r_CFStiffness | bilateral | 995400\C4031-05\C99-5100-124\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-124, ShearWave Elastography Stiffness sample 1 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | Conf. Threshold | meas | m_thyles_EQI_2 | r_CFStiffness | bilateral | 995400\C4031-05\C99-5100-125\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-125, ShearWave Elastography Stiffness sample 2 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | Conf. Threshold | meas | m_thyles_EQI_3 | r_CFStiffness | bilateral | 995400\C4031-05\C99-5100-126\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-126, ShearWave Elastography Stiffness sample 3 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | Circ | meas | m_thyles_EQI_1 | r_Circumference | bilateral | 995400\C4031-05\C99-5100-134\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-134, ShearWave Elastography sample 1 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | Circ | meas | m_thyles_EQI_2 | r_Circumference | bilateral | 995400\C4031-05\C99-5100-135\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-135, ShearWave Elastography sample 2 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | Circ | meas | m_thyles_EQI_3 | r_Circumference | bilateral | 995400\C4031-05\C99-5100-136\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-136, ShearWave Elastography sample 3 Circumference |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | Area | meas | m_thyles_EQI_1 | r_Area | bilateral | 995400\C4031-05\C99-5100-144\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-144, ShearWave Elastography sample 1 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | Area | meas | m_thyles_EQI_2 | r_Area | bilateral | 995400\C4031-05\C99-5100-145\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-145, ShearWave Elastography sample 2 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | Area | meas | m_thyles_EQI_3 | r_Area | bilateral | 995400\C4031-05\C99-5100-146\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-146, ShearWave Elastography sample 3 Area |  | UCUM, mm2, mm2 | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | EQI Min | meas | m_thyles_EQI_1 | r_MinStiffness_p | bilateral | 995400\C4031-05\C99-5100-154\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-154, ShearWave Elastography Stiffness sample 1 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | EQI Min | meas | m_thyles_EQI_2 | r_MinStiffness_p | bilateral | 995400\C4031-05\C99-5100-155\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-155, ShearWave Elastography Stiffness sample 2 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | EQI Min | meas | m_thyles_EQI_3 | r_MinStiffness_p | bilateral | 995400\C4031-05\C99-5100-156\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-156, ShearWave Elastography Stiffness sample 3 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | EQI Min | meas | m_thyles_EQI_4 | r_MinStiffness_p | bilateral | 995400\C4031-05\C99-5100-157\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-157, ShearWave Elastography Stiffness sample 4 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | EQI Min | meas | m_thyles_EQI_5 | r_MinStiffness_p | bilateral | 995400\C4031-05\C99-5100-158\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-158, ShearWave Elastography Stiffness sample 5 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | EQI Min | meas | m_thyles_EQI_6 | r_MinStiffness_p | bilateral | 995400\C4031-05\C99-5100-159\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-159, ShearWave Elastography Stiffness sample 6 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | EQI Min | meas | m_thyles_EQI_7 | r_MinStiffness_p | bilateral | 995400\C4031-05\C99-5100-160\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-160, ShearWave Elastography Stiffness sample 7 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | EQI Min | meas | m_thyles_EQI_8 | r_MinStiffness_p | bilateral | 995400\C4031-05\C99-5100-161\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-161, ShearWave Elastography Stiffness sample 8 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | EQI Min | meas | m_thyles_EQI_9 | r_MinStiffness_p | bilateral | 995400\C4031-05\C99-5100-162\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-162, ShearWave Elastography Stiffness sample 9 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | EQI Min | meas | m_thyles_EQI_10 | r_MinStiffness_p | bilateral | 995400\C4031-05\C99-5100-163\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-163, ShearWave Elastography Stiffness sample 10 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | EQI Min Vel | meas | m_thyles_EQI_1 | r_MinStiffness_v | bilateral | 995400\C4031-05\C99-5100-164\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-164, ShearWave Elastography Velocity sample 1 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | EQI Min Vel | meas | m_thyles_EQI_2 | r_MinStiffness_v | bilateral | 995400\C4031-05\C99-5100-165\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-165, ShearWave Elastography Velocity sample 2 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | EQI Min Vel | meas | m_thyles_EQI_3 | r_MinStiffness_v | bilateral | 995400\C4031-05\C99-5100-166\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-166, ShearWave Elastography Velocity sample 3 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | EQI Min Vel | meas | m_thyles_EQI_4 | r_MinStiffness_v | bilateral | 995400\C4031-05\C99-5100-167\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-167, ShearWave Elastography Velocity sample 4 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | EQI Min Vel | meas | m_thyles_EQI_5 | r_MinStiffness_v | bilateral | 995400\C4031-05\C99-5100-168\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-168, ShearWave Elastography Velocity sample 5 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | EQI Min Vel | meas | m_thyles_EQI_6 | r_MinStiffness_v | bilateral | 995400\C4031-05\C99-5100-169\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-169, ShearWave Elastography Velocity sample 6 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | EQI Min Vel | meas | m_thyles_EQI_7 | r_MinStiffness_v | bilateral | 995400\C4031-05\C99-5100-170\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-170, ShearWave Elastography Velocity sample 7 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | EQI Min Vel | meas | m_thyles_EQI_8 | r_MinStiffness_v | bilateral | 995400\C4031-05\C99-5100-171 \T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-171 , ShearWave Elastography Velocity sample 8 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | EQI Min Vel | meas | m_thyles_EQI_9 | r_MinStiffness_v | bilateral | 995400\C4031-05\C99-5100-172\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-172, ShearWave Elastography Velocity sample 9 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | EQI Min Vel | meas | m_thyles_EQI_10 | r_MinStiffness_v | bilateral | 995400\C4031-05\C99-5100-173\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-173, ShearWave Elastography Velocity sample 10 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | EQI Max | meas | m_thyles_EQI_1 | r_MaxStiffness_p | bilateral | 995400\C4031-05\C99-5100-174\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-174, ShearWave Elastography Stiffness sample 1 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | EQI Max | meas | m_thyles_EQI_2 | r_MaxStiffness_p | bilateral | 995400\C4031-05\C99-5100-175\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-175, ShearWave Elastography Stiffness sample 2 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | EQI Max | meas | m_thyles_EQI_3 | r_MaxStiffness_p | bilateral | 995400\C4031-05\C99-5100-176\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-176, ShearWave Elastography Stiffness sample 3 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | EQI Max | meas | m_thyles_EQI_4 | r_MaxStiffness_p | bilateral | 995400\C4031-05\C99-5100-177\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-177, ShearWave Elastography Stiffness sample 4 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | EQI Max | meas | m_thyles_EQI_5 | r_MaxStiffness_p | bilateral | 995400\C4031-05\C99-5100-178\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-178, ShearWave Elastography Stiffness sample 5 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | EQI Max | meas | m_thyles_EQI_6 | r_MaxStiffness_p | bilateral | 995400\C4031-05\C99-5100-179\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-179, ShearWave Elastography Stiffness sample 6 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | EQI Max | meas | m_thyles_EQI_7 | r_MaxStiffness_p | bilateral | 995400\C4031-05\C99-5100-180\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-180, ShearWave Elastography Stiffness sample 7 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | EQI Max | meas | m_thyles_EQI_8 | r_MaxStiffness_p | bilateral | 995400\C4031-05\C99-5100-181\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-181, ShearWave Elastography Stiffness sample 8 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | EQI Max | meas | m_thyles_EQI_9 | r_MaxStiffness_p | bilateral | 995400\C4031-05\C99-5100-182\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-182, ShearWave Elastography Stiffness sample 9 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | EQI Max | meas | m_thyles_EQI_10 | r_MaxStiffness_p | bilateral | 995400\C4031-05\C99-5100-183\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-183, ShearWave Elastography Stiffness sample 10 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 1 | EQI Max Vel | meas | m_thyles_EQI_1 | r_MaxStiffness_v | bilateral | 995400\C4031-05\C99-5100-184\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-184, ShearWave Elastography Velocity sample 1 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 2 | EQI Max Vel | meas | m_thyles_EQI_2 | r_MaxStiffness_v | bilateral | 995400\C4031-05\C99-5100-185\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-185, ShearWave Elastography Velocity sample 2 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 3 | EQI Max Vel | meas | m_thyles_EQI_3 | r_MaxStiffness_v | bilateral | 995400\C4031-05\C99-5100-186\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-186, ShearWave Elastography Velocity sample 3 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 4 | EQI Max Vel | meas | m_thyles_EQI_4 | r_MaxStiffness_v | bilateral | 995400\C4031-05\C99-5100-187\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-187, ShearWave Elastography Velocity sample 4 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 5 | EQI Max Vel | meas | m_thyles_EQI_5 | r_MaxStiffness_v | bilateral | 995400\C4031-05\C99-5100-188\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-188, ShearWave Elastography Velocity sample 5 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 6 | EQI Max Vel | meas | m_thyles_EQI_6 | r_MaxStiffness_v | bilateral | 995400\C4031-05\C99-5100-189\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-189, ShearWave Elastography Velocity sample 6 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 7 | EQI Max Vel | meas | m_thyles_EQI_7 | r_MaxStiffness_v | bilateral | 995400\C4031-05\C99-5100-190\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-190, ShearWave Elastography Velocity sample 7 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 8 | EQI Max Vel | meas | m_thyles_EQI_8 | r_MaxStiffness_v | bilateral | 995400\C4031-05\C99-5100-191 \T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-191 , ShearWave Elastography Velocity sample 8 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 9 | EQI Max Vel | meas | m_thyles_EQI_9 | r_MaxStiffness_v | bilateral | 995400\C4031-05\C99-5100-192\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-192, ShearWave Elastography Velocity sample 9 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI 10 | EQI Max Vel | meas | m_thyles_EQI_10 | r_MaxStiffness_v | bilateral | 995400\C4031-05\C99-5100-193\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-193, ShearWave Elastography Velocity sample 10 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | EQI Avg | meas | m_breast_EQI_4 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-004\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-004, ShearWave Elastography Stiffness sample 4 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | EQI Avg | meas | m_breast_EQI_5 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-005\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-005, ShearWave Elastography Stiffness sample 5 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | EQI Avg | meas | m_breast_EQI_6 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-006\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-006, ShearWave Elastography Stiffness sample 6 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | EQI Avg | meas | m_breast_EQI_7 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-007\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-007, ShearWave Elastography Stiffness sample 7 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | EQI Avg | meas | m_breast_EQI_8 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-008\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-008, ShearWave Elastography Stiffness sample 8 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | EQI Avg | meas | m_breast_EQI_9 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-009\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-009, ShearWave Elastography Stiffness sample 9 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | EQI Avg | meas | m_breast_EQI_10 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-010\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-010, ShearWave Elastography Stiffness sample 10 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | EQI Avg Vel | meas | m_breast_EQI_1 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-011\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-011, ShearWave Elastography Velocity sample 1 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | EQI Avg Vel | meas | m_breast_EQI_2 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-012\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-012, ShearWave Elastography Velocity sample 2 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | EQI Avg Vel | meas | m_breast_EQI_3 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-013\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-013, ShearWave Elastography Velocity sample 3 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | EQI Avg Vel | meas | m_breast_EQI_4 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-014\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-014, ShearWave Elastography Velocity sample 4 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | EQI Avg Vel | meas | m_breast_EQI_5 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-015\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-015, ShearWave Elastography Velocity sample 5 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | EQI Avg Vel | meas | m_breast_EQI_6 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-016\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-016, ShearWave Elastography Velocity sample 6 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | EQI Avg Vel | meas | m_breast_EQI_7 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-017\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-017, ShearWave Elastography Velocity sample 7 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | EQI Avg Vel | meas | m_breast_EQI_8 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-018 \T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-018 , ShearWave Elastography Velocity sample 8 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | EQI Avg Vel | meas | m_breast_EQI_9 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-019\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-019, ShearWave Elastography Velocity sample 9 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | EQI Avg Vel | meas | m_breast_EQI_10 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-020\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-020, ShearWave Elastography Velocity sample 10 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | EQI Std | meas | m_breast_EQI_4 | r_StdStiffness_p | bilateral | 995400\ C99-5100-047\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-047, ShearWave Elastography Stiffness sample 4 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | EQI Std | meas | m_breast_EQI_5 | r_StdStiffness_p | bilateral | 995400\ C99-5100-048\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-048, ShearWave Elastography Stiffness sample 5 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | EQI Std | meas | m_breast_EQI_6 | r_StdStiffness_p | bilateral | 995400\ C99-5100-049\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-049, ShearWave Elastography Stiffness sample 6 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | EQI Std | meas | m_breast_EQI_7 | r_StdStiffness_p | bilateral | 995400\ C99-5100-050\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-050, ShearWave Elastography Stiffness sample 7 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | EQI Std | meas | m_breast_EQI_8 | r_StdStiffness_p | bilateral | 995400\ C99-5100-051\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-051, ShearWave Elastography Stiffness sample 8 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | EQI Std | meas | m_breast_EQI_9 | r_StdStiffness_p | bilateral | 995400\ C99-5100-052\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-052, ShearWave Elastography Stiffness sample 9 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | EQI Std | meas | m_breast_EQI_10 | r_StdStiffness_p | bilateral | 995400\ C99-5100-053\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-053, ShearWave Elastography Stiffness sample 10 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | EQI Std Vel | meas | m_breast_EQI_1 | r_StdStiffness_v | bilateral | 995400\ C99-5100-054\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-054, ShearWave Elastography Velocity sample 1 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | EQI Std Vel | meas | m_breast_EQI_2 | r_StdStiffness_v | bilateral | 995400\ C99-5100-055\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-055, ShearWave Elastography Velocity sample 2 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | EQI Std Vel | meas | m_breast_EQI_3 | r_StdStiffness_v | bilateral | 995400\ C99-5100-056\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-056, ShearWave Elastography Velocity sample 3 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | EQI Std Vel | meas | m_breast_EQI_4 | r_StdStiffness_v | bilateral | 995400\ C99-5100-057\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-057, ShearWave Elastography Velocity sample 4 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | EQI Std Vel | meas | m_breast_EQI_5 | r_StdStiffness_v | bilateral | 995400\ C99-5100-058\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-058, ShearWave Elastography Velocity sample 5 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | EQI Std Vel | meas | m_breast_EQI_6 | r_StdStiffness_v | bilateral | 995400\ C99-5100-059\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-059, ShearWave Elastography Velocity sample 6 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | EQI Std Vel | meas | m_breast_EQI_7 | r_StdStiffness_v | bilateral | 995400\ C99-5100-060\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-060, ShearWave Elastography Velocity sample 7 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | EQI Std Vel | meas | m_breast_EQI_8 | r_StdStiffness_v | bilateral | 995400\ C99-5100-061\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-061, ShearWave Elastography Velocity sample 8 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | EQI Std Vel | meas | m_breast_EQI_9 | r_StdStiffness_v | bilateral | 995400\ C99-5100-062\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-062, ShearWave Elastography Velocity sample 9 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | EQI Std Vel | meas | m_breast_EQI_10 | r_StdStiffness_v | bilateral | 995400\ C99-5100-063\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-063, ShearWave Elastography Velocity sample 10 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | EQI Med | meas | m_breast_EQI_4 | r_MedStiffness_p | bilateral | 995400\ C99-5100-067\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-067, ShearWave Elastography Stiffness sample 4 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | EQI Med | meas | m_breast_EQI_5 | r_MedStiffness_p | bilateral | 995400\ C99-5100-068\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-068, ShearWave Elastography Stiffness sample 5 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | EQI Med | meas | m_breast_EQI_6 | r_MedStiffness_p | bilateral | 995400\ C99-5100-069\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-069, ShearWave Elastography Stiffness sample 6 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | EQI Med | meas | m_breast_EQI_7 | r_MedStiffness_p | bilateral | 995400\ C99-5100-070\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-070, ShearWave Elastography Stiffness sample 7 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | EQI Med | meas | m_breast_EQI_8 | r_MedStiffness_p | bilateral | 995400\ C99-5100-071\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-071, ShearWave Elastography Stiffness sample 8 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | EQI Med | meas | m_breast_EQI_9 | r_MedStiffness_p | bilateral | 995400\ C99-5100-072\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-072, ShearWave Elastography Stiffness sample 9 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | EQI Med | meas | m_breast_EQI_10 | r_MedStiffness_p | bilateral | 995400\ C99-5100-073\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-073, ShearWave Elastography Stiffness sample 10 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | EQI Med Vel | meas | m_breast_EQI_1 | r_MedStiffness_v | bilateral | 995400\ C99-5100-074\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-074, ShearWave Elastography Velocity sample 1 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | EQI Med Vel | meas | m_breast_EQI_2 | r_MedStiffness_v | bilateral | 995400\ C99-5100-075\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-075, ShearWave Elastography Velocity sample 2 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | EQI Med Vel | meas | m_breast_EQI_3 | r_MedStiffness_v | bilateral | 995400\ C99-5100-076\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-076, ShearWave Elastography Velocity sample 3 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | EQI Med Vel | meas | m_breast_EQI_4 | r_MedStiffness_v | bilateral | 995400\ C99-5100-077\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-077, ShearWave Elastography Velocity sample 4 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | EQI Med Vel | meas | m_breast_EQI_5 | r_MedStiffness_v | bilateral | 995400\ C99-5100-078\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-078, ShearWave Elastography Velocity sample 5 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | EQI Med Vel | meas | m_breast_EQI_6 | r_MedStiffness_v | bilateral | 995400\ C99-5100-079\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-079, ShearWave Elastography Velocity sample 6 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | EQI Med Vel | meas | m_breast_EQI_7 | r_MedStiffness_v | bilateral | 995400\ C99-5100-080\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-080, ShearWave Elastography Velocity sample 7 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | EQI Med Vel | meas | m_breast_EQI_8 | r_MedStiffness_v | bilateral | 995400\ C99-5100-081\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-081, ShearWave Elastography Velocity sample 8 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | EQI Med Vel | meas | m_breast_EQI_9 | r_MedStiffness_v | bilateral | 995400\ C99-5100-082\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-082, ShearWave Elastography Velocity sample 9 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | EQI Med Vel | meas | m_breast_EQI_10 | r_MedStiffness_v | bilateral | 995400\ C99-5100-083\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-083, ShearWave Elastography Velocity sample 10 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | EQI IQR | meas | m_breast_EQI_4 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-087\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-087, ShearWave Elastography Stiffness sample 4 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | EQI IQR | meas | m_breast_EQI_5 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-088\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-088, ShearWave Elastography Stiffness sample 5 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | EQI IQR | meas | m_breast_EQI_6 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-089\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-089, ShearWave Elastography Stiffness sample 6 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | EQI IQR | meas | m_breast_EQI_7 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-090\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-090, ShearWave Elastography Stiffness sample 7 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | EQI IQR | meas | m_breast_EQI_8 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-091\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-091, ShearWave Elastography Stiffness sample 8 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | EQI IQR | meas | m_breast_EQI_9 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-092\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-092, ShearWave Elastography Stiffness sample 9 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | EQI IQR | meas | m_breast_EQI_10 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-093\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-093, ShearWave Elastography Stiffness sample 10 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | EQI IQR Vel | meas | m_breast_EQI_1 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-094\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-094, ShearWave Elastography Velocity sample 1 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | EQI IQR Vel | meas | m_breast_EQI_2 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-095\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-095, ShearWave Elastography Velocity sample 2 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | EQI IQR Vel | meas | m_breast_EQI_3 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-096\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-096, ShearWave Elastography Velocity sample 3 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | EQI IQR Vel | meas | m_breast_EQI_4 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-097\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-097, ShearWave Elastography Velocity sample 4 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | EQI IQR Vel | meas | m_breast_EQI_5 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-098\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-098, ShearWave Elastography Velocity sample 5 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | EQI IQR Vel | meas | m_breast_EQI_6 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-099\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-099, ShearWave Elastography Velocity sample 6 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | EQI IQR Vel | meas | m_breast_EQI_7 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-100\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-100, ShearWave Elastography Velocity sample 7 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | EQI IQR Vel | meas | m_breast_EQI_8 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-101\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-101, ShearWave Elastography Velocity sample 8 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | EQI IQR Vel | meas | m_breast_EQI_9 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-102\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-102, ShearWave Elastography Velocity sample 9 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | EQI IQR Vel | meas | m_breast_EQI_10 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-103\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-103, ShearWave Elastography Velocity sample 10 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | EQI IQR/Med | meas | m_breast_EQI_4 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-107\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-107, ShearWave Elasto Stiffness sample 4 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | EQI IQR/Med | meas | m_breast_EQI_5 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-108\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-108, ShearWave Elasto Stiffness sample 5 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | EQI IQR/Med | meas | m_breast_EQI_6 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-109\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-109, ShearWave Elasto Stiffness sample 6 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | EQI IQR/Med | meas | m_breast_EQI_7 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-110\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-110, ShearWave Elasto Stiffness sample 7 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | EQI IQR/Med | meas | m_breast_EQI_8 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-111\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-111, ShearWave Elasto Stiffness sample 8 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | EQI IQR/Med | meas | m_breast_EQI_9 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-112\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-112, ShearWave Elasto Stiffness sample 9 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | EQI IQR/Med | meas | m_breast_EQI_10 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-113\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-113, ShearWave Elasto Stiffness sample 10 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | EQI IQR/Med Vel | meas | m_breast_EQI_1 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-114\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-114, ShearWave Elasto Velocity sample 1 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | EQI IQR/Med Vel | meas | m_breast_EQI_2 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-115\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-115, ShearWave Elasto Velocity sample 2 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | EQI IQR/Med Vel | meas | m_breast_EQI_3 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-116\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-116, ShearWave Elasto Velocity sample 3 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | EQI IQR/Med Vel | meas | m_breast_EQI_4 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-117\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-117, ShearWave Elasto Velocity sample 4 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | EQI IQR/Med Vel | meas | m_breast_EQI_5 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-118\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-118, ShearWave Elasto Velocity sample 5 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | EQI IQR/Med Vel | meas | m_breast_EQI_6 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-119\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-119, ShearWave Elasto Velocity sample 6 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | EQI IQR/Med Vel | meas | m_breast_EQI_7 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-120\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-120, ShearWave Elasto Velocity sample 7 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | EQI IQR/Med Vel | meas | m_breast_EQI_8 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-121\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-121, ShearWave Elasto Velocity sample 8 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | EQI IQR/Med Vel | meas | m_breast_EQI_9 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-122\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-122, ShearWave Elasto Velocity sample 9 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | EQI IQR/Med Vel | meas | m_breast_EQI_10 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-123\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-123, ShearWave Elasto Velocity sample 10 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | Conf. Threshold | meas | m_breast_EQI_4 | r_CFStiffness | bilateral | 995400\ C99-5100-127\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-127, ShearWave Elastography Stiffness sample 4 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | Conf. Threshold | meas | m_breast_EQI_5 | r_CFStiffness | bilateral | 995400\ C99-5100-128\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-128, ShearWave Elastography Stiffness sample 5 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | Conf. Threshold | meas | m_breast_EQI_6 | r_CFStiffness | bilateral | 995400\ C99-5100-129\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-129, ShearWave Elastography Stiffness sample 6 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | Conf. Threshold | meas | m_breast_EQI_7 | r_CFStiffness | bilateral | 995400\ C99-5100-130\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-130, ShearWave Elastography Stiffness sample 7 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | Conf. Threshold | meas | m_breast_EQI_8 | r_CFStiffness | bilateral | 995400\ C99-5100-131\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-131, ShearWave Elastography Stiffness sample 8 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | Conf. Threshold | meas | m_breast_EQI_9 | r_CFStiffness | bilateral | 995400\ C99-5100-132\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-132, ShearWave Elastography Stiffness sample 9 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | Conf. Threshold | meas | m_breast_EQI_10 | r_CFStiffness | bilateral | 995400\ C99-5100-133\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-133, ShearWave Elastography Stiffness sample 10 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | Circ | meas | m_breast_EQI_4 | r_Circumference | bilateral | 995400\ C99-5100-137\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-137, ShearWave Elastography sample 4 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | Circ | meas | m_breast_EQI_5 | r_Circumference | bilateral | 995400\ C99-5100-138\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-138, ShearWave Elastography sample 5 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | Circ | meas | m_breast_EQI_6 | r_Circumference | bilateral | 995400\ C99-5100-139\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-139, ShearWave Elastography sample 6 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | Circ | meas | m_breast_EQI_7 | r_Circumference | bilateral | 995400\ C99-5100-140\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-140, ShearWave Elastography sample 7 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | Circ | meas | m_breast_EQI_8 | r_Circumference | bilateral | 995400\ C99-5100-141\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-141, ShearWave Elastography sample 8 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | Circ | meas | m_breast_EQI_9 | r_Circumference | bilateral | 995400\ C99-5100-142\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-142, ShearWave Elastography sample 9 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | Circ | meas | m_breast_EQI_10 | r_Circumference | bilateral | 995400\ C99-5100-143\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-143, ShearWave Elastography sample 10 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | Area | meas | m_breast_EQI_4 | r_Area | bilateral | 995400\ C99-5100-147\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-147, ShearWave Elastography sample 4 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | Area | meas | m_breast_EQI_5 | r_Area | bilateral | 995400\ C99-5100-148\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-148, ShearWave Elastography sample 5 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | Area | meas | m_breast_EQI_6 | r_Area | bilateral | 995400\ C99-5100-149\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-149, ShearWave Elastography sample 6 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | Area | meas | m_breast_EQI_7 | r_Area | bilateral | 995400\ C99-5100-150\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-150, ShearWave Elastography sample 7 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | Area | meas | m_breast_EQI_8 | r_Area | bilateral | 995400\ C99-5100-151\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-151, ShearWave Elastography sample 8 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | Area | meas | m_breast_EQI_9 | r_Area | bilateral | 995400\ C99-5100-152\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-152, ShearWave Elastography sample 9 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | Area | meas | m_breast_EQI_10 | r_Area | bilateral | 995400\ C99-5100-153\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-153, ShearWave Elastography sample 10 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | EQI Avg | meas | m_breast_EQI_1 | r_AvgStiffness_p | bilateral | 995400\C99-5100-001\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-001, ShearWave Elastography Stiffness sample 1 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | EQI Avg | meas | m_breast_EQI_2 | r_AvgStiffness_p | bilateral | 995400\C99-5100-002\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-002, ShearWave Elastography Stiffness sample 2 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | EQI Avg | meas | m_breast_EQI_3 | r_AvgStiffness_p | bilateral | 995400\C99-5100-003\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-003, ShearWave Elastography Stiffness sample 3 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | EQI Std | meas | m_breast_EQI_1 | r_StdStiffness_p | bilateral | 995400\C99-5100-044\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-044, ShearWave Elastography Stiffness sample 1 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | EQI Std | meas | m_breast_EQI_2 | r_StdStiffness_p | bilateral | 995400\C99-5100-045\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-045, ShearWave Elastography Stiffness sample 2 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | EQI Std | meas | m_breast_EQI_3 | r_StdStiffness_p | bilateral | 995400\C99-5100-046\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-046, ShearWave Elastography Stiffness sample 3 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | EQI Med | meas | m_breast_EQI_1 | r_MedStiffness_p | bilateral | 995400\C99-5100-064\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-064, ShearWave Elastography Stiffness sample 1 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | EQI Med | meas | m_breast_EQI_2 | r_MedStiffness_p | bilateral | 995400\C99-5100-065\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-065, ShearWave Elastography Stiffness sample 2 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | EQI Med | meas | m_breast_EQI_3 | r_MedStiffness_p | bilateral | 995400\C99-5100-066\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-066, ShearWave Elastography Stiffness sample 3 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | EQI IQR | meas | m_breast_EQI_1 | r_IqrStiffness_p | bilateral | 995400\C99-5100-084\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-084, ShearWave Elastography Stiffness sample 1 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | EQI IQR | meas | m_breast_EQI_2 | r_IqrStiffness_p | bilateral | 995400\C99-5100-085\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-085, ShearWave Elastography Stiffness sample 2 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | EQI IQR | meas | m_breast_EQI_3 | r_IqrStiffness_p | bilateral | 995400\C99-5100-086\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-086, ShearWave Elastography Stiffness sample 3 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | EQI IQR/Med | meas | m_breast_EQI_1 | r_IqrMedStiffness_p | bilateral | 995400\C99-5100-104\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-104, ShearWave Elasto Stiffness sample 1 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | EQI IQR/Med | meas | m_breast_EQI_2 | r_IqrMedStiffness_p | bilateral | 995400\C99-5100-105\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-105, ShearWave Elasto Stiffness sample 2 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | EQI IQR/Med | meas | m_breast_EQI_3 | r_IqrMedStiffness_p | bilateral | 995400\C99-5100-106\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-106, ShearWave Elasto Stiffness sample 3 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | Conf. Threshold | meas | m_breast_EQI_1 | r_CFStiffness | bilateral | 995400\C99-5100-124\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-124, ShearWave Elastography Stiffness sample 1 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | Conf. Threshold | meas | m_breast_EQI_2 | r_CFStiffness | bilateral | 995400\C99-5100-125\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-125, ShearWave Elastography Stiffness sample 2 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | Conf. Threshold | meas | m_breast_EQI_3 | r_CFStiffness | bilateral | 995400\C99-5100-126\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-126, ShearWave Elastography Stiffness sample 3 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | Circ | meas | m_breast_EQI_1 | r_Circumference | bilateral | 995400\C99-5100-134\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-134, ShearWave Elastography sample 1 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | Circ | meas | m_breast_EQI_2 | r_Circumference | bilateral | 995400\C99-5100-135\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-135, ShearWave Elastography sample 2 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | Circ | meas | m_breast_EQI_3 | r_Circumference | bilateral | 995400\C99-5100-136\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-136, ShearWave Elastography sample 3 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | Area | meas | m_breast_EQI_1 | r_Area | bilateral | 995400\C99-5100-144\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-144, ShearWave Elastography sample 1 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | Area | meas | m_breast_EQI_2 | r_Area | bilateral | 995400\C99-5100-145\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-145, ShearWave Elastography sample 2 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | Area | meas | m_breast_EQI_3 | r_Area | bilateral | 995400\C99-5100-146\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-146, ShearWave Elastography sample 3 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | EQI Min | meas | m_breast_EQI_1 | r_MinStiffness_p | bilateral | 995400\C99-5100-154\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-154, ShearWave Elastography Stiffness sample 1 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | EQI Min | meas | m_breast_EQI_2 | r_MinStiffness_p | bilateral | 995400\C99-5100-155\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-155, ShearWave Elastography Stiffness sample 2 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | EQI Min | meas | m_breast_EQI_3 | r_MinStiffness_p | bilateral | 995400\C99-5100-156\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-156, ShearWave Elastography Stiffness sample 3 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | EQI Min | meas | m_breast_EQI_4 | r_MinStiffness_p | bilateral | 995400\C99-5100-157\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-157, ShearWave Elastography Stiffness sample 4 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | EQI Min | meas | m_breast_EQI_5 | r_MinStiffness_p | bilateral | 995400\C99-5100-158\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-158, ShearWave Elastography Stiffness sample 5 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | EQI Min | meas | m_breast_EQI_6 | r_MinStiffness_p | bilateral | 995400\C99-5100-159\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-159, ShearWave Elastography Stiffness sample 6 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | EQI Min | meas | m_breast_EQI_7 | r_MinStiffness_p | bilateral | 995400\C99-5100-160\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-160, ShearWave Elastography Stiffness sample 7 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | EQI Min | meas | m_breast_EQI_8 | r_MinStiffness_p | bilateral | 995400\C99-5100-161\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-161, ShearWave Elastography Stiffness sample 8 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | EQI Min | meas | m_breast_EQI_9 | r_MinStiffness_p | bilateral | 995400\C99-5100-162\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-162, ShearWave Elastography Stiffness sample 9 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | EQI Min | meas | m_breast_EQI_10 | r_MinStiffness_p | bilateral | 995400\C99-5100-163\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-163, ShearWave Elastography Stiffness sample 10 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | EQI Min Vel | meas | m_breast_EQI_1 | r_MinStiffness_v | bilateral | 995400\C99-5100-164\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-164, ShearWave Elastography Velocity sample 1 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | EQI Min Vel | meas | m_breast_EQI_2 | r_MinStiffness_v | bilateral | 995400\C99-5100-165\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-165, ShearWave Elastography Velocity sample 2 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | EQI Min Vel | meas | m_breast_EQI_3 | r_MinStiffness_v | bilateral | 995400\C99-5100-166\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-166, ShearWave Elastography Velocity sample 3 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | EQI Min Vel | meas | m_breast_EQI_4 | r_MinStiffness_v | bilateral | 995400\C99-5100-167\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-167, ShearWave Elastography Velocity sample 4 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | EQI Min Vel | meas | m_breast_EQI_5 | r_MinStiffness_v | bilateral | 995400\C99-5100-168\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-168, ShearWave Elastography Velocity sample 5 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | EQI Min Vel | meas | m_breast_EQI_6 | r_MinStiffness_v | bilateral | 995400\C99-5100-169\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-169, ShearWave Elastography Velocity sample 6 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | EQI Min Vel | meas | m_breast_EQI_7 | r_MinStiffness_v | bilateral | 995400\C99-5100-170\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-170, ShearWave Elastography Velocity sample 7 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | EQI Min Vel | meas | m_breast_EQI_8 | r_MinStiffness_v | bilateral | 995400\C99-5100-171 \T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-171 , ShearWave Elastography Velocity sample 8 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | EQI Min Vel | meas | m_breast_EQI_9 | r_MinStiffness_v | bilateral | 995400\C99-5100-172\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-172, ShearWave Elastography Velocity sample 9 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | EQI Min Vel | meas | m_breast_EQI_10 | r_MinStiffness_v | bilateral | 995400\C99-5100-173\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-173, ShearWave Elastography Velocity sample 10 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | EQI Max | meas | m_breast_EQI_1 | r_MaxStiffness_p | bilateral | 995400\C99-5100-174\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-174, ShearWave Elastography Stiffness sample 1 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | EQI Max | meas | m_breast_EQI_2 | r_MaxStiffness_p | bilateral | 995400\C99-5100-175\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-175, ShearWave Elastography Stiffness sample 2 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | EQI Max | meas | m_breast_EQI_3 | r_MaxStiffness_p | bilateral | 995400\C99-5100-176\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-176, ShearWave Elastography Stiffness sample 3 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | EQI Max | meas | m_breast_EQI_4 | r_MaxStiffness_p | bilateral | 995400\C99-5100-177\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-177, ShearWave Elastography Stiffness sample 4 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | EQI Max | meas | m_breast_EQI_5 | r_MaxStiffness_p | bilateral | 995400\C99-5100-178\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-178, ShearWave Elastography Stiffness sample 5 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | EQI Max | meas | m_breast_EQI_6 | r_MaxStiffness_p | bilateral | 995400\C99-5100-179\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-179, ShearWave Elastography Stiffness sample 6 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | EQI Max | meas | m_breast_EQI_7 | r_MaxStiffness_p | bilateral | 995400\C99-5100-180\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-180, ShearWave Elastography Stiffness sample 7 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | EQI Max | meas | m_breast_EQI_8 | r_MaxStiffness_p | bilateral | 995400\C99-5100-181\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-181, ShearWave Elastography Stiffness sample 8 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | EQI Max | meas | m_breast_EQI_9 | r_MaxStiffness_p | bilateral | 995400\C99-5100-182\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-182, ShearWave Elastography Stiffness sample 9 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | EQI Max | meas | m_breast_EQI_10 | r_MaxStiffness_p | bilateral | 995400\C99-5100-183\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-183, ShearWave Elastography Stiffness sample 10 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 1 | EQI Max Vel | meas | m_breast_EQI_1 | r_MaxStiffness_v | bilateral | 995400\C99-5100-184\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-184, ShearWave Elastography Velocity sample 1 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 2 | EQI Max Vel | meas | m_breast_EQI_2 | r_MaxStiffness_v | bilateral | 995400\C99-5100-185\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-185, ShearWave Elastography Velocity sample 2 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 3 | EQI Max Vel | meas | m_breast_EQI_3 | r_MaxStiffness_v | bilateral | 995400\C99-5100-186\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-186, ShearWave Elastography Velocity sample 3 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 4 | EQI Max Vel | meas | m_breast_EQI_4 | r_MaxStiffness_v | bilateral | 995400\C99-5100-187\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-187, ShearWave Elastography Velocity sample 4 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 5 | EQI Max Vel | meas | m_breast_EQI_5 | r_MaxStiffness_v | bilateral | 995400\C99-5100-188\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-188, ShearWave Elastography Velocity sample 5 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 6 | EQI Max Vel | meas | m_breast_EQI_6 | r_MaxStiffness_v | bilateral | 995400\C99-5100-189\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-189, ShearWave Elastography Velocity sample 6 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 7 | EQI Max Vel | meas | m_breast_EQI_7 | r_MaxStiffness_v | bilateral | 995400\C99-5100-190\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-190, ShearWave Elastography Velocity sample 7 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 8 | EQI Max Vel | meas | m_breast_EQI_8 | r_MaxStiffness_v | bilateral | 995400\C99-5100-191 \T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-191 , ShearWave Elastography Velocity sample 8 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 9 | EQI Max Vel | meas | m_breast_EQI_9 | r_MaxStiffness_v | bilateral | 995400\C99-5100-192\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-192, ShearWave Elastography Velocity sample 9 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI 10 | EQI Max Vel | meas | m_breast_EQI_10 | r_MaxStiffness_v | bilateral | 995400\C99-5100-193\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-193, ShearWave Elastography Velocity sample 10 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | EQI Avg | meas | m_thy_EQI_4 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-004\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-004, ShearWave Elastography Stiffness sample 4 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | EQI Avg | meas | m_thy_EQI_5 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-005\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-005, ShearWave Elastography Stiffness sample 5 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | EQI Avg | meas | m_thy_EQI_6 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-006\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-006, ShearWave Elastography Stiffness sample 6 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | EQI Avg | meas | m_thy_EQI_7 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-007\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-007, ShearWave Elastography Stiffness sample 7 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | EQI Avg | meas | m_thy_EQI_8 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-008\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-008, ShearWave Elastography Stiffness sample 8 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | EQI Avg | meas | m_thy_EQI_9 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-009\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-009, ShearWave Elastography Stiffness sample 9 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | EQI Avg | meas | m_thy_EQI_10 | r_AvgStiffness_p | bilateral | 995400\ C99-5100-010\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-010, ShearWave Elastography Stiffness sample 10 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | EQI Avg Vel | meas | m_thy_EQI_1 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-011\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-011, ShearWave Elastography Velocity sample 1 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | EQI Avg Vel | meas | m_thy_EQI_2 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-012\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-012, ShearWave Elastography Velocity sample 2 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | EQI Avg Vel | meas | m_thy_EQI_3 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-013\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-013, ShearWave Elastography Velocity sample 3 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | EQI Avg Vel | meas | m_thy_EQI_4 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-014\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-014, ShearWave Elastography Velocity sample 4 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | EQI Avg Vel | meas | m_thy_EQI_5 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-015\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-015, ShearWave Elastography Velocity sample 5 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | EQI Avg Vel | meas | m_thy_EQI_6 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-016\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-016, ShearWave Elastography Velocity sample 6 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | EQI Avg Vel | meas | m_thy_EQI_7 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-017\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-017, ShearWave Elastography Velocity sample 7 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | EQI Avg Vel | meas | m_thy_EQI_8 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-018 \T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-018 , ShearWave Elastography Velocity sample 8 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | EQI Avg Vel | meas | m_thy_EQI_9 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-019\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-019, ShearWave Elastography Velocity sample 9 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | EQI Avg Vel | meas | m_thy_EQI_10 | r_AvgStiffness_v | bilateral | 995400\ C99-5100-020\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-020, ShearWave Elastography Velocity sample 10 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | EQI Std | meas | m_thy_EQI_4 | r_StdStiffness_p | bilateral | 995400\ C99-5100-047\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-047, ShearWave Elastography Stiffness sample 4 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | EQI Std | meas | m_thy_EQI_5 | r_StdStiffness_p | bilateral | 995400\ C99-5100-048\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-048, ShearWave Elastography Stiffness sample 5 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | EQI Std | meas | m_thy_EQI_6 | r_StdStiffness_p | bilateral | 995400\ C99-5100-049\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-049, ShearWave Elastography Stiffness sample 6 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | EQI Std | meas | m_thy_EQI_7 | r_StdStiffness_p | bilateral | 995400\ C99-5100-050\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-050, ShearWave Elastography Stiffness sample 7 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | EQI Std | meas | m_thy_EQI_8 | r_StdStiffness_p | bilateral | 995400\ C99-5100-051\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-051, ShearWave Elastography Stiffness sample 8 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | EQI Std | meas | m_thy_EQI_9 | r_StdStiffness_p | bilateral | 995400\ C99-5100-052\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-052, ShearWave Elastography Stiffness sample 9 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | EQI Std | meas | m_thy_EQI_10 | r_StdStiffness_p | bilateral | 995400\ C99-5100-053\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-053, ShearWave Elastography Stiffness sample 10 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | EQI Std Vel | meas | m_thy_EQI_1 | r_StdStiffness_v | bilateral | 995400\ C99-5100-054\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-054, ShearWave Elastography Velocity sample 1 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | EQI Std Vel | meas | m_thy_EQI_2 | r_StdStiffness_v | bilateral | 995400\ C99-5100-055\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-055, ShearWave Elastography Velocity sample 2 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | EQI Std Vel | meas | m_thy_EQI_3 | r_StdStiffness_v | bilateral | 995400\ C99-5100-056\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-056, ShearWave Elastography Velocity sample 3 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | EQI Std Vel | meas | m_thy_EQI_4 | r_StdStiffness_v | bilateral | 995400\ C99-5100-057\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-057, ShearWave Elastography Velocity sample 4 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | EQI Std Vel | meas | m_thy_EQI_5 | r_StdStiffness_v | bilateral | 995400\ C99-5100-058\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-058, ShearWave Elastography Velocity sample 5 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | EQI Std Vel | meas | m_thy_EQI_6 | r_StdStiffness_v | bilateral | 995400\ C99-5100-059\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-059, ShearWave Elastography Velocity sample 6 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | EQI Std Vel | meas | m_thy_EQI_7 | r_StdStiffness_v | bilateral | 995400\ C99-5100-060\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-060, ShearWave Elastography Velocity sample 7 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | EQI Std Vel | meas | m_thy_EQI_8 | r_StdStiffness_v | bilateral | 995400\ C99-5100-061\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-061, ShearWave Elastography Velocity sample 8 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | EQI Std Vel | meas | m_thy_EQI_9 | r_StdStiffness_v | bilateral | 995400\ C99-5100-062\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-062, ShearWave Elastography Velocity sample 9 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | EQI Std Vel | meas | m_thy_EQI_10 | r_StdStiffness_v | bilateral | 995400\ C99-5100-063\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-063, ShearWave Elastography Velocity sample 10 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | EQI Med | meas | m_thy_EQI_4 | r_MedStiffness_p | bilateral | 995400\ C99-5100-067\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-067, ShearWave Elastography Stiffness sample 4 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | EQI Med | meas | m_thy_EQI_5 | r_MedStiffness_p | bilateral | 995400\ C99-5100-068\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-068, ShearWave Elastography Stiffness sample 5 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | EQI Med | meas | m_thy_EQI_6 | r_MedStiffness_p | bilateral | 995400\ C99-5100-069\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-069, ShearWave Elastography Stiffness sample 6 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | EQI Med | meas | m_thy_EQI_7 | r_MedStiffness_p | bilateral | 995400\ C99-5100-070\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-070, ShearWave Elastography Stiffness sample 7 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | EQI Med | meas | m_thy_EQI_8 | r_MedStiffness_p | bilateral | 995400\ C99-5100-071\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-071, ShearWave Elastography Stiffness sample 8 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | EQI Med | meas | m_thy_EQI_9 | r_MedStiffness_p | bilateral | 995400\ C99-5100-072\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-072, ShearWave Elastography Stiffness sample 9 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | EQI Med | meas | m_thy_EQI_10 | r_MedStiffness_p | bilateral | 995400\ C99-5100-073\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-073, ShearWave Elastography Stiffness sample 10 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | EQI Med Vel | meas | m_thy_EQI_1 | r_MedStiffness_v | bilateral | 995400\ C99-5100-074\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-074, ShearWave Elastography Velocity sample 1 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | EQI Med Vel | meas | m_thy_EQI_2 | r_MedStiffness_v | bilateral | 995400\ C99-5100-075\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-075, ShearWave Elastography Velocity sample 2 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | EQI Med Vel | meas | m_thy_EQI_3 | r_MedStiffness_v | bilateral | 995400\ C99-5100-076\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-076, ShearWave Elastography Velocity sample 3 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | EQI Med Vel | meas | m_thy_EQI_4 | r_MedStiffness_v | bilateral | 995400\ C99-5100-077\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-077, ShearWave Elastography Velocity sample 4 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | EQI Med Vel | meas | m_thy_EQI_5 | r_MedStiffness_v | bilateral | 995400\ C99-5100-078\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-078, ShearWave Elastography Velocity sample 5 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | EQI Med Vel | meas | m_thy_EQI_6 | r_MedStiffness_v | bilateral | 995400\ C99-5100-079\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-079, ShearWave Elastography Velocity sample 6 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | EQI Med Vel | meas | m_thy_EQI_7 | r_MedStiffness_v | bilateral | 995400\ C99-5100-080\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-080, ShearWave Elastography Velocity sample 7 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | EQI Med Vel | meas | m_thy_EQI_8 | r_MedStiffness_v | bilateral | 995400\ C99-5100-081\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-081, ShearWave Elastography Velocity sample 8 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | EQI Med Vel | meas | m_thy_EQI_9 | r_MedStiffness_v | bilateral | 995400\ C99-5100-082\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-082, ShearWave Elastography Velocity sample 9 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | EQI Med Vel | meas | m_thy_EQI_10 | r_MedStiffness_v | bilateral | 995400\ C99-5100-083\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-083, ShearWave Elastography Velocity sample 10 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | EQI IQR | meas | m_thy_EQI_4 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-087\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-087, ShearWave Elastography Stiffness sample 4 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | EQI IQR | meas | m_thy_EQI_5 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-088\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-088, ShearWave Elastography Stiffness sample 5 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | EQI IQR | meas | m_thy_EQI_6 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-089\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-089, ShearWave Elastography Stiffness sample 6 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | EQI IQR | meas | m_thy_EQI_7 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-090\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-090, ShearWave Elastography Stiffness sample 7 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | EQI IQR | meas | m_thy_EQI_8 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-091\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-091, ShearWave Elastography Stiffness sample 8 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | EQI IQR | meas | m_thy_EQI_9 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-092\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-092, ShearWave Elastography Stiffness sample 9 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | EQI IQR | meas | m_thy_EQI_10 | r_IqrStiffness_p | bilateral | 995400\ C99-5100-093\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-093, ShearWave Elastography Stiffness sample 10 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | EQI IQR Vel | meas | m_thy_EQI_1 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-094\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-094, ShearWave Elastography Velocity sample 1 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | EQI IQR Vel | meas | m_thy_EQI_2 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-095\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-095, ShearWave Elastography Velocity sample 2 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | EQI IQR Vel | meas | m_thy_EQI_3 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-096\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-096, ShearWave Elastography Velocity sample 3 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | EQI IQR Vel | meas | m_thy_EQI_4 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-097\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-097, ShearWave Elastography Velocity sample 4 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | EQI IQR Vel | meas | m_thy_EQI_5 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-098\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-098, ShearWave Elastography Velocity sample 5 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | EQI IQR Vel | meas | m_thy_EQI_6 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-099\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-099, ShearWave Elastography Velocity sample 6 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | EQI IQR Vel | meas | m_thy_EQI_7 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-100\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-100, ShearWave Elastography Velocity sample 7 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | EQI IQR Vel | meas | m_thy_EQI_8 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-101\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-101, ShearWave Elastography Velocity sample 8 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | EQI IQR Vel | meas | m_thy_EQI_9 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-102\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-102, ShearWave Elastography Velocity sample 9 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | EQI IQR Vel | meas | m_thy_EQI_10 | r_IqrStiffness_v | bilateral | 995400\ C99-5100-103\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-103, ShearWave Elastography Velocity sample 10 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | EQI IQR/Med | meas | m_thy_EQI_4 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-107\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-107, ShearWave Elasto Stiffness sample 4 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | EQI IQR/Med | meas | m_thy_EQI_5 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-108\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-108, ShearWave Elasto Stiffness sample 5 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | EQI IQR/Med | meas | m_thy_EQI_6 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-109\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-109, ShearWave Elasto Stiffness sample 6 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | EQI IQR/Med | meas | m_thy_EQI_7 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-110\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-110, ShearWave Elasto Stiffness sample 7 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | EQI IQR/Med | meas | m_thy_EQI_8 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-111\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-111, ShearWave Elasto Stiffness sample 8 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | EQI IQR/Med | meas | m_thy_EQI_9 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-112\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-112, ShearWave Elasto Stiffness sample 9 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | EQI IQR/Med | meas | m_thy_EQI_10 | r_IqrMedStiffness_p | bilateral | 995400\ C99-5100-113\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-113, ShearWave Elasto Stiffness sample 10 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | EQI IQR/Med Vel | meas | m_thy_EQI_1 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-114\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-114, ShearWave Elasto Velocity sample 1 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | EQI IQR/Med Vel | meas | m_thy_EQI_2 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-115\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-115, ShearWave Elasto Velocity sample 2 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | EQI IQR/Med Vel | meas | m_thy_EQI_3 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-116\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-116, ShearWave Elasto Velocity sample 3 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | EQI IQR/Med Vel | meas | m_thy_EQI_4 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-117\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-117, ShearWave Elasto Velocity sample 4 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | EQI IQR/Med Vel | meas | m_thy_EQI_5 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-118\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-118, ShearWave Elasto Velocity sample 5 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | EQI IQR/Med Vel | meas | m_thy_EQI_6 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-119\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-119, ShearWave Elasto Velocity sample 6 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | EQI IQR/Med Vel | meas | m_thy_EQI_7 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-120\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-120, ShearWave Elasto Velocity sample 7 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | EQI IQR/Med Vel | meas | m_thy_EQI_8 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-121\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-121, ShearWave Elasto Velocity sample 8 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | EQI IQR/Med Vel | meas | m_thy_EQI_9 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-122\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-122, ShearWave Elasto Velocity sample 9 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | EQI IQR/Med Vel | meas | m_thy_EQI_10 | r_IqrMedStiffness_v | bilateral | 995400\ C99-5100-123\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-123, ShearWave Elasto Velocity sample 10 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | Conf. Threshold | meas | m_thy_EQI_4 | r_CFStiffness | bilateral | 995400\ C99-5100-127\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-127, ShearWave Elastography Stiffness sample 4 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | Conf. Threshold | meas | m_thy_EQI_5 | r_CFStiffness | bilateral | 995400\ C99-5100-128\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-128, ShearWave Elastography Stiffness sample 5 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | Conf. Threshold | meas | m_thy_EQI_6 | r_CFStiffness | bilateral | 995400\ C99-5100-129\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-129, ShearWave Elastography Stiffness sample 6 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | Conf. Threshold | meas | m_thy_EQI_7 | r_CFStiffness | bilateral | 995400\ C99-5100-130\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-130, ShearWave Elastography Stiffness sample 7 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | Conf. Threshold | meas | m_thy_EQI_8 | r_CFStiffness | bilateral | 995400\ C99-5100-131\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-131, ShearWave Elastography Stiffness sample 8 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | Conf. Threshold | meas | m_thy_EQI_9 | r_CFStiffness | bilateral | 995400\ C99-5100-132\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-132, ShearWave Elastography Stiffness sample 9 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | Conf. Threshold | meas | m_thy_EQI_10 | r_CFStiffness | bilateral | 995400\ C99-5100-133\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-133, ShearWave Elastography Stiffness sample 10 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | Circ | meas | m_thy_EQI_4 | r_Circumference | bilateral | 995400\ C99-5100-137\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-137, ShearWave Elastography sample 4 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | Circ | meas | m_thy_EQI_5 | r_Circumference | bilateral | 995400\ C99-5100-138\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-138, ShearWave Elastography sample 5 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | Circ | meas | m_thy_EQI_6 | r_Circumference | bilateral | 995400\ C99-5100-139\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-139, ShearWave Elastography sample 6 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | Circ | meas | m_thy_EQI_7 | r_Circumference | bilateral | 995400\ C99-5100-140\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-140, ShearWave Elastography sample 7 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | Circ | meas | m_thy_EQI_8 | r_Circumference | bilateral | 995400\ C99-5100-141\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-141, ShearWave Elastography sample 8 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | Circ | meas | m_thy_EQI_9 | r_Circumference | bilateral | 995400\ C99-5100-142\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-142, ShearWave Elastography sample 9 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | Circ | meas | m_thy_EQI_10 | r_Circumference | bilateral | 995400\ C99-5100-143\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-143, ShearWave Elastography sample 10 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | Area | meas | m_thy_EQI_4 | r_Area | bilateral | 995400\ C99-5100-147\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-147, ShearWave Elastography sample 4 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | Area | meas | m_thy_EQI_5 | r_Area | bilateral | 995400\ C99-5100-148\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-148, ShearWave Elastography sample 5 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | Area | meas | m_thy_EQI_6 | r_Area | bilateral | 995400\ C99-5100-149\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-149, ShearWave Elastography sample 6 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | Area | meas | m_thy_EQI_7 | r_Area | bilateral | 995400\ C99-5100-150\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-150, ShearWave Elastography sample 7 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | Area | meas | m_thy_EQI_8 | r_Area | bilateral | 995400\ C99-5100-151\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-151, ShearWave Elastography sample 8 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | Area | meas | m_thy_EQI_9 | r_Area | bilateral | 995400\ C99-5100-152\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-152, ShearWave Elastography sample 9 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | Area | meas | m_thy_EQI_10 | r_Area | bilateral | 995400\ C99-5100-153\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-153, ShearWave Elastography sample 10 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | EQI Avg | meas | m_thy_EQI_1 | r_AvgStiffness_p | bilateral | 995400\C99-5100-001\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-001, ShearWave Elastography Stiffness sample 1 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | EQI Avg | meas | m_thy_EQI_2 | r_AvgStiffness_p | bilateral | 995400\C99-5100-002\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-002, ShearWave Elastography Stiffness sample 2 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | EQI Avg | meas | m_thy_EQI_3 | r_AvgStiffness_p | bilateral | 995400\C99-5100-003\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-003, ShearWave Elastography Stiffness sample 3 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | EQI Std | meas | m_thy_EQI_1 | r_StdStiffness_p | bilateral | 995400\C99-5100-044\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-044, ShearWave Elastography Stiffness sample 1 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | EQI Std | meas | m_thy_EQI_2 | r_StdStiffness_p | bilateral | 995400\C99-5100-045\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-045, ShearWave Elastography Stiffness sample 2 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | EQI Std | meas | m_thy_EQI_3 | r_StdStiffness_p | bilateral | 995400\C99-5100-046\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-046, ShearWave Elastography Stiffness sample 3 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | EQI Med | meas | m_thy_EQI_1 | r_MedStiffness_p | bilateral | 995400\C99-5100-064\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-064, ShearWave Elastography Stiffness sample 1 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | EQI Med | meas | m_thy_EQI_2 | r_MedStiffness_p | bilateral | 995400\C99-5100-065\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-065, ShearWave Elastography Stiffness sample 2 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | EQI Med | meas | m_thy_EQI_3 | r_MedStiffness_p | bilateral | 995400\C99-5100-066\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-066, ShearWave Elastography Stiffness sample 3 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | EQI IQR | meas | m_thy_EQI_1 | r_IqrStiffness_p | bilateral | 995400\C99-5100-084\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-084, ShearWave Elastography Stiffness sample 1 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | EQI IQR | meas | m_thy_EQI_2 | r_IqrStiffness_p | bilateral | 995400\C99-5100-085\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-085, ShearWave Elastography Stiffness sample 2 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | EQI IQR | meas | m_thy_EQI_3 | r_IqrStiffness_p | bilateral | 995400\C99-5100-086\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-086, ShearWave Elastography Stiffness sample 3 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | EQI IQR/Med | meas | m_thy_EQI_1 | r_IqrMedStiffness_p | bilateral | 995400\C99-5100-104\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-104, ShearWave Elasto Stiffness sample 1 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | EQI IQR/Med | meas | m_thy_EQI_2 | r_IqrMedStiffness_p | bilateral | 995400\C99-5100-105\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-105, ShearWave Elasto Stiffness sample 2 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | EQI IQR/Med | meas | m_thy_EQI_3 | r_IqrMedStiffness_p | bilateral | 995400\C99-5100-106\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-106, ShearWave Elasto Stiffness sample 3 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | Conf. Threshold | meas | m_thy_EQI_1 | r_CFStiffness | bilateral | 995400\C99-5100-124\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-124, ShearWave Elastography Stiffness sample 1 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | Conf. Threshold | meas | m_thy_EQI_2 | r_CFStiffness | bilateral | 995400\C99-5100-125\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-125, ShearWave Elastography Stiffness sample 2 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | Conf. Threshold | meas | m_thy_EQI_3 | r_CFStiffness | bilateral | 995400\C99-5100-126\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-126, ShearWave Elastography Stiffness sample 3 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | Circ | meas | m_thy_EQI_1 | r_Circumference | bilateral | 995400\C99-5100-134\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-134, ShearWave Elastography sample 1 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | Circ | meas | m_thy_EQI_2 | r_Circumference | bilateral | 995400\C99-5100-135\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-135, ShearWave Elastography sample 2 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | Circ | meas | m_thy_EQI_3 | r_Circumference | bilateral | 995400\C99-5100-136\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-136, ShearWave Elastography sample 3 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | Area | meas | m_thy_EQI_1 | r_Area | bilateral | 995400\C99-5100-144\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-144, ShearWave Elastography sample 1 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | Area | meas | m_thy_EQI_2 | r_Area | bilateral | 995400\C99-5100-145\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-145, ShearWave Elastography sample 2 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | Area | meas | m_thy_EQI_3 | r_Area | bilateral | 995400\C99-5100-146\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-146, ShearWave Elastography sample 3 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | EQI Min | meas | m_thy_EQI_1 | r_MinStiffness_p | bilateral | 995400\C99-5100-154\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-154, ShearWave Elastography Stiffness sample 1 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | EQI Min | meas | m_thy_EQI_2 | r_MinStiffness_p | bilateral | 995400\C99-5100-155\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-155, ShearWave Elastography Stiffness sample 2 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | EQI Min | meas | m_thy_EQI_3 | r_MinStiffness_p | bilateral | 995400\C99-5100-156\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-156, ShearWave Elastography Stiffness sample 3 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | EQI Min | meas | m_thy_EQI_4 | r_MinStiffness_p | bilateral | 995400\C99-5100-157\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-157, ShearWave Elastography Stiffness sample 4 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | EQI Min | meas | m_thy_EQI_5 | r_MinStiffness_p | bilateral | 995400\C99-5100-158\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-158, ShearWave Elastography Stiffness sample 5 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | EQI Min | meas | m_thy_EQI_6 | r_MinStiffness_p | bilateral | 995400\C99-5100-159\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-159, ShearWave Elastography Stiffness sample 6 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | EQI Min | meas | m_thy_EQI_7 | r_MinStiffness_p | bilateral | 995400\C99-5100-160\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-160, ShearWave Elastography Stiffness sample 7 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | EQI Min | meas | m_thy_EQI_8 | r_MinStiffness_p | bilateral | 995400\C99-5100-161\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-161, ShearWave Elastography Stiffness sample 8 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | EQI Min | meas | m_thy_EQI_9 | r_MinStiffness_p | bilateral | 995400\C99-5100-162\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-162, ShearWave Elastography Stiffness sample 9 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | EQI Min | meas | m_thy_EQI_10 | r_MinStiffness_p | bilateral | 995400\C99-5100-163\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-163, ShearWave Elastography Stiffness sample 10 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | EQI Min Vel | meas | m_thy_EQI_1 | r_MinStiffness_v | bilateral | 995400\C99-5100-164\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-164, ShearWave Elastography Velocity sample 1 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | EQI Min Vel | meas | m_thy_EQI_2 | r_MinStiffness_v | bilateral | 995400\C99-5100-165\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-165, ShearWave Elastography Velocity sample 2 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | EQI Min Vel | meas | m_thy_EQI_3 | r_MinStiffness_v | bilateral | 995400\C99-5100-166\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-166, ShearWave Elastography Velocity sample 3 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | EQI Min Vel | meas | m_thy_EQI_4 | r_MinStiffness_v | bilateral | 995400\C99-5100-167\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-167, ShearWave Elastography Velocity sample 4 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | EQI Min Vel | meas | m_thy_EQI_5 | r_MinStiffness_v | bilateral | 995400\C99-5100-168\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-168, ShearWave Elastography Velocity sample 5 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | EQI Min Vel | meas | m_thy_EQI_6 | r_MinStiffness_v | bilateral | 995400\C99-5100-169\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-169, ShearWave Elastography Velocity sample 6 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | EQI Min Vel | meas | m_thy_EQI_7 | r_MinStiffness_v | bilateral | 995400\C99-5100-170\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-170, ShearWave Elastography Velocity sample 7 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | EQI Min Vel | meas | m_thy_EQI_8 | r_MinStiffness_v | bilateral | 995400\C99-5100-171 \T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-171 , ShearWave Elastography Velocity sample 8 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | EQI Min Vel | meas | m_thy_EQI_9 | r_MinStiffness_v | bilateral | 995400\C99-5100-172\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-172, ShearWave Elastography Velocity sample 9 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | EQI Min Vel | meas | m_thy_EQI_10 | r_MinStiffness_v | bilateral | 995400\C99-5100-173\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-173, ShearWave Elastography Velocity sample 10 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | EQI Max | meas | m_thy_EQI_1 | r_MaxStiffness_p | bilateral | 995400\C99-5100-174\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-174, ShearWave Elastography Stiffness sample 1 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | EQI Max | meas | m_thy_EQI_2 | r_MaxStiffness_p | bilateral | 995400\C99-5100-175\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-175, ShearWave Elastography Stiffness sample 2 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | EQI Max | meas | m_thy_EQI_3 | r_MaxStiffness_p | bilateral | 995400\C99-5100-176\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-176, ShearWave Elastography Stiffness sample 3 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | EQI Max | meas | m_thy_EQI_4 | r_MaxStiffness_p | bilateral | 995400\C99-5100-177\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-177, ShearWave Elastography Stiffness sample 4 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | EQI Max | meas | m_thy_EQI_5 | r_MaxStiffness_p | bilateral | 995400\C99-5100-178\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-178, ShearWave Elastography Stiffness sample 5 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | EQI Max | meas | m_thy_EQI_6 | r_MaxStiffness_p | bilateral | 995400\C99-5100-179\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-179, ShearWave Elastography Stiffness sample 6 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | EQI Max | meas | m_thy_EQI_7 | r_MaxStiffness_p | bilateral | 995400\C99-5100-180\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-180, ShearWave Elastography Stiffness sample 7 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | EQI Max | meas | m_thy_EQI_8 | r_MaxStiffness_p | bilateral | 995400\C99-5100-181\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-181, ShearWave Elastography Stiffness sample 8 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | EQI Max | meas | m_thy_EQI_9 | r_MaxStiffness_p | bilateral | 995400\C99-5100-182\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-182, ShearWave Elastography Stiffness sample 9 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | EQI Max | meas | m_thy_EQI_10 | r_MaxStiffness_p | bilateral | 995400\C99-5100-183\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-183, ShearWave Elastography Stiffness sample 10 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 1 | EQI Max Vel | meas | m_thy_EQI_1 | r_MaxStiffness_v | bilateral | 995400\C99-5100-184\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-184, ShearWave Elastography Velocity sample 1 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 2 | EQI Max Vel | meas | m_thy_EQI_2 | r_MaxStiffness_v | bilateral | 995400\C99-5100-185\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-185, ShearWave Elastography Velocity sample 2 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 3 | EQI Max Vel | meas | m_thy_EQI_3 | r_MaxStiffness_v | bilateral | 995400\C99-5100-186\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-186, ShearWave Elastography Velocity sample 3 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 4 | EQI Max Vel | meas | m_thy_EQI_4 | r_MaxStiffness_v | bilateral | 995400\C99-5100-187\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-187, ShearWave Elastography Velocity sample 4 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 5 | EQI Max Vel | meas | m_thy_EQI_5 | r_MaxStiffness_v | bilateral | 995400\C99-5100-188\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-188, ShearWave Elastography Velocity sample 5 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 6 | EQI Max Vel | meas | m_thy_EQI_6 | r_MaxStiffness_v | bilateral | 995400\C99-5100-189\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-189, ShearWave Elastography Velocity sample 6 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 7 | EQI Max Vel | meas | m_thy_EQI_7 | r_MaxStiffness_v | bilateral | 995400\C99-5100-190\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-190, ShearWave Elastography Velocity sample 7 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 8 | EQI Max Vel | meas | m_thy_EQI_8 | r_MaxStiffness_v | bilateral | 995400\C99-5100-191 \T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-191 , ShearWave Elastography Velocity sample 8 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 9 | EQI Max Vel | meas | m_thy_EQI_9 | r_MaxStiffness_v | bilateral | 995400\C99-5100-192\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-192, ShearWave Elastography Velocity sample 9 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI 10 | EQI Max Vel | meas | m_thy_EQI_10 | r_MaxStiffness_v | bilateral | 995400\C99-5100-193\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-193, ShearWave Elastography Velocity sample 10 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Breast | Lesion 1 L | Dist | meas | m_brleslen1 | r_Distance | bilateral | 995400\C7470-30\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-30, Lesion Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 2 L | Dist | meas | m_brleslen2 | r_Distance | bilateral | 995400\C7470-30\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-30, Lesion Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 3 L | Dist | meas | m_brleslen3 | r_Distance | bilateral | 995400\C7470-30\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-30, Lesion Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 4 L | Dist | meas | m_brleslen4 | r_Distance | bilateral | 995400\C7470-30\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-30, Lesion Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 5 L | Dist | meas | m_brleslen5 | r_Distance | bilateral | 995400\C7470-30\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-30, Lesion Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 6 L | Dist | meas | m_brleslen6 | r_Distance | bilateral | 995400\C7470-30\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-30, Lesion Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 7 L | Dist | meas | m_brleslen7 | r_Distance | bilateral | 995400\C7470-30\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-30, Lesion Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 8 L | Dist | meas | m_brleslen8 | r_Distance | bilateral | 995400\C7470-30\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-30, Lesion Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 1 W | Dist | meas | m_brlesw1 | r_Distance | bilateral | 995400\C7470-31\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-31, Lesion Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 2 W | Dist | meas | m_brlesw2 | r_Distance | bilateral | 995400\C7470-31\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-31, Lesion Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 3 W | Dist | meas | m_brlesw3 | r_Distance | bilateral | 995400\C7470-31\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-31, Lesion Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 4 W | Dist | meas | m_brlesw4 | r_Distance | bilateral | 995400\C7470-31\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-31, Lesion Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 5 W | Dist | meas | m_brlesw5 | r_Distance | bilateral | 995400\C7470-31\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-31, Lesion Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 6 W | Dist | meas | m_brlesw6 | r_Distance | bilateral | 995400\C7470-31\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-31, Lesion Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 7 W | Dist | meas | m_brlesw7 | r_Distance | bilateral | 995400\C7470-31\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-31, Lesion Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 8 W | Dist | meas | m_brlesw8 | r_Distance | bilateral | 995400\C7470-31\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-31, Lesion Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 1 H | Dist | meas | m_brlesht1 | r_Distance | bilateral | 995400\C7470-32\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-32, Lesion Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 2 H | Dist | meas | m_brlesht2 | r_Distance | bilateral | 995400\C7470-32\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-32, Lesion Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 3 H | Dist | meas | m_brlesht3 | r_Distance | bilateral | 995400\C7470-32\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-32, Lesion Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 4 H | Dist | meas | m_brlesht4 | r_Distance | bilateral | 995400\C7470-32\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-32, Lesion Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 5 H | Dist | meas | m_brlesht5 | r_Distance | bilateral | 995400\C7470-32\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-32, Lesion Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 6 H | Dist | meas | m_brlesht6 | r_Distance | bilateral | 995400\C7470-32\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-32, Lesion Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 7 H | Dist | meas | m_brlesht7 | r_Distance | bilateral | 995400\C7470-32\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-32, Lesion Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion 8 H | Dist | meas | m_brlesht8 | r_Distance | bilateral | 995400\C7470-32\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7470-32, Lesion Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Prostate H | Dist | meas | m_prostateht | r_Distance | unilateral | 995400\121207\G-03A2\T-9200B | SRT, T-9200B, Prostate |  | DCM, 121207, Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Prostate W | Dist | meas | m_prostatewidth | r_Distance | unilateral | 995400\G-03A2\G-A220\T-9200B | SRT, T-9200B, Prostate |  | SRT, G-A220, Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Prostate L | Dist | meas | m_prostatelen | r_Distance | unilateral | 995400\G-03A2\G-A22A\T-9200B | SRT, T-9200B, Prostate |  | SRT, G-A22A, Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Testicle H1 | Dist | meas | m_tesht1 | r_Distance | bilateral | 995400\121207\G-03A2\T-94000 | SRT, T-94000, Testis |  | DCM, 121207, Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Testicle W1 | Dist | meas | m_tesWth1 | r_Distance | bilateral | 995400\G-03A2\G-A220\T-94000 | SRT, T-94000, Testis |  | SRT, G-A220, Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Testicle L1 | Dist | meas | m_teslen1 | r_Distance | bilateral | 995400\G-03A2\G-A22A\T-94000 | SRT, T-94000, Testis |  | SRT, G-A22A, Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule L1 | Dist | meas | m_thynoduleL1 | r_Distance | bilateral | 995400\C7470-33\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-33, Nodule Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule L2 | Dist | meas | m_thynoduleL2 | r_Distance | bilateral | 995400\C7470-33\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-33, Nodule Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule L3 | Dist | meas | m_thynoduleL3 | r_Distance | bilateral | 995400\C7470-33\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-33, Nodule Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule L4 | Dist | meas | m_thynoduleL4 | r_Distance | bilateral | 995400\C7470-33\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-33, Nodule Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule L5 | Dist | meas | m_thynoduleL5 | r_Distance | bilateral | 995400\C7470-33\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-33, Nodule Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule L6 | Dist | meas | m_thynoduleL6 | r_Distance | bilateral | 995400\C7470-33\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-33, Nodule Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule L7 | Dist | meas | m_thynoduleL7 | r_Distance | bilateral | 995400\C7470-33\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-33, Nodule Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule L8 | Dist | meas | m_thynoduleL8 | r_Distance | bilateral | 995400\C7470-33\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-33, Nodule Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule W1 | Dist | meas | m_thynoduleW1 | r_Distance | bilateral | 995400\C7470-34\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-34, Nodule Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule W2 | Dist | meas | m_thynoduleW2 | r_Distance | bilateral | 995400\C7470-34\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-34, Nodule Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule W3 | Dist | meas | m_thynoduleW3 | r_Distance | bilateral | 995400\C7470-34\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-34, Nodule Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule W4 | Dist | meas | m_thynoduleW4 | r_Distance | bilateral | 995400\C7470-34\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-34, Nodule Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule W5 | Dist | meas | m_thynoduleW5 | r_Distance | bilateral | 995400\C7470-34\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-34, Nodule Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule W6 | Dist | meas | m_thynoduleW6 | r_Distance | bilateral | 995400\C7470-34\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-34, Nodule Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule W7 | Dist | meas | m_thynoduleW7 | r_Distance | bilateral | 995400\C7470-34\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-34, Nodule Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule W8 | Dist | meas | m_thynoduleW8 | r_Distance | bilateral | 995400\C7470-34\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-34, Nodule Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule H1 | Dist | meas | m_thynoduleH1 | r_Distance | bilateral | 995400\C7470-35\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-35, Nodule Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule H2 | Dist | meas | m_thynoduleH2 | r_Distance | bilateral | 995400\C7470-35\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-35, Nodule Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule H3 | Dist | meas | m_thynoduleH3 | r_Distance | bilateral | 995400\C7470-35\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-35, Nodule Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule H4 | Dist | meas | m_thynoduleH4 | r_Distance | bilateral | 995400\C7470-35\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-35, Nodule Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule H5 | Dist | meas | m_thynoduleH5 | r_Distance | bilateral | 995400\C7470-35\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-35, Nodule Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule H6 | Dist | meas | m_thynoduleH6 | r_Distance | bilateral | 995400\C7470-35\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-35, Nodule Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule H7 | Dist | meas | m_thynoduleH7 | r_Distance | bilateral | 995400\C7470-35\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-35, Nodule Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule H8 | Dist | meas | m_thynoduleH8 | r_Distance | bilateral | 995400\C7470-35\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7470-35, Nodule Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid Height | Dist | meas | m_thyroidht | r_Distance | bilateral | 995400\121207\G-03A2\T-B6000 | SRT, T-B6000, Thyroid |  | DCM, 121207, Height |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid Width | Dist | meas | m_thyroidwidth | r_Distance | bilateral | 995400\G-03A2\G-A220\T-B6000 | SRT, T-B6000, Thyroid |  | SRT, G-A220, Width |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid Length | Dist | meas | m_thyroidlen | r_Distance | bilateral | 995400\G-03A2\G-A22A\T-B6000 | SRT, T-B6000, Thyroid |  | SRT, G-A22A, Length |  | UCUM, mm, mm |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI Avg | none | calc | c_brles_EQI_Avg_p | none | bilateral | 995400\ C99-5100-021 \C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-021 , Average of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI Std | none | calc | c_brles_EQI_Std_p | none | bilateral | 995400\ C99-5100-022\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-022, Standard Deviation of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI Med | none | calc | c_brles_EQI_Med_p | none | bilateral | 995400\ C99-5100-023\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-023, Median of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI IQR | none | calc | c_brles_EQI_Iqr_p | none | bilateral | 995400\ C99-5100-024\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-024, Interquartile Range of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI IQR/Med | none | calc | c_brles_EQI_Iqrmed_p | none | bilateral | 995400\ C99-5100-025\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-025, InterquartileRange over Median of ShearWave Elastography samples |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI Avg Vel | none | calc | c_brles_EQI_Avg_v | none | bilateral | 995400\ C99-5100-026\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-026, Average of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI Std Vel | none | calc | c_brles_EQI_Std_v | none | bilateral | 995400\ C99-5100-027 \C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-027 , Standard Deviation of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI Med Vel | none | calc | c_brles_EQI_Med_v | none | bilateral | 995400\ C99-5100-028\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-028, Median of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI IQR Vel | none | calc | c_brles_EQI_Iqr_v | none | bilateral | 995400\ C99-5100-029\C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-029, Interquartile Range of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Breast | Bre Les EQI IQR/Med Vel | none | calc | c_brles_EQI_Iqrmed_v | none | bilateral | 995400\ C99-5100-030 \C4031-04\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-030 , Interquartile Range over Median of ShearWave Velocity samples |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-04, Breast Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI Avg | none | calc | c_thyles_EQI_Avg_p | none | bilateral | 995400\ C99-5100-021 \C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-021 , Average of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI Std | none | calc | c_thyles_EQI_Std_p | none | bilateral | 995400\ C99-5100-022\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-022, Standard Deviation of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI Med | none | calc | c_thyles_EQI_Med_p | none | bilateral | 995400\ C99-5100-023\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-023, Median of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI IQR | none | calc | c_thyles_EQI_Iqr_p | none | bilateral | 995400\ C99-5100-024\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-024, Interquartile Range of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI IQR/Med | none | calc | c_thyles_EQI_IqrM_p | none | bilateral | 995400\ C99-5100-025\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-025, InterquartileRange over Median of ShearWave Elastography samples |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI Avg Vel | none | calc | c_thyles_EQI_Avg_v | none | bilateral | 995400\ C99-5100-026\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-026, Average of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI Std Vel | none | calc | c_thyles_EQI_Std_v | none | bilateral | 995400\ C99-5100-027 \C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-027 , Standard Deviation of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI Med Vel | none | calc | c_thyles_EQI_Med_v | none | bilateral | 995400\ C99-5100-028\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-028, Median of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI IQR Vel | none | calc | c_thyles_EQI_Iqr_v | none | bilateral | 995400\ C99-5100-029\C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-029, Interquartile Range of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Small Parts | Thy Les EQI IQR/Med Vel | none | calc | c_thyles_EQI_IqrM_v | none | bilateral | 995400\ C99-5100-030 \C4031-05\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-030 , Interquartile Range over Median of ShearWave Velocity samples |  | UCUM, %, Percent | [7 empty] | 99PMSBLUS, C4031-05, Thyroid Lesion
Small Parts Ultrasound Procedure Report | Breast | Breast EQI Avg | none | calc | c_br_EQI_Avg_p | none | bilateral | 995400\ C99-5100-021 \T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-021 , Average of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI Std | none | calc | c_br_EQI_Std_p | none | bilateral | 995400\ C99-5100-022\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-022, Standard Deviation of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI Med | none | calc | c_br_EQI_Med_p | none | bilateral | 995400\ C99-5100-023\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-023, Median of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI IQR | none | calc | c_br_EQI_Iqr_p | none | bilateral | 995400\ C99-5100-024\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-024, Interquartile Range of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI IQR/Med | none | calc | c_br_EQI_Iqrmed_p | none | bilateral | 995400\ C99-5100-025\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-025, InterquartileRange over Median of ShearWave Elastography samples |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI Avg Vel | none | calc | c_br_EQI_Avg_v | none | bilateral | 995400\ C99-5100-026\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-026, Average of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI Std Vel | none | calc | c_br_EQI_Std_v | none | bilateral | 995400\ C99-5100-027 \T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-027 , Standard Deviation of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI Med Vel | none | calc | c_br_EQI_Med_v | none | bilateral | 995400\ C99-5100-028\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-028, Median of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI IQR Vel | none | calc | c_br_EQI_Iqr_v | none | bilateral | 995400\ C99-5100-029\T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-029, Interquartile Range of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Breast | Breast EQI IQR/Med Vel | none | calc | c_br_EQI_Iqrmed_v | none | bilateral | 995400\ C99-5100-030 \T-04000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-030 , Interquartile Range over Median of ShearWave Velocity samples |  | UCUM, %, Percent | [7 empty] | SRT, T-04000, Breast
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI Avg | none | calc | c_thy_EQI_Avg_p | none | bilateral | 995400\ C99-5100-021 \T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-021 , Average of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI Std | none | calc | c_thy_EQI_Std_p | none | bilateral | 995400\ C99-5100-022\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-022, Standard Deviation of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI Med | none | calc | c_thy_EQI_Med_p | none | bilateral | 995400\ C99-5100-023\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-023, Median of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI IQR | none | calc | c_thy_EQI_Iqr_p | none | bilateral | 995400\ C99-5100-024\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-024, Interquartile Range of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI IQR/Med | none | calc | c_thy_EQI_Iqrmed_p | none | bilateral | 995400\ C99-5100-025\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-025, InterquartileRange over Median of ShearWave Elastography samples |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI Avg Vel | none | calc | c_thy_EQI_Avg_v | none | bilateral | 995400\ C99-5100-026\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-026, Average of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI Std Vel | none | calc | c_thy_EQI_Std_v | none | bilateral | 995400\ C99-5100-027 \T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-027 , Standard Deviation of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI Med Vel | none | calc | c_thy_EQI_Med_v | none | bilateral | 995400\ C99-5100-028\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-028, Median of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI IQR Vel | none | calc | c_thy_EQI_Iqr_v | none | bilateral | 995400\ C99-5100-029\T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-029, Interquartile Range of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid EQI IQR/Med Vel | none | calc | c_thy_EQI_Iqrmed_v | none | bilateral | 995400\ C99-5100-030 \T5100-01\T-B6000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-030 , Interquartile Range over Median of ShearWave Velocity samples |  | UCUM, %, Percent | [7 empty] | SRT, T-B6000, Thyroid
Small Parts Ultrasound Procedure Report | Breast | Lesion1V | none | calc | c_brlesvol1 | none | bilateral | 995400\C7472-03\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7472-03, Lesion Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion2V | none | calc | c_brlesvol2 | none | bilateral | 995400\C7472-03\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7472-03, Lesion Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion3V | none | calc | c_brlesvol3 | none | bilateral | 995400\C7472-03\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7472-03, Lesion Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion4V | none | calc | c_brlesvol4 | none | bilateral | 995400\C7472-03\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7472-03, Lesion Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion5V | none | calc | c_brlesvol5 | none | bilateral | 995400\C7472-03\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7472-03, Lesion Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion6V | none | calc | c_brlesvol6 | none | bilateral | 995400\C7472-03\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7472-03, Lesion Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion7V | none | calc | c_brlesvol7 | none | bilateral | 995400\C7472-03\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7472-03, Lesion Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Breast | Lesion8V | none | calc | c_brlesvol8 | none | bilateral | 995400\C7472-03\G-03A2\T-04000\xxx | SRT, T-04000, Breast |  | 99PMSBLUS, C7472-03, Lesion Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | PSA Density | none | calc | c_prostatePsaDensity | none | unilateral | 995400\C995400-1-02\G-03A2\T-9200B | SRT, T-9200B, Prostate |  | 99PMSBLUS, C995400-1-02, Prostate specific antigen density |  | 99PMSBLUS, ng/mL/cm3, ng/mL/cm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Prostate Volume | none | calc | c_prostateVol | none | unilateral | 995400\G-03A2\G-D705\T-9200B | SRT, T-9200B, Prostate |  | SRT, G-D705, Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Testicle Vol1 | none | calc | c_tesVol1 | none | bilateral | 995400\G-03A2\G-D705\T-94000 | SRT, T-94000, Testis |  | SRT, G-D705, Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule Vol1 | none | calc | c_thynoduledvol1 | none | bilateral | 995400\C7472-04\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7472-04, Nodule Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule Vol2 | none | calc | c_thynoduledvol2 | none | bilateral | 995400\C7472-04\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7472-04, Nodule Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule Vol3 | none | calc | c_thynoduledvol3 | none | bilateral | 995400\C7472-04\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7472-04, Nodule Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule Vol4 | none | calc | c_thynoduledvol4 | none | bilateral | 995400\C7472-04\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7472-04, Nodule Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule Vol5 | none | calc | c_thynoduledvol5 | none | bilateral | 995400\C7472-04\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7472-04, Nodule Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule Vol6 | none | calc | c_thynoduledvol6 | none | bilateral | 995400\C7472-04\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7472-04, Nodule Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule Vol7 | none | calc | c_thynoduledvol7 | none | bilateral | 995400\C7472-04\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7472-04, Nodule Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thy Nodule Vol8 | none | calc | c_thynoduledvol8 | none | bilateral | 995400\C7472-04\G-03A2\T-B6000\xxx | SRT, T-B6000, Thyroid |  | 99PMSBLUS, C7472-04, Nodule Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Small Parts Ultrasound Procedure Report | Small Parts | Thyroid Volume | none | calc | c_thyroidvol | none | bilateral | 995400\G-03A2\G-D705\T-B6000 | SRT, T-B6000, Thyroid |  | SRT, G-D705, Volume |  | UCUM, mm3, mm3 |  | SRT, G-03A2, 2D mode
Vascular Ultrasound Procedure Report | Abdominal | Liver Vasc Angle | Angle | meas | m_LiverVascAngle | r_Angle | unilateral | 125100\C12221-05\T5100-01\T-62000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C12221-05, Liver Vasc Angle |  | UCUM, deg, deg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | EQI Avg | meas | m_kidney_EQI_4 | r_AvgStiffness_p | bilateral | 125100\ C99-5100-004\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-004, ShearWave Elastography Stiffness sample 4 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | EQI Avg | meas | m_kidney_EQI_5 | r_AvgStiffness_p | bilateral | 125100\ C99-5100-005\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-005, ShearWave Elastography Stiffness sample 5 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | EQI Avg | meas | m_kidney_EQI_6 | r_AvgStiffness_p | bilateral | 125100\ C99-5100-006\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-006, ShearWave Elastography Stiffness sample 6 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | EQI Avg | meas | m_kidney_EQI_7 | r_AvgStiffness_p | bilateral | 125100\ C99-5100-007\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-007, ShearWave Elastography Stiffness sample 7 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | EQI Avg | meas | m_kidney_EQI_8 | r_AvgStiffness_p | bilateral | 125100\ C99-5100-008\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-008, ShearWave Elastography Stiffness sample 8 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | EQI Avg | meas | m_kidney_EQI_9 | r_AvgStiffness_p | bilateral | 125100\ C99-5100-009\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-009, ShearWave Elastography Stiffness sample 9 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | EQI Avg | meas | m_kidney_EQI_10 | r_AvgStiffness_p | bilateral | 125100\ C99-5100-010\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-010, ShearWave Elastography Stiffness sample 10 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | EQI Avg Vel | meas | m_kidney_EQI_1 | r_AvgStiffness_v | bilateral | 125100\ C99-5100-011\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-011, ShearWave Elastography Velocity sample 1 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | EQI Avg Vel | meas | m_kidney_EQI_2 | r_AvgStiffness_v | bilateral | 125100\ C99-5100-012\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-012, ShearWave Elastography Velocity sample 2 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | EQI Avg Vel | meas | m_kidney_EQI_3 | r_AvgStiffness_v | bilateral | 125100\ C99-5100-013\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-013, ShearWave Elastography Velocity sample 3 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | EQI Avg Vel | meas | m_kidney_EQI_4 | r_AvgStiffness_v | bilateral | 125100\ C99-5100-014\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-014, ShearWave Elastography Velocity sample 4 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | EQI Avg Vel | meas | m_kidney_EQI_5 | r_AvgStiffness_v | bilateral | 125100\ C99-5100-015\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-015, ShearWave Elastography Velocity sample 5 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | EQI Avg Vel | meas | m_kidney_EQI_6 | r_AvgStiffness_v | bilateral | 125100\ C99-5100-016\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-016, ShearWave Elastography Velocity sample 6 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | EQI Avg Vel | meas | m_kidney_EQI_7 | r_AvgStiffness_v | bilateral | 125100\ C99-5100-017\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-017, ShearWave Elastography Velocity sample 7 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | EQI Avg Vel | meas | m_kidney_EQI_8 | r_AvgStiffness_v | bilateral | 125100\ C99-5100-018 \T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-018 , ShearWave Elastography Velocity sample 8 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | EQI Avg Vel | meas | m_kidney_EQI_9 | r_AvgStiffness_v | bilateral | 125100\ C99-5100-019\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-019, ShearWave Elastography Velocity sample 9 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | EQI Avg Vel | meas | m_kidney_EQI_10 | r_AvgStiffness_v | bilateral | 125100\ C99-5100-020\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-020, ShearWave Elastography Velocity sample 10 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | EQI Std | meas | m_kidney_EQI_4 | r_StdStiffness_p | bilateral | 125100\ C99-5100-047\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-047, ShearWave Elastography Stiffness sample 4 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | EQI Std | meas | m_kidney_EQI_5 | r_StdStiffness_p | bilateral | 125100\ C99-5100-048\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-048, ShearWave Elastography Stiffness sample 5 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | EQI Std | meas | m_kidney_EQI_6 | r_StdStiffness_p | bilateral | 125100\ C99-5100-049\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-049, ShearWave Elastography Stiffness sample 6 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | EQI Std | meas | m_kidney_EQI_7 | r_StdStiffness_p | bilateral | 125100\ C99-5100-050\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-050, ShearWave Elastography Stiffness sample 7 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | EQI Std | meas | m_kidney_EQI_8 | r_StdStiffness_p | bilateral | 125100\ C99-5100-051\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-051, ShearWave Elastography Stiffness sample 8 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | EQI Std | meas | m_kidney_EQI_9 | r_StdStiffness_p | bilateral | 125100\ C99-5100-052\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-052, ShearWave Elastography Stiffness sample 9 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | EQI Std | meas | m_kidney_EQI_10 | r_StdStiffness_p | bilateral | 125100\ C99-5100-053\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-053, ShearWave Elastography Stiffness sample 10 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | EQI Std Vel | meas | m_kidney_EQI_1 | r_StdStiffness_v | bilateral | 125100\ C99-5100-054\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-054, ShearWave Elastography Velocity sample 1 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | EQI Std Vel | meas | m_kidney_EQI_2 | r_StdStiffness_v | bilateral | 125100\ C99-5100-055\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-055, ShearWave Elastography Velocity sample 2 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | EQI Std Vel | meas | m_kidney_EQI_3 | r_StdStiffness_v | bilateral | 125100\ C99-5100-056\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-056, ShearWave Elastography Velocity sample 3 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | EQI Std Vel | meas | m_kidney_EQI_4 | r_StdStiffness_v | bilateral | 125100\ C99-5100-057\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-057, ShearWave Elastography Velocity sample 4 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | EQI Std Vel | meas | m_kidney_EQI_5 | r_StdStiffness_v | bilateral | 125100\ C99-5100-058\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-058, ShearWave Elastography Velocity sample 5 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | EQI Std Vel | meas | m_kidney_EQI_6 | r_StdStiffness_v | bilateral | 125100\ C99-5100-059\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-059, ShearWave Elastography Velocity sample 6 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | EQI Std Vel | meas | m_kidney_EQI_7 | r_StdStiffness_v | bilateral | 125100\ C99-5100-060\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-060, ShearWave Elastography Velocity sample 7 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | EQI Std Vel | meas | m_kidney_EQI_8 | r_StdStiffness_v | bilateral | 125100\ C99-5100-061\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-061, ShearWave Elastography Velocity sample 8 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | EQI Std Vel | meas | m_kidney_EQI_9 | r_StdStiffness_v | bilateral | 125100\ C99-5100-062\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-062, ShearWave Elastography Velocity sample 9 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | EQI Std Vel | meas | m_kidney_EQI_10 | r_StdStiffness_v | bilateral | 125100\ C99-5100-063\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-063, ShearWave Elastography Velocity sample 10 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | EQI Med | meas | m_kidney_EQI_4 | r_MedStiffness_p | bilateral | 125100\ C99-5100-067\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-067, ShearWave Elastography Stiffness sample 4 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | EQI Med | meas | m_kidney_EQI_5 | r_MedStiffness_p | bilateral | 125100\ C99-5100-068\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-068, ShearWave Elastography Stiffness sample 5 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | EQI Med | meas | m_kidney_EQI_6 | r_MedStiffness_p | bilateral | 125100\ C99-5100-069\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-069, ShearWave Elastography Stiffness sample 6 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | EQI Med | meas | m_kidney_EQI_7 | r_MedStiffness_p | bilateral | 125100\ C99-5100-070\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-070, ShearWave Elastography Stiffness sample 7 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | EQI Med | meas | m_kidney_EQI_8 | r_MedStiffness_p | bilateral | 125100\ C99-5100-071\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-071, ShearWave Elastography Stiffness sample 8 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | EQI Med | meas | m_kidney_EQI_9 | r_MedStiffness_p | bilateral | 125100\ C99-5100-072\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-072, ShearWave Elastography Stiffness sample 9 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | EQI Med | meas | m_kidney_EQI_10 | r_MedStiffness_p | bilateral | 125100\ C99-5100-073\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-073, ShearWave Elastography Stiffness sample 10 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | EQI Med Vel | meas | m_kidney_EQI_1 | r_MedStiffness_v | bilateral | 125100\ C99-5100-074\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-074, ShearWave Elastography Velocity sample 1 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | EQI Med Vel | meas | m_kidney_EQI_2 | r_MedStiffness_v | bilateral | 125100\ C99-5100-075\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-075, ShearWave Elastography Velocity sample 2 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | EQI Med Vel | meas | m_kidney_EQI_3 | r_MedStiffness_v | bilateral | 125100\ C99-5100-076\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-076, ShearWave Elastography Velocity sample 3 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | EQI Med Vel | meas | m_kidney_EQI_4 | r_MedStiffness_v | bilateral | 125100\ C99-5100-077\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-077, ShearWave Elastography Velocity sample 4 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | EQI Med Vel | meas | m_kidney_EQI_5 | r_MedStiffness_v | bilateral | 125100\ C99-5100-078\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-078, ShearWave Elastography Velocity sample 5 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | EQI Med Vel | meas | m_kidney_EQI_6 | r_MedStiffness_v | bilateral | 125100\ C99-5100-079\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-079, ShearWave Elastography Velocity sample 6 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | EQI Med Vel | meas | m_kidney_EQI_7 | r_MedStiffness_v | bilateral | 125100\ C99-5100-080\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-080, ShearWave Elastography Velocity sample 7 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | EQI Med Vel | meas | m_kidney_EQI_8 | r_MedStiffness_v | bilateral | 125100\ C99-5100-081\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-081, ShearWave Elastography Velocity sample 8 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | EQI Med Vel | meas | m_kidney_EQI_9 | r_MedStiffness_v | bilateral | 125100\ C99-5100-082\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-082, ShearWave Elastography Velocity sample 9 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | EQI Med Vel | meas | m_kidney_EQI_10 | r_MedStiffness_v | bilateral | 125100\ C99-5100-083\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-083, ShearWave Elastography Velocity sample 10 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | EQI IQR | meas | m_kidney_EQI_4 | r_IqrStiffness_p | bilateral | 125100\ C99-5100-087\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-087, ShearWave Elastography Stiffness sample 4 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | EQI IQR | meas | m_kidney_EQI_5 | r_IqrStiffness_p | bilateral | 125100\ C99-5100-088\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-088, ShearWave Elastography Stiffness sample 5 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | EQI IQR | meas | m_kidney_EQI_6 | r_IqrStiffness_p | bilateral | 125100\ C99-5100-089\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-089, ShearWave Elastography Stiffness sample 6 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | EQI IQR | meas | m_kidney_EQI_7 | r_IqrStiffness_p | bilateral | 125100\ C99-5100-090\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-090, ShearWave Elastography Stiffness sample 7 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | EQI IQR | meas | m_kidney_EQI_8 | r_IqrStiffness_p | bilateral | 125100\ C99-5100-091\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-091, ShearWave Elastography Stiffness sample 8 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | EQI IQR | meas | m_kidney_EQI_9 | r_IqrStiffness_p | bilateral | 125100\ C99-5100-092\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-092, ShearWave Elastography Stiffness sample 9 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | EQI IQR | meas | m_kidney_EQI_10 | r_IqrStiffness_p | bilateral | 125100\ C99-5100-093\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-093, ShearWave Elastography Stiffness sample 10 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | EQI IQR Vel | meas | m_kidney_EQI_1 | r_IqrStiffness_v | bilateral | 125100\ C99-5100-094\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-094, ShearWave Elastography Velocity sample 1 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | EQI IQR Vel | meas | m_kidney_EQI_2 | r_IqrStiffness_v | bilateral | 125100\ C99-5100-095\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-095, ShearWave Elastography Velocity sample 2 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | EQI IQR Vel | meas | m_kidney_EQI_3 | r_IqrStiffness_v | bilateral | 125100\ C99-5100-096\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-096, ShearWave Elastography Velocity sample 3 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | EQI IQR Vel | meas | m_kidney_EQI_4 | r_IqrStiffness_v | bilateral | 125100\ C99-5100-097\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-097, ShearWave Elastography Velocity sample 4 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | EQI IQR Vel | meas | m_kidney_EQI_5 | r_IqrStiffness_v | bilateral | 125100\ C99-5100-098\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-098, ShearWave Elastography Velocity sample 5 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | EQI IQR Vel | meas | m_kidney_EQI_6 | r_IqrStiffness_v | bilateral | 125100\ C99-5100-099\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-099, ShearWave Elastography Velocity sample 6 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | EQI IQR Vel | meas | m_kidney_EQI_7 | r_IqrStiffness_v | bilateral | 125100\ C99-5100-100\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-100, ShearWave Elastography Velocity sample 7 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | EQI IQR Vel | meas | m_kidney_EQI_8 | r_IqrStiffness_v | bilateral | 125100\ C99-5100-101\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-101, ShearWave Elastography Velocity sample 8 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | EQI IQR Vel | meas | m_kidney_EQI_9 | r_IqrStiffness_v | bilateral | 125100\ C99-5100-102\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-102, ShearWave Elastography Velocity sample 9 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | EQI IQR Vel | meas | m_kidney_EQI_10 | r_IqrStiffness_v | bilateral | 125100\ C99-5100-103\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-103, ShearWave Elastography Velocity sample 10 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | EQI IQR/Med | meas | m_kidney_EQI_4 | r_IqrMedStiffness_p | bilateral | 125100\ C99-5100-107\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-107, ShearWave Elasto Stiffness sample 4 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | EQI IQR/Med | meas | m_kidney_EQI_5 | r_IqrMedStiffness_p | bilateral | 125100\ C99-5100-108\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-108, ShearWave Elasto Stiffness sample 5 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | EQI IQR/Med | meas | m_kidney_EQI_6 | r_IqrMedStiffness_p | bilateral | 125100\ C99-5100-109\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-109, ShearWave Elasto Stiffness sample 6 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | EQI IQR/Med | meas | m_kidney_EQI_7 | r_IqrMedStiffness_p | bilateral | 125100\ C99-5100-110\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-110, ShearWave Elasto Stiffness sample 7 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | EQI IQR/Med | meas | m_kidney_EQI_8 | r_IqrMedStiffness_p | bilateral | 125100\ C99-5100-111\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-111, ShearWave Elasto Stiffness sample 8 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | EQI IQR/Med | meas | m_kidney_EQI_9 | r_IqrMedStiffness_p | bilateral | 125100\ C99-5100-112\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-112, ShearWave Elasto Stiffness sample 9 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | EQI IQR/Med | meas | m_kidney_EQI_10 | r_IqrMedStiffness_p | bilateral | 125100\ C99-5100-113\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-113, ShearWave Elasto Stiffness sample 10 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | EQI IQR/Med Vel | meas | m_kidney_EQI_1 | r_IqrMedStiffness_v | bilateral | 125100\ C99-5100-114\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-114, ShearWave Elasto Velocity sample 1 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | EQI IQR/Med Vel | meas | m_kidney_EQI_2 | r_IqrMedStiffness_v | bilateral | 125100\ C99-5100-115\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-115, ShearWave Elasto Velocity sample 2 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | EQI IQR/Med Vel | meas | m_kidney_EQI_3 | r_IqrMedStiffness_v | bilateral | 125100\ C99-5100-116\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-116, ShearWave Elasto Velocity sample 3 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | EQI IQR/Med Vel | meas | m_kidney_EQI_4 | r_IqrMedStiffness_v | bilateral | 125100\ C99-5100-117\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-117, ShearWave Elasto Velocity sample 4 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | EQI IQR/Med Vel | meas | m_kidney_EQI_5 | r_IqrMedStiffness_v | bilateral | 125100\ C99-5100-118\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-118, ShearWave Elasto Velocity sample 5 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | EQI IQR/Med Vel | meas | m_kidney_EQI_6 | r_IqrMedStiffness_v | bilateral | 125100\ C99-5100-119\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-119, ShearWave Elasto Velocity sample 6 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | EQI IQR/Med Vel | meas | m_kidney_EQI_7 | r_IqrMedStiffness_v | bilateral | 125100\ C99-5100-120\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-120, ShearWave Elasto Velocity sample 7 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | EQI IQR/Med Vel | meas | m_kidney_EQI_8 | r_IqrMedStiffness_v | bilateral | 125100\ C99-5100-121\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-121, ShearWave Elasto Velocity sample 8 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | EQI IQR/Med Vel | meas | m_kidney_EQI_9 | r_IqrMedStiffness_v | bilateral | 125100\ C99-5100-122\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-122, ShearWave Elasto Velocity sample 9 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | EQI IQR/Med Vel | meas | m_kidney_EQI_10 | r_IqrMedStiffness_v | bilateral | 125100\ C99-5100-123\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-123, ShearWave Elasto Velocity sample 10 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | Conf. Threshold | meas | m_kidney_EQI_4 | r_CFStiffness | bilateral | 125100\ C99-5100-127\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-127, ShearWave Elastography Stiffness sample 4 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | Conf. Threshold | meas | m_kidney_EQI_5 | r_CFStiffness | bilateral | 125100\ C99-5100-128\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-128, ShearWave Elastography Stiffness sample 5 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | Conf. Threshold | meas | m_kidney_EQI_6 | r_CFStiffness | bilateral | 125100\ C99-5100-129\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-129, ShearWave Elastography Stiffness sample 6 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | Conf. Threshold | meas | m_kidney_EQI_7 | r_CFStiffness | bilateral | 125100\ C99-5100-130\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-130, ShearWave Elastography Stiffness sample 7 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | Conf. Threshold | meas | m_kidney_EQI_8 | r_CFStiffness | bilateral | 125100\ C99-5100-131\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-131, ShearWave Elastography Stiffness sample 8 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | Conf. Threshold | meas | m_kidney_EQI_9 | r_CFStiffness | bilateral | 125100\ C99-5100-132\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-132, ShearWave Elastography Stiffness sample 9 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | Conf. Threshold | meas | m_kidney_EQI_10 | r_CFStiffness | bilateral | 125100\ C99-5100-133\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-133, ShearWave Elastography Stiffness sample 10 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | Circ | meas | m_kidney_EQI_4 | r_Circumference | bilateral | 125100\ C99-5100-137\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-137, ShearWave Elastography sample 4 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | Circ | meas | m_kidney_EQI_5 | r_Circumference | bilateral | 125100\ C99-5100-138\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-138, ShearWave Elastography sample 5 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | Circ | meas | m_kidney_EQI_6 | r_Circumference | bilateral | 125100\ C99-5100-139\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-139, ShearWave Elastography sample 6 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | Circ | meas | m_kidney_EQI_7 | r_Circumference | bilateral | 125100\ C99-5100-140\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-140, ShearWave Elastography sample 7 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | Circ | meas | m_kidney_EQI_8 | r_Circumference | bilateral | 125100\ C99-5100-141\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-141, ShearWave Elastography sample 8 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | Circ | meas | m_kidney_EQI_9 | r_Circumference | bilateral | 125100\ C99-5100-142\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-142, ShearWave Elastography sample 9 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | Circ | meas | m_kidney_EQI_10 | r_Circumference | bilateral | 125100\ C99-5100-143\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-143, ShearWave Elastography sample 10 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | Area | meas | m_kidney_EQI_4 | r_Area | bilateral | 125100\ C99-5100-147\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-147, ShearWave Elastography sample 4 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | Area | meas | m_kidney_EQI_5 | r_Area | bilateral | 125100\ C99-5100-148\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-148, ShearWave Elastography sample 5 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | Area | meas | m_kidney_EQI_6 | r_Area | bilateral | 125100\ C99-5100-149\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-149, ShearWave Elastography sample 6 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | Area | meas | m_kidney_EQI_7 | r_Area | bilateral | 125100\ C99-5100-150\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-150, ShearWave Elastography sample 7 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | Area | meas | m_kidney_EQI_8 | r_Area | bilateral | 125100\ C99-5100-151\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-151, ShearWave Elastography sample 8 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | Area | meas | m_kidney_EQI_9 | r_Area | bilateral | 125100\ C99-5100-152\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-152, ShearWave Elastography sample 9 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | Area | meas | m_kidney_EQI_10 | r_Area | bilateral | 125100\ C99-5100-153\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-153, ShearWave Elastography sample 10 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kid Height | Dist | meas | m_kidht | r_Distance | bilateral | 125100\C7470-01\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C7470-01, Height |  | UCUM, mm, mm | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | EQI Avg | meas | m_kidney_EQI_1 | r_AvgStiffness_p | bilateral | 125100\C99-5100-001\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-001, ShearWave Elastography Stiffness sample 1 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | EQI Avg | meas | m_kidney_EQI_2 | r_AvgStiffness_p | bilateral | 125100\C99-5100-002\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-002, ShearWave Elastography Stiffness sample 2 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | EQI Avg | meas | m_kidney_EQI_3 | r_AvgStiffness_p | bilateral | 125100\C99-5100-003\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-003, ShearWave Elastography Stiffness sample 3 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | EQI Std | meas | m_kidney_EQI_1 | r_StdStiffness_p | bilateral | 125100\C99-5100-044\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-044, ShearWave Elastography Stiffness sample 1 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | EQI Std | meas | m_kidney_EQI_2 | r_StdStiffness_p | bilateral | 125100\C99-5100-045\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-045, ShearWave Elastography Stiffness sample 2 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | EQI Std | meas | m_kidney_EQI_3 | r_StdStiffness_p | bilateral | 125100\C99-5100-046\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-046, ShearWave Elastography Stiffness sample 3 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | EQI Med | meas | m_kidney_EQI_1 | r_MedStiffness_p | bilateral | 125100\C99-5100-064\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-064, ShearWave Elastography Stiffness sample 1 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | EQI Med | meas | m_kidney_EQI_2 | r_MedStiffness_p | bilateral | 125100\C99-5100-065\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-065, ShearWave Elastography Stiffness sample 2 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | EQI Med | meas | m_kidney_EQI_3 | r_MedStiffness_p | bilateral | 125100\C99-5100-066\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-066, ShearWave Elastography Stiffness sample 3 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | EQI IQR | meas | m_kidney_EQI_1 | r_IqrStiffness_p | bilateral | 125100\C99-5100-084\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-084, ShearWave Elastography Stiffness sample 1 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | EQI IQR | meas | m_kidney_EQI_2 | r_IqrStiffness_p | bilateral | 125100\C99-5100-085\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-085, ShearWave Elastography Stiffness sample 2 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | EQI IQR | meas | m_kidney_EQI_3 | r_IqrStiffness_p | bilateral | 125100\C99-5100-086\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-086, ShearWave Elastography Stiffness sample 3 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | EQI IQR/Med | meas | m_kidney_EQI_1 | r_IqrMedStiffness_p | bilateral | 125100\C99-5100-104\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-104, ShearWave Elasto Stiffness sample 1 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | EQI IQR/Med | meas | m_kidney_EQI_2 | r_IqrMedStiffness_p | bilateral | 125100\C99-5100-105\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-105, ShearWave Elasto Stiffness sample 2 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | EQI IQR/Med | meas | m_kidney_EQI_3 | r_IqrMedStiffness_p | bilateral | 125100\C99-5100-106\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-106, ShearWave Elasto Stiffness sample 3 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | Conf. Threshold | meas | m_kidney_EQI_1 | r_CFStiffness | bilateral | 125100\C99-5100-124\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-124, ShearWave Elastography Stiffness sample 1 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | Conf. Threshold | meas | m_kidney_EQI_2 | r_CFStiffness | bilateral | 125100\C99-5100-125\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-125, ShearWave Elastography Stiffness sample 2 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | Conf. Threshold | meas | m_kidney_EQI_3 | r_CFStiffness | bilateral | 125100\C99-5100-126\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-126, ShearWave Elastography Stiffness sample 3 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | Circ | meas | m_kidney_EQI_1 | r_Circumference | bilateral | 125100\C99-5100-134\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-134, ShearWave Elastography sample 1 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | Circ | meas | m_kidney_EQI_2 | r_Circumference | bilateral | 125100\C99-5100-135\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-135, ShearWave Elastography sample 2 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | Circ | meas | m_kidney_EQI_3 | r_Circumference | bilateral | 125100\C99-5100-136\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-136, ShearWave Elastography sample 3 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | Area | meas | m_kidney_EQI_1 | r_Area | bilateral | 125100\C99-5100-144\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-144, ShearWave Elastography sample 1 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | Area | meas | m_kidney_EQI_2 | r_Area | bilateral | 125100\C99-5100-145\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-145, ShearWave Elastography sample 2 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | Area | meas | m_kidney_EQI_3 | r_Area | bilateral | 125100\C99-5100-146\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-146, ShearWave Elastography sample 3 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | EQI Min | meas | m_kidney_EQI_1 | r_MinStiffness_p | bilateral | 125100\C99-5100-154\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-154, ShearWave Elastography Stiffness sample 1 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | EQI Min | meas | m_kidney_EQI_2 | r_MinStiffness_p | bilateral | 125100\C99-5100-155\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-155, ShearWave Elastography Stiffness sample 2 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | EQI Min | meas | m_kidney_EQI_3 | r_MinStiffness_p | bilateral | 125100\C99-5100-156\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-156, ShearWave Elastography Stiffness sample 3 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | EQI Min | meas | m_kidney_EQI_4 | r_MinStiffness_p | bilateral | 125100\C99-5100-157\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-157, ShearWave Elastography Stiffness sample 4 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | EQI Min | meas | m_kidney_EQI_5 | r_MinStiffness_p | bilateral | 125100\C99-5100-158\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-158, ShearWave Elastography Stiffness sample 5 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | EQI Min | meas | m_kidney_EQI_6 | r_MinStiffness_p | bilateral | 125100\C99-5100-159\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-159, ShearWave Elastography Stiffness sample 6 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | EQI Min | meas | m_kidney_EQI_7 | r_MinStiffness_p | bilateral | 125100\C99-5100-160\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-160, ShearWave Elastography Stiffness sample 7 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | EQI Min | meas | m_kidney_EQI_8 | r_MinStiffness_p | bilateral | 125100\C99-5100-161\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-161, ShearWave Elastography Stiffness sample 8 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | EQI Min | meas | m_kidney_EQI_9 | r_MinStiffness_p | bilateral | 125100\C99-5100-162\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-162, ShearWave Elastography Stiffness sample 9 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | EQI Min | meas | m_kidney_EQI_10 | r_MinStiffness_p | bilateral | 125100\C99-5100-163\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-163, ShearWave Elastography Stiffness sample 10 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | EQI Min Vel | meas | m_kidney_EQI_1 | r_MinStiffness_v | bilateral | 125100\C99-5100-164\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-164, ShearWave Elastography Velocity sample 1 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | EQI Min Vel | meas | m_kidney_EQI_2 | r_MinStiffness_v | bilateral | 125100\C99-5100-165\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-165, ShearWave Elastography Velocity sample 2 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | EQI Min Vel | meas | m_kidney_EQI_3 | r_MinStiffness_v | bilateral | 125100\C99-5100-166\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-166, ShearWave Elastography Velocity sample 3 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | EQI Min Vel | meas | m_kidney_EQI_4 | r_MinStiffness_v | bilateral | 125100\C99-5100-167\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-167, ShearWave Elastography Velocity sample 4 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | EQI Min Vel | meas | m_kidney_EQI_5 | r_MinStiffness_v | bilateral | 125100\C99-5100-168\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-168, ShearWave Elastography Velocity sample 5 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | EQI Min Vel | meas | m_kidney_EQI_6 | r_MinStiffness_v | bilateral | 125100\C99-5100-169\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-169, ShearWave Elastography Velocity sample 6 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | EQI Min Vel | meas | m_kidney_EQI_7 | r_MinStiffness_v | bilateral | 125100\C99-5100-170\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-170, ShearWave Elastography Velocity sample 7 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | EQI Min Vel | meas | m_kidney_EQI_8 | r_MinStiffness_v | bilateral | 125100\C99-5100-171 \T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-171 , ShearWave Elastography Velocity sample 8 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | EQI Min Vel | meas | m_kidney_EQI_9 | r_MinStiffness_v | bilateral | 125100\C99-5100-172\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-172, ShearWave Elastography Velocity sample 9 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | EQI Min Vel | meas | m_kidney_EQI_10 | r_MinStiffness_v | bilateral | 125100\C99-5100-173\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-173, ShearWave Elastography Velocity sample 10 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | EQI Max | meas | m_kidney_EQI_1 | r_MaxStiffness_p | bilateral | 125100\C99-5100-174\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-174, ShearWave Elastography Stiffness sample 1 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | EQI Max | meas | m_kidney_EQI_2 | r_MaxStiffness_p | bilateral | 125100\C99-5100-175\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-175, ShearWave Elastography Stiffness sample 2 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | EQI Max | meas | m_kidney_EQI_3 | r_MaxStiffness_p | bilateral | 125100\C99-5100-176\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-176, ShearWave Elastography Stiffness sample 3 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | EQI Max | meas | m_kidney_EQI_4 | r_MaxStiffness_p | bilateral | 125100\C99-5100-177\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-177, ShearWave Elastography Stiffness sample 4 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | EQI Max | meas | m_kidney_EQI_5 | r_MaxStiffness_p | bilateral | 125100\C99-5100-178\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-178, ShearWave Elastography Stiffness sample 5 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | EQI Max | meas | m_kidney_EQI_6 | r_MaxStiffness_p | bilateral | 125100\C99-5100-179\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-179, ShearWave Elastography Stiffness sample 6 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | EQI Max | meas | m_kidney_EQI_7 | r_MaxStiffness_p | bilateral | 125100\C99-5100-180\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-180, ShearWave Elastography Stiffness sample 7 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | EQI Max | meas | m_kidney_EQI_8 | r_MaxStiffness_p | bilateral | 125100\C99-5100-181\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-181, ShearWave Elastography Stiffness sample 8 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | EQI Max | meas | m_kidney_EQI_9 | r_MaxStiffness_p | bilateral | 125100\C99-5100-182\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-182, ShearWave Elastography Stiffness sample 9 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | EQI Max | meas | m_kidney_EQI_10 | r_MaxStiffness_p | bilateral | 125100\C99-5100-183\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-183, ShearWave Elastography Stiffness sample 10 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 1 | EQI Max Vel | meas | m_kidney_EQI_1 | r_MaxStiffness_v | bilateral | 125100\C99-5100-184\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-184, ShearWave Elastography Velocity sample 1 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 2 | EQI Max Vel | meas | m_kidney_EQI_2 | r_MaxStiffness_v | bilateral | 125100\C99-5100-185\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-185, ShearWave Elastography Velocity sample 2 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 3 | EQI Max Vel | meas | m_kidney_EQI_3 | r_MaxStiffness_v | bilateral | 125100\C99-5100-186\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-186, ShearWave Elastography Velocity sample 3 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 4 | EQI Max Vel | meas | m_kidney_EQI_4 | r_MaxStiffness_v | bilateral | 125100\C99-5100-187\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-187, ShearWave Elastography Velocity sample 4 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 5 | EQI Max Vel | meas | m_kidney_EQI_5 | r_MaxStiffness_v | bilateral | 125100\C99-5100-188\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-188, ShearWave Elastography Velocity sample 5 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 6 | EQI Max Vel | meas | m_kidney_EQI_6 | r_MaxStiffness_v | bilateral | 125100\C99-5100-189\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-189, ShearWave Elastography Velocity sample 6 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 7 | EQI Max Vel | meas | m_kidney_EQI_7 | r_MaxStiffness_v | bilateral | 125100\C99-5100-190\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-190, ShearWave Elastography Velocity sample 7 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 8 | EQI Max Vel | meas | m_kidney_EQI_8 | r_MaxStiffness_v | bilateral | 125100\C99-5100-191 \T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-191 , ShearWave Elastography Velocity sample 8 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 9 | EQI Max Vel | meas | m_kidney_EQI_9 | r_MaxStiffness_v | bilateral | 125100\C99-5100-192\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-192, ShearWave Elastography Velocity sample 9 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI 10 | EQI Max Vel | meas | m_kidney_EQI_10 | r_MaxStiffness_v | bilateral | 125100\C99-5100-193\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-193, ShearWave Elastography Velocity sample 10 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kid Width | Dist | meas | m_kidwidth | r_Distance | bilateral | 125100\G-A220\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | SNM3, G-A220, Width |  | UCUM, mm, mm | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kid Length | Dist | meas | m_kidlen | r_Distance | bilateral | 125100\G-A22A\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | SRT, G-A22A, Length |  | UCUM, mm, mm | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Lung Angle | Angle | meas | m_LungAngle | r_Angle | unilateral | 125100\C7470-22\T-28000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C7470-22, Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-28000, Lung
Vascular Ultrasound Procedure Report | Abdominal | Prox Aorta AP | Dist | meas | m_proxAo_ap | r_Distance | unilateral | 125100\122675\G-A118\M-02550\T-42000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | SNM3, M-02550, Diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-42000, Aorta |  | DCM, 122675, Anterior-Posterior |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Aorta Transverse | Dist | meas | m_proxAo_trans | r_Distance | unilateral | 125100\G-A117\G-A118\M-02550\T-42000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | SNM3, M-02550, Diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-42000, Aorta |  | SRT, G-A117, Transverse |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Dist Aorta AP | Dist | meas | m_distAo_ap | r_Distance | unilateral | 125100\122675\G-A119\M-02550\T-42000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | SNM3, M-02550, Diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-42000, Aorta |  | DCM, 122675, Anterior-Posterior |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Aorta Transverse | Dist | meas | m_distAo_trans | r_Distance | unilateral | 125100\G-A117\G-A119\M-02550\T-42000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | SNM3, M-02550, Diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-42000, Aorta |  | SRT, G-A117, Transverse |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Mid Aorta AP | Dist | meas | m_midAo_ap | r_Distance | unilateral | 125100\122675\G-A188\M-02550\T-42000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | SNM3, M-02550, Diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-42000, Aorta |  | DCM, 122675, Anterior-Posterior |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Aorta Transverse | Dist | meas | m_midAo_trans | r_Distance | unilateral | 125100\G-A117\G-A188\M-02550\T-42000\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | SNM3, M-02550, Diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-42000, Aorta |  | SRT, G-A117, Transverse |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Iliac Art AP | Dist | meas | m_iliacAo_ap | r_Distance | bilateral | 125100\122675\M-02550\T-46700\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | SNM3, M-02550, Diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46700, Iliac artery |  | DCM, 122675, Anterior-Posterior
Vascular Ultrasound Procedure Report | Abdominal | Iliac Art Transverse | Dist | meas | m_iliacAo_trans | r_Distance | bilateral | 125100\G-A117\M-02550\T-46700\T5100-01 | 99PMSBLUS, T5100-01, Anatomic Structures |  | SNM3, M-02550, Diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46700, Iliac artery |  | SRT, G-A117, Transverse
Vascular Ultrasound Procedure Report | Abdominal | PV Bladder L | Dist | meas | m_pvbladder_len | r_Distance | unilateral | 125100\C12011-05\C4-01\T5100-02 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C12011-05, Post Void Bladder Length |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4-01, Post Void Bladder
Vascular Ultrasound Procedure Report | Abdominal | PV Bladder W | Dist | meas | m_pvbladder_wid | r_Distance | unilateral | 125100\C12011-06\C4-01\T5100-02 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C12011-06, Post Void Bladder Width |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4-01, Post Void Bladder
Vascular Ultrasound Procedure Report | Abdominal | PV Bladder H | Dist | meas | m_pvbladder_ht | r_Distance | unilateral | 125100\C12011-07\C4-01\T5100-02 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C12011-07, Post Void Bladder Height |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C4-01, Post Void Bladder
Vascular Ultrasound Procedure Report | Abdominal | CBD Diam | Dist | meas | m_cbd | r_Distance | unilateral | 125100\121206\T5100-02\T-60610 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | DCM, 121206, Distance |  | UCUM, mm, mm | [7 empty] | SNM3, T-60610, Bileduct
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | EQI Avg | meas | m_liver_EQI_4 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-004\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-004, ShearWave Elastography Stiffness sample 4 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | EQI Avg | meas | m_liver_EQI_5 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-005\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-005, ShearWave Elastography Stiffness sample 5 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | EQI Avg | meas | m_liver_EQI_6 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-006\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-006, ShearWave Elastography Stiffness sample 6 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | EQI Avg | meas | m_liver_EQI_7 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-007\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-007, ShearWave Elastography Stiffness sample 7 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | EQI Avg | meas | m_liver_EQI_8 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-008\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-008, ShearWave Elastography Stiffness sample 8 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | EQI Avg | meas | m_liver_EQI_9 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-009\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-009, ShearWave Elastography Stiffness sample 9 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | EQI Avg | meas | m_liver_EQI_10 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-010\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-010, ShearWave Elastography Stiffness sample 10 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | EQI Avg Vel | meas | m_liver_EQI_1 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-011\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-011, ShearWave Elastography Velocity sample 1 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | EQI Avg Vel | meas | m_liver_EQI_2 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-012\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-012, ShearWave Elastography Velocity sample 2 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | EQI Avg Vel | meas | m_liver_EQI_3 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-013\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-013, ShearWave Elastography Velocity sample 3 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | EQI Avg Vel | meas | m_liver_EQI_4 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-014\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-014, ShearWave Elastography Velocity sample 4 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | EQI Avg Vel | meas | m_liver_EQI_5 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-015\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-015, ShearWave Elastography Velocity sample 5 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | EQI Avg Vel | meas | m_liver_EQI_6 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-016\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-016, ShearWave Elastography Velocity sample 6 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | EQI Avg Vel | meas | m_liver_EQI_7 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-017\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-017, ShearWave Elastography Velocity sample 7 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | EQI Avg Vel | meas | m_liver_EQI_8 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-018 \T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-018 , ShearWave Elastography Velocity sample 8 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | EQI Avg Vel | meas | m_liver_EQI_9 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-019\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-019, ShearWave Elastography Velocity sample 9 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | EQI Avg Vel | meas | m_liver_EQI_10 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-020\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-020, ShearWave Elastography Velocity sample 10 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | EQI Std | meas | m_liver_EQI_4 | r_StdStiffness_p | unilateral | 125100\ C99-5100-047\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-047, ShearWave Elastography Stiffness sample 4 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | EQI Std | meas | m_liver_EQI_5 | r_StdStiffness_p | unilateral | 125100\ C99-5100-048\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-048, ShearWave Elastography Stiffness sample 5 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | EQI Std | meas | m_liver_EQI_6 | r_StdStiffness_p | unilateral | 125100\ C99-5100-049\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-049, ShearWave Elastography Stiffness sample 6 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | EQI Std | meas | m_liver_EQI_7 | r_StdStiffness_p | unilateral | 125100\ C99-5100-050\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-050, ShearWave Elastography Stiffness sample 7 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | EQI Std | meas | m_liver_EQI_8 | r_StdStiffness_p | unilateral | 125100\ C99-5100-051\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-051, ShearWave Elastography Stiffness sample 8 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | EQI Std | meas | m_liver_EQI_9 | r_StdStiffness_p | unilateral | 125100\ C99-5100-052\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-052, ShearWave Elastography Stiffness sample 9 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | EQI Std | meas | m_liver_EQI_10 | r_StdStiffness_p | unilateral | 125100\ C99-5100-053\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-053, ShearWave Elastography Stiffness sample 10 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | EQI Std Vel | meas | m_liver_EQI_1 | r_StdStiffness_v | unilateral | 125100\ C99-5100-054\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-054, ShearWave Elastography Velocity sample 1 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | EQI Std Vel | meas | m_liver_EQI_2 | r_StdStiffness_v | unilateral | 125100\ C99-5100-055\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-055, ShearWave Elastography Velocity sample 2 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | EQI Std Vel | meas | m_liver_EQI_3 | r_StdStiffness_v | unilateral | 125100\ C99-5100-056\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-056, ShearWave Elastography Velocity sample 3 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | EQI Std Vel | meas | m_liver_EQI_4 | r_StdStiffness_v | unilateral | 125100\ C99-5100-057\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-057, ShearWave Elastography Velocity sample 4 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | EQI Std Vel | meas | m_liver_EQI_5 | r_StdStiffness_v | unilateral | 125100\ C99-5100-058\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-058, ShearWave Elastography Velocity sample 5 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | EQI Std Vel | meas | m_liver_EQI_6 | r_StdStiffness_v | unilateral | 125100\ C99-5100-059\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-059, ShearWave Elastography Velocity sample 6 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | EQI Std Vel | meas | m_liver_EQI_7 | r_StdStiffness_v | unilateral | 125100\ C99-5100-060\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-060, ShearWave Elastography Velocity sample 7 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | EQI Std Vel | meas | m_liver_EQI_8 | r_StdStiffness_v | unilateral | 125100\ C99-5100-061\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-061, ShearWave Elastography Velocity sample 8 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | EQI Std Vel | meas | m_liver_EQI_9 | r_StdStiffness_v | unilateral | 125100\ C99-5100-062\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-062, ShearWave Elastography Velocity sample 9 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | EQI Std Vel | meas | m_liver_EQI_10 | r_StdStiffness_v | unilateral | 125100\ C99-5100-063\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-063, ShearWave Elastography Velocity sample 10 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | EQI Med | meas | m_liver_EQI_4 | r_MedStiffness_p | unilateral | 125100\ C99-5100-067\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-067, ShearWave Elastography Stiffness sample 4 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | EQI Med | meas | m_liver_EQI_5 | r_MedStiffness_p | unilateral | 125100\ C99-5100-068\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-068, ShearWave Elastography Stiffness sample 5 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | EQI Med | meas | m_liver_EQI_6 | r_MedStiffness_p | unilateral | 125100\ C99-5100-069\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-069, ShearWave Elastography Stiffness sample 6 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | EQI Med | meas | m_liver_EQI_7 | r_MedStiffness_p | unilateral | 125100\ C99-5100-070\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-070, ShearWave Elastography Stiffness sample 7 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | EQI Med | meas | m_liver_EQI_8 | r_MedStiffness_p | unilateral | 125100\ C99-5100-071\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-071, ShearWave Elastography Stiffness sample 8 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | EQI Med | meas | m_liver_EQI_9 | r_MedStiffness_p | unilateral | 125100\ C99-5100-072\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-072, ShearWave Elastography Stiffness sample 9 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | EQI Med | meas | m_liver_EQI_10 | r_MedStiffness_p | unilateral | 125100\ C99-5100-073\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-073, ShearWave Elastography Stiffness sample 10 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | EQI Med Vel | meas | m_liver_EQI_1 | r_MedStiffness_v | unilateral | 125100\ C99-5100-074\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-074, ShearWave Elastography Velocity sample 1 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | EQI Med Vel | meas | m_liver_EQI_2 | r_MedStiffness_v | unilateral | 125100\ C99-5100-075\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-075, ShearWave Elastography Velocity sample 2 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | EQI Med Vel | meas | m_liver_EQI_3 | r_MedStiffness_v | unilateral | 125100\ C99-5100-076\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-076, ShearWave Elastography Velocity sample 3 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | EQI Med Vel | meas | m_liver_EQI_4 | r_MedStiffness_v | unilateral | 125100\ C99-5100-077\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-077, ShearWave Elastography Velocity sample 4 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | EQI Med Vel | meas | m_liver_EQI_5 | r_MedStiffness_v | unilateral | 125100\ C99-5100-078\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-078, ShearWave Elastography Velocity sample 5 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | EQI Med Vel | meas | m_liver_EQI_6 | r_MedStiffness_v | unilateral | 125100\ C99-5100-079\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-079, ShearWave Elastography Velocity sample 6 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | EQI Med Vel | meas | m_liver_EQI_7 | r_MedStiffness_v | unilateral | 125100\ C99-5100-080\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-080, ShearWave Elastography Velocity sample 7 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | EQI Med Vel | meas | m_liver_EQI_8 | r_MedStiffness_v | unilateral | 125100\ C99-5100-081\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-081, ShearWave Elastography Velocity sample 8 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | EQI Med Vel | meas | m_liver_EQI_9 | r_MedStiffness_v | unilateral | 125100\ C99-5100-082\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-082, ShearWave Elastography Velocity sample 9 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | EQI Med Vel | meas | m_liver_EQI_10 | r_MedStiffness_v | unilateral | 125100\ C99-5100-083\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-083, ShearWave Elastography Velocity sample 10 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | EQI IQR | meas | m_liver_EQI_4 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-087\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-087, ShearWave Elastography Stiffness sample 4 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | EQI IQR | meas | m_liver_EQI_5 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-088\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-088, ShearWave Elastography Stiffness sample 5 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | EQI IQR | meas | m_liver_EQI_6 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-089\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-089, ShearWave Elastography Stiffness sample 6 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | EQI IQR | meas | m_liver_EQI_7 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-090\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-090, ShearWave Elastography Stiffness sample 7 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | EQI IQR | meas | m_liver_EQI_8 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-091\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-091, ShearWave Elastography Stiffness sample 8 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | EQI IQR | meas | m_liver_EQI_9 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-092\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-092, ShearWave Elastography Stiffness sample 9 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | EQI IQR | meas | m_liver_EQI_10 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-093\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-093, ShearWave Elastography Stiffness sample 10 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | EQI IQR Vel | meas | m_liver_EQI_1 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-094\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-094, ShearWave Elastography Velocity sample 1 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | EQI IQR Vel | meas | m_liver_EQI_2 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-095\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-095, ShearWave Elastography Velocity sample 2 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | EQI IQR Vel | meas | m_liver_EQI_3 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-096\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-096, ShearWave Elastography Velocity sample 3 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | EQI IQR Vel | meas | m_liver_EQI_4 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-097\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-097, ShearWave Elastography Velocity sample 4 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | EQI IQR Vel | meas | m_liver_EQI_5 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-098\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-098, ShearWave Elastography Velocity sample 5 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | EQI IQR Vel | meas | m_liver_EQI_6 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-099\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-099, ShearWave Elastography Velocity sample 6 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | EQI IQR Vel | meas | m_liver_EQI_7 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-100\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-100, ShearWave Elastography Velocity sample 7 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | EQI IQR Vel | meas | m_liver_EQI_8 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-101\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-101, ShearWave Elastography Velocity sample 8 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | EQI IQR Vel | meas | m_liver_EQI_9 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-102\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-102, ShearWave Elastography Velocity sample 9 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | EQI IQR Vel | meas | m_liver_EQI_10 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-103\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-103, ShearWave Elastography Velocity sample 10 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | EQI IQR/Med | meas | m_liver_EQI_4 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-107\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-107, ShearWave Elasto Stiffness sample 4 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | EQI IQR/Med | meas | m_liver_EQI_5 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-108\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-108, ShearWave Elasto Stiffness sample 5 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | EQI IQR/Med | meas | m_liver_EQI_6 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-109\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-109, ShearWave Elasto Stiffness sample 6 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | EQI IQR/Med | meas | m_liver_EQI_7 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-110\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-110, ShearWave Elasto Stiffness sample 7 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | EQI IQR/Med | meas | m_liver_EQI_8 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-111\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-111, ShearWave Elasto Stiffness sample 8 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | EQI IQR/Med | meas | m_liver_EQI_9 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-112\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-112, ShearWave Elasto Stiffness sample 9 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | EQI IQR/Med | meas | m_liver_EQI_10 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-113\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-113, ShearWave Elasto Stiffness sample 10 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | EQI IQR/Med Vel | meas | m_liver_EQI_1 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-114\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-114, ShearWave Elasto Velocity sample 1 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | EQI IQR/Med Vel | meas | m_liver_EQI_2 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-115\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-115, ShearWave Elasto Velocity sample 2 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | EQI IQR/Med Vel | meas | m_liver_EQI_3 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-116\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-116, ShearWave Elasto Velocity sample 3 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | EQI IQR/Med Vel | meas | m_liver_EQI_4 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-117\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-117, ShearWave Elasto Velocity sample 4 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | EQI IQR/Med Vel | meas | m_liver_EQI_5 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-118\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-118, ShearWave Elasto Velocity sample 5 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | EQI IQR/Med Vel | meas | m_liver_EQI_6 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-119\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-119, ShearWave Elasto Velocity sample 6 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | EQI IQR/Med Vel | meas | m_liver_EQI_7 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-120\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-120, ShearWave Elasto Velocity sample 7 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | EQI IQR/Med Vel | meas | m_liver_EQI_8 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-121\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-121, ShearWave Elasto Velocity sample 8 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | EQI IQR/Med Vel | meas | m_liver_EQI_9 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-122\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-122, ShearWave Elasto Velocity sample 9 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | EQI IQR/Med Vel | meas | m_liver_EQI_10 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-123\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-123, ShearWave Elasto Velocity sample 10 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | Conf. Threshold | meas | m_liver_EQI_4 | r_CFStiffness | unilateral | 125100\ C99-5100-127\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-127, ShearWave Elastography Stiffness sample 4 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | Conf. Threshold | meas | m_liver_EQI_5 | r_CFStiffness | unilateral | 125100\ C99-5100-128\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-128, ShearWave Elastography Stiffness sample 5 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | Conf. Threshold | meas | m_liver_EQI_6 | r_CFStiffness | unilateral | 125100\ C99-5100-129\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-129, ShearWave Elastography Stiffness sample 6 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | Conf. Threshold | meas | m_liver_EQI_7 | r_CFStiffness | unilateral | 125100\ C99-5100-130\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-130, ShearWave Elastography Stiffness sample 7 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | Conf. Threshold | meas | m_liver_EQI_8 | r_CFStiffness | unilateral | 125100\ C99-5100-131\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-131, ShearWave Elastography Stiffness sample 8 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | Conf. Threshold | meas | m_liver_EQI_9 | r_CFStiffness | unilateral | 125100\ C99-5100-132\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-132, ShearWave Elastography Stiffness sample 9 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | Conf. Threshold | meas | m_liver_EQI_10 | r_CFStiffness | unilateral | 125100\ C99-5100-133\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-133, ShearWave Elastography Stiffness sample 10 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | Circ | meas | m_liver_EQI_4 | r_Circumference | unilateral | 125100\ C99-5100-137\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-137, ShearWave Elastography sample 4 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | Circ | meas | m_liver_EQI_5 | r_Circumference | unilateral | 125100\ C99-5100-138\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-138, ShearWave Elastography sample 5 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | Circ | meas | m_liver_EQI_6 | r_Circumference | unilateral | 125100\ C99-5100-139\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-139, ShearWave Elastography sample 6 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | Circ | meas | m_liver_EQI_7 | r_Circumference | unilateral | 125100\ C99-5100-140\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-140, ShearWave Elastography sample 7 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | Circ | meas | m_liver_EQI_8 | r_Circumference | unilateral | 125100\ C99-5100-141\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-141, ShearWave Elastography sample 8 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | Circ | meas | m_liver_EQI_9 | r_Circumference | unilateral | 125100\ C99-5100-142\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-142, ShearWave Elastography sample 9 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | Circ | meas | m_liver_EQI_10 | r_Circumference | unilateral | 125100\ C99-5100-143\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-143, ShearWave Elastography sample 10 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | Area | meas | m_liver_EQI_4 | r_Area | unilateral | 125100\ C99-5100-147\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-147, ShearWave Elastography sample 4 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | Area | meas | m_liver_EQI_5 | r_Area | unilateral | 125100\ C99-5100-148\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-148, ShearWave Elastography sample 5 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | Area | meas | m_liver_EQI_6 | r_Area | unilateral | 125100\ C99-5100-149\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-149, ShearWave Elastography sample 6 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | Area | meas | m_liver_EQI_7 | r_Area | unilateral | 125100\ C99-5100-150\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-150, ShearWave Elastography sample 7 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | Area | meas | m_liver_EQI_8 | r_Area | unilateral | 125100\ C99-5100-151\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-151, ShearWave Elastography sample 8 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | Area | meas | m_liver_EQI_9 | r_Area | unilateral | 125100\ C99-5100-152\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-152, ShearWave Elastography sample 9 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | Area | meas | m_liver_EQI_10 | r_Area | unilateral | 125100\ C99-5100-153\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-153, ShearWave Elastography sample 10 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | EQI Avg | meas | m_liver_EQI_1 | r_AvgStiffness_p | unilateral | 125100\C99-5100-001\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-001, ShearWave Elastography Stiffness sample 1 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | EQI Avg | meas | m_liver_EQI_2 | r_AvgStiffness_p | unilateral | 125100\C99-5100-002\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-002, ShearWave Elastography Stiffness sample 2 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | EQI Avg | meas | m_liver_EQI_3 | r_AvgStiffness_p | unilateral | 125100\C99-5100-003\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-003, ShearWave Elastography Stiffness sample 3 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | EQI Std | meas | m_liver_EQI_1 | r_StdStiffness_p | unilateral | 125100\C99-5100-044\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-044, ShearWave Elastography Stiffness sample 1 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | EQI Std | meas | m_liver_EQI_2 | r_StdStiffness_p | unilateral | 125100\C99-5100-045\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-045, ShearWave Elastography Stiffness sample 2 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | EQI Std | meas | m_liver_EQI_3 | r_StdStiffness_p | unilateral | 125100\C99-5100-046\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-046, ShearWave Elastography Stiffness sample 3 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | EQI Med | meas | m_liver_EQI_1 | r_MedStiffness_p | unilateral | 125100\C99-5100-064\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-064, ShearWave Elastography Stiffness sample 1 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | EQI Med | meas | m_liver_EQI_2 | r_MedStiffness_p | unilateral | 125100\C99-5100-065\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-065, ShearWave Elastography Stiffness sample 2 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | EQI Med | meas | m_liver_EQI_3 | r_MedStiffness_p | unilateral | 125100\C99-5100-066\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-066, ShearWave Elastography Stiffness sample 3 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | EQI IQR | meas | m_liver_EQI_1 | r_IqrStiffness_p | unilateral | 125100\C99-5100-084\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-084, ShearWave Elastography Stiffness sample 1 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | EQI IQR | meas | m_liver_EQI_2 | r_IqrStiffness_p | unilateral | 125100\C99-5100-085\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-085, ShearWave Elastography Stiffness sample 2 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | EQI IQR | meas | m_liver_EQI_3 | r_IqrStiffness_p | unilateral | 125100\C99-5100-086\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-086, ShearWave Elastography Stiffness sample 3 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | EQI IQR/Med | meas | m_liver_EQI_1 | r_IqrMedStiffness_p | unilateral | 125100\C99-5100-104\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-104, ShearWave Elasto Stiffness sample 1 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | EQI IQR/Med | meas | m_liver_EQI_2 | r_IqrMedStiffness_p | unilateral | 125100\C99-5100-105\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-105, ShearWave Elasto Stiffness sample 2 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | EQI IQR/Med | meas | m_liver_EQI_3 | r_IqrMedStiffness_p | unilateral | 125100\C99-5100-106\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-106, ShearWave Elasto Stiffness sample 3 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | Conf. Threshold | meas | m_liver_EQI_1 | r_CFStiffness | unilateral | 125100\C99-5100-124\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-124, ShearWave Elastography Stiffness sample 1 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | Conf. Threshold | meas | m_liver_EQI_2 | r_CFStiffness | unilateral | 125100\C99-5100-125\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-125, ShearWave Elastography Stiffness sample 2 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | Conf. Threshold | meas | m_liver_EQI_3 | r_CFStiffness | unilateral | 125100\C99-5100-126\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-126, ShearWave Elastography Stiffness sample 3 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | Circ | meas | m_liver_EQI_1 | r_Circumference | unilateral | 125100\C99-5100-134\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-134, ShearWave Elastography sample 1 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | Circ | meas | m_liver_EQI_2 | r_Circumference | unilateral | 125100\C99-5100-135\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-135, ShearWave Elastography sample 2 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | Circ | meas | m_liver_EQI_3 | r_Circumference | unilateral | 125100\C99-5100-136\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-136, ShearWave Elastography sample 3 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | Area | meas | m_liver_EQI_1 | r_Area | unilateral | 125100\C99-5100-144\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-144, ShearWave Elastography sample 1 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | Area | meas | m_liver_EQI_2 | r_Area | unilateral | 125100\C99-5100-145\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-145, ShearWave Elastography sample 2 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | Area | meas | m_liver_EQI_3 | r_Area | unilateral | 125100\C99-5100-146\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-146, ShearWave Elastography sample 3 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | EQI Min | meas | m_liver_EQI_1 | r_MinStiffness_p | unilateral | 125100\C99-5100-154\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-154, ShearWave Elastography Stiffness sample 1 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | EQI Min | meas | m_liver_EQI_2 | r_MinStiffness_p | unilateral | 125100\C99-5100-155\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-155, ShearWave Elastography Stiffness sample 2 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | EQI Min | meas | m_liver_EQI_3 | r_MinStiffness_p | unilateral | 125100\C99-5100-156\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-156, ShearWave Elastography Stiffness sample 3 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | EQI Min | meas | m_liver_EQI_4 | r_MinStiffness_p | unilateral | 125100\C99-5100-157\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-157, ShearWave Elastography Stiffness sample 4 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | EQI Min | meas | m_liver_EQI_5 | r_MinStiffness_p | unilateral | 125100\C99-5100-158\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-158, ShearWave Elastography Stiffness sample 5 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | EQI Min | meas | m_liver_EQI_6 | r_MinStiffness_p | unilateral | 125100\C99-5100-159\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-159, ShearWave Elastography Stiffness sample 6 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | EQI Min | meas | m_liver_EQI_7 | r_MinStiffness_p | unilateral | 125100\C99-5100-160\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-160, ShearWave Elastography Stiffness sample 7 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | EQI Min | meas | m_liver_EQI_8 | r_MinStiffness_p | unilateral | 125100\C99-5100-161\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-161, ShearWave Elastography Stiffness sample 8 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | EQI Min | meas | m_liver_EQI_9 | r_MinStiffness_p | unilateral | 125100\C99-5100-162\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-162, ShearWave Elastography Stiffness sample 9 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | EQI Min | meas | m_liver_EQI_10 | r_MinStiffness_p | unilateral | 125100\C99-5100-163\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-163, ShearWave Elastography Stiffness sample 10 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | EQI Min Vel | meas | m_liver_EQI_1 | r_MinStiffness_v | unilateral | 125100\C99-5100-164\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-164, ShearWave Elastography Velocity sample 1 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | EQI Min Vel | meas | m_liver_EQI_2 | r_MinStiffness_v | unilateral | 125100\C99-5100-165\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-165, ShearWave Elastography Velocity sample 2 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | EQI Min Vel | meas | m_liver_EQI_3 | r_MinStiffness_v | unilateral | 125100\C99-5100-166\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-166, ShearWave Elastography Velocity sample 3 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | EQI Min Vel | meas | m_liver_EQI_4 | r_MinStiffness_v | unilateral | 125100\C99-5100-167\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-167, ShearWave Elastography Velocity sample 4 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | EQI Min Vel | meas | m_liver_EQI_5 | r_MinStiffness_v | unilateral | 125100\C99-5100-168\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-168, ShearWave Elastography Velocity sample 5 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | EQI Min Vel | meas | m_liver_EQI_6 | r_MinStiffness_v | unilateral | 125100\C99-5100-169\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-169, ShearWave Elastography Velocity sample 6 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | EQI Min Vel | meas | m_liver_EQI_7 | r_MinStiffness_v | unilateral | 125100\C99-5100-170\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-170, ShearWave Elastography Velocity sample 7 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | EQI Min Vel | meas | m_liver_EQI_8 | r_MinStiffness_v | unilateral | 125100\C99-5100-171 \T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-171 , ShearWave Elastography Velocity sample 8 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | EQI Min Vel | meas | m_liver_EQI_9 | r_MinStiffness_v | unilateral | 125100\C99-5100-172\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-172, ShearWave Elastography Velocity sample 9 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | EQI Min Vel | meas | m_liver_EQI_10 | r_MinStiffness_v | unilateral | 125100\C99-5100-173\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-173, ShearWave Elastography Velocity sample 10 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | EQI Max | meas | m_liver_EQI_1 | r_MaxStiffness_p | unilateral | 125100\C99-5100-174\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-174, ShearWave Elastography Stiffness sample 1 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | EQI Max | meas | m_liver_EQI_2 | r_MaxStiffness_p | unilateral | 125100\C99-5100-175\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-175, ShearWave Elastography Stiffness sample 2 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | EQI Max | meas | m_liver_EQI_3 | r_MaxStiffness_p | unilateral | 125100\C99-5100-176\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-176, ShearWave Elastography Stiffness sample 3 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | EQI Max | meas | m_liver_EQI_4 | r_MaxStiffness_p | unilateral | 125100\C99-5100-177\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-177, ShearWave Elastography Stiffness sample 4 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | EQI Max | meas | m_liver_EQI_5 | r_MaxStiffness_p | unilateral | 125100\C99-5100-178\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-178, ShearWave Elastography Stiffness sample 5 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | EQI Max | meas | m_liver_EQI_6 | r_MaxStiffness_p | unilateral | 125100\C99-5100-179\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-179, ShearWave Elastography Stiffness sample 6 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | EQI Max | meas | m_liver_EQI_7 | r_MaxStiffness_p | unilateral | 125100\C99-5100-180\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-180, ShearWave Elastography Stiffness sample 7 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | EQI Max | meas | m_liver_EQI_8 | r_MaxStiffness_p | unilateral | 125100\C99-5100-181\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-181, ShearWave Elastography Stiffness sample 8 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | EQI Max | meas | m_liver_EQI_9 | r_MaxStiffness_p | unilateral | 125100\C99-5100-182\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-182, ShearWave Elastography Stiffness sample 9 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | EQI Max | meas | m_liver_EQI_10 | r_MaxStiffness_p | unilateral | 125100\C99-5100-183\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-183, ShearWave Elastography Stiffness sample 10 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 1 | EQI Max Vel | meas | m_liver_EQI_1 | r_MaxStiffness_v | unilateral | 125100\C99-5100-184\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-184, ShearWave Elastography Velocity sample 1 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 2 | EQI Max Vel | meas | m_liver_EQI_2 | r_MaxStiffness_v | unilateral | 125100\C99-5100-185\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-185, ShearWave Elastography Velocity sample 2 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 3 | EQI Max Vel | meas | m_liver_EQI_3 | r_MaxStiffness_v | unilateral | 125100\C99-5100-186\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-186, ShearWave Elastography Velocity sample 3 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 4 | EQI Max Vel | meas | m_liver_EQI_4 | r_MaxStiffness_v | unilateral | 125100\C99-5100-187\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-187, ShearWave Elastography Velocity sample 4 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 5 | EQI Max Vel | meas | m_liver_EQI_5 | r_MaxStiffness_v | unilateral | 125100\C99-5100-188\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-188, ShearWave Elastography Velocity sample 5 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 6 | EQI Max Vel | meas | m_liver_EQI_6 | r_MaxStiffness_v | unilateral | 125100\C99-5100-189\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-189, ShearWave Elastography Velocity sample 6 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 7 | EQI Max Vel | meas | m_liver_EQI_7 | r_MaxStiffness_v | unilateral | 125100\C99-5100-190\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-190, ShearWave Elastography Velocity sample 7 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 8 | EQI Max Vel | meas | m_liver_EQI_8 | r_MaxStiffness_v | unilateral | 125100\C99-5100-191 \T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-191 , ShearWave Elastography Velocity sample 8 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 9 | EQI Max Vel | meas | m_liver_EQI_9 | r_MaxStiffness_v | unilateral | 125100\C99-5100-192\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-192, ShearWave Elastography Velocity sample 9 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI 10 | EQI Max Vel | meas | m_liver_EQI_10 | r_MaxStiffness_v | unilateral | 125100\C99-5100-193\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-193, ShearWave Elastography Velocity sample 10 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver Length | Dist | meas | m_liverLen | r_Distance | unilateral | 125100\G-A22A\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | SRT, G-A22A, Length |  | UCUM, mm, mm | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | GB Wall Diam | Dist | meas | m_gbWall | r_Distance | unilateral | 125100\121206\T5100-02\T-63000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | DCM, 121206, Distance |  | UCUM, mm, mm | [7 empty] | SNM3, T-63000, Gallbladder
Vascular Ultrasound Procedure Report | Abdominal | Bladder L | Dist | meas | m_bladder_len | r_Distance | unilateral | 125100\C12011-01\T5100-02\T-74000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C12011-01, Bladder Length |  | UCUM, mm, mm | [7 empty] | SNM3, T-74000, Bladder
Vascular Ultrasound Procedure Report | Abdominal | Bladder W | Dist | meas | m_bladder_wid | r_Distance | unilateral | 125100\C12011-02\T5100-02\T-74000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C12011-02, Bladder Width |  | UCUM, mm, mm | [7 empty] | SNM3, T-74000, Bladder
Vascular Ultrasound Procedure Report | Abdominal | Bladder H | Dist | meas | m_bladder_ht | r_Distance | unilateral | 125100\C12011-03\T5100-02\T-74000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C12011-03, Bladder Height |  | UCUM, mm, mm | [7 empty] | SNM3, T-74000, Bladder
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | EQI Avg | meas | m_spleen_EQI_4 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-004\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-004, ShearWave Elastography Stiffness sample 4 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | EQI Avg | meas | m_spleen_EQI_5 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-005\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-005, ShearWave Elastography Stiffness sample 5 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | EQI Avg | meas | m_spleen_EQI_6 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-006\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-006, ShearWave Elastography Stiffness sample 6 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | EQI Avg | meas | m_spleen_EQI_7 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-007\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-007, ShearWave Elastography Stiffness sample 7 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | EQI Avg | meas | m_spleen_EQI_8 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-008\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-008, ShearWave Elastography Stiffness sample 8 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | EQI Avg | meas | m_spleen_EQI_9 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-009\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-009, ShearWave Elastography Stiffness sample 9 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | EQI Avg | meas | m_spleen_EQI_10 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-010\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-010, ShearWave Elastography Stiffness sample 10 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | EQI Avg Vel | meas | m_spleen_EQI_1 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-011\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-011, ShearWave Elastography Velocity sample 1 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | EQI Avg Vel | meas | m_spleen_EQI_2 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-012\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-012, ShearWave Elastography Velocity sample 2 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | EQI Avg Vel | meas | m_spleen_EQI_3 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-013\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-013, ShearWave Elastography Velocity sample 3 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | EQI Avg Vel | meas | m_spleen_EQI_4 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-014\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-014, ShearWave Elastography Velocity sample 4 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | EQI Avg Vel | meas | m_spleen_EQI_5 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-015\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-015, ShearWave Elastography Velocity sample 5 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | EQI Avg Vel | meas | m_spleen_EQI_6 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-016\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-016, ShearWave Elastography Velocity sample 6 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | EQI Avg Vel | meas | m_spleen_EQI_7 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-017\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-017, ShearWave Elastography Velocity sample 7 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | EQI Avg Vel | meas | m_spleen_EQI_8 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-018 \T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-018 , ShearWave Elastography Velocity sample 8 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | EQI Avg Vel | meas | m_spleen_EQI_9 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-019\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-019, ShearWave Elastography Velocity sample 9 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | EQI Avg Vel | meas | m_spleen_EQI_10 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-020\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-020, ShearWave Elastography Velocity sample 10 Average |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | EQI Std | meas | m_spleen_EQI_4 | r_StdStiffness_p | unilateral | 125100\ C99-5100-047\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-047, ShearWave Elastography Stiffness sample 4 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | EQI Std | meas | m_spleen_EQI_5 | r_StdStiffness_p | unilateral | 125100\ C99-5100-048\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-048, ShearWave Elastography Stiffness sample 5 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | EQI Std | meas | m_spleen_EQI_6 | r_StdStiffness_p | unilateral | 125100\ C99-5100-049\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-049, ShearWave Elastography Stiffness sample 6 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | EQI Std | meas | m_spleen_EQI_7 | r_StdStiffness_p | unilateral | 125100\ C99-5100-050\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-050, ShearWave Elastography Stiffness sample 7 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | EQI Std | meas | m_spleen_EQI_8 | r_StdStiffness_p | unilateral | 125100\ C99-5100-051\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-051, ShearWave Elastography Stiffness sample 8 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | EQI Std | meas | m_spleen_EQI_9 | r_StdStiffness_p | unilateral | 125100\ C99-5100-052\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-052, ShearWave Elastography Stiffness sample 9 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | EQI Std | meas | m_spleen_EQI_10 | r_StdStiffness_p | unilateral | 125100\ C99-5100-053\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-053, ShearWave Elastography Stiffness sample 10 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | EQI Std Vel | meas | m_spleen_EQI_1 | r_StdStiffness_v | unilateral | 125100\ C99-5100-054\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-054, ShearWave Elastography Velocity sample 1 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | EQI Std Vel | meas | m_spleen_EQI_2 | r_StdStiffness_v | unilateral | 125100\ C99-5100-055\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-055, ShearWave Elastography Velocity sample 2 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | EQI Std Vel | meas | m_spleen_EQI_3 | r_StdStiffness_v | unilateral | 125100\ C99-5100-056\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-056, ShearWave Elastography Velocity sample 3 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | EQI Std Vel | meas | m_spleen_EQI_4 | r_StdStiffness_v | unilateral | 125100\ C99-5100-057\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-057, ShearWave Elastography Velocity sample 4 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | EQI Std Vel | meas | m_spleen_EQI_5 | r_StdStiffness_v | unilateral | 125100\ C99-5100-058\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-058, ShearWave Elastography Velocity sample 5 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | EQI Std Vel | meas | m_spleen_EQI_6 | r_StdStiffness_v | unilateral | 125100\ C99-5100-059\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-059, ShearWave Elastography Velocity sample 6 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | EQI Std Vel | meas | m_spleen_EQI_7 | r_StdStiffness_v | unilateral | 125100\ C99-5100-060\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-060, ShearWave Elastography Velocity sample 7 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | EQI Std Vel | meas | m_spleen_EQI_8 | r_StdStiffness_v | unilateral | 125100\ C99-5100-061\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-061, ShearWave Elastography Velocity sample 8 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | EQI Std Vel | meas | m_spleen_EQI_9 | r_StdStiffness_v | unilateral | 125100\ C99-5100-062\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-062, ShearWave Elastography Velocity sample 9 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | EQI Std Vel | meas | m_spleen_EQI_10 | r_StdStiffness_v | unilateral | 125100\ C99-5100-063\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-063, ShearWave Elastography Velocity sample 10 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | EQI Med | meas | m_spleen_EQI_4 | r_MedStiffness_p | unilateral | 125100\ C99-5100-067\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-067, ShearWave Elastography Stiffness sample 4 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | EQI Med | meas | m_spleen_EQI_5 | r_MedStiffness_p | unilateral | 125100\ C99-5100-068\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-068, ShearWave Elastography Stiffness sample 5 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | EQI Med | meas | m_spleen_EQI_6 | r_MedStiffness_p | unilateral | 125100\ C99-5100-069\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-069, ShearWave Elastography Stiffness sample 6 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | EQI Med | meas | m_spleen_EQI_7 | r_MedStiffness_p | unilateral | 125100\ C99-5100-070\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-070, ShearWave Elastography Stiffness sample 7 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | EQI Med | meas | m_spleen_EQI_8 | r_MedStiffness_p | unilateral | 125100\ C99-5100-071\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-071, ShearWave Elastography Stiffness sample 8 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | EQI Med | meas | m_spleen_EQI_9 | r_MedStiffness_p | unilateral | 125100\ C99-5100-072\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-072, ShearWave Elastography Stiffness sample 9 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | EQI Med | meas | m_spleen_EQI_10 | r_MedStiffness_p | unilateral | 125100\ C99-5100-073\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-073, ShearWave Elastography Stiffness sample 10 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | EQI Med Vel | meas | m_spleen_EQI_1 | r_MedStiffness_v | unilateral | 125100\ C99-5100-074\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-074, ShearWave Elastography Velocity sample 1 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | EQI Med Vel | meas | m_spleen_EQI_2 | r_MedStiffness_v | unilateral | 125100\ C99-5100-075\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-075, ShearWave Elastography Velocity sample 2 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | EQI Med Vel | meas | m_spleen_EQI_3 | r_MedStiffness_v | unilateral | 125100\ C99-5100-076\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-076, ShearWave Elastography Velocity sample 3 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | EQI Med Vel | meas | m_spleen_EQI_4 | r_MedStiffness_v | unilateral | 125100\ C99-5100-077\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-077, ShearWave Elastography Velocity sample 4 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | EQI Med Vel | meas | m_spleen_EQI_5 | r_MedStiffness_v | unilateral | 125100\ C99-5100-078\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-078, ShearWave Elastography Velocity sample 5 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | EQI Med Vel | meas | m_spleen_EQI_6 | r_MedStiffness_v | unilateral | 125100\ C99-5100-079\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-079, ShearWave Elastography Velocity sample 6 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | EQI Med Vel | meas | m_spleen_EQI_7 | r_MedStiffness_v | unilateral | 125100\ C99-5100-080\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-080, ShearWave Elastography Velocity sample 7 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | EQI Med Vel | meas | m_spleen_EQI_8 | r_MedStiffness_v | unilateral | 125100\ C99-5100-081\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-081, ShearWave Elastography Velocity sample 8 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | EQI Med Vel | meas | m_spleen_EQI_9 | r_MedStiffness_v | unilateral | 125100\ C99-5100-082\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-082, ShearWave Elastography Velocity sample 9 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | EQI Med Vel | meas | m_spleen_EQI_10 | r_MedStiffness_v | unilateral | 125100\ C99-5100-083\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-083, ShearWave Elastography Velocity sample 10 Median |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | EQI IQR | meas | m_spleen_EQI_4 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-087\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-087, ShearWave Elastography Stiffness sample 4 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | EQI IQR | meas | m_spleen_EQI_5 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-088\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-088, ShearWave Elastography Stiffness sample 5 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | EQI IQR | meas | m_spleen_EQI_6 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-089\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-089, ShearWave Elastography Stiffness sample 6 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | EQI IQR | meas | m_spleen_EQI_7 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-090\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-090, ShearWave Elastography Stiffness sample 7 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | EQI IQR | meas | m_spleen_EQI_8 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-091\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-091, ShearWave Elastography Stiffness sample 8 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | EQI IQR | meas | m_spleen_EQI_9 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-092\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-092, ShearWave Elastography Stiffness sample 9 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | EQI IQR | meas | m_spleen_EQI_10 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-093\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-093, ShearWave Elastography Stiffness sample 10 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | EQI IQR Vel | meas | m_spleen_EQI_1 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-094\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-094, ShearWave Elastography Velocity sample 1 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | EQI IQR Vel | meas | m_spleen_EQI_2 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-095\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-095, ShearWave Elastography Velocity sample 2 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | EQI IQR Vel | meas | m_spleen_EQI_3 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-096\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-096, ShearWave Elastography Velocity sample 3 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | EQI IQR Vel | meas | m_spleen_EQI_4 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-097\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-097, ShearWave Elastography Velocity sample 4 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | EQI IQR Vel | meas | m_spleen_EQI_5 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-098\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-098, ShearWave Elastography Velocity sample 5 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | EQI IQR Vel | meas | m_spleen_EQI_6 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-099\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-099, ShearWave Elastography Velocity sample 6 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | EQI IQR Vel | meas | m_spleen_EQI_7 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-100\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-100, ShearWave Elastography Velocity sample 7 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | EQI IQR Vel | meas | m_spleen_EQI_8 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-101\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-101, ShearWave Elastography Velocity sample 8 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | EQI IQR Vel | meas | m_spleen_EQI_9 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-102\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-102, ShearWave Elastography Velocity sample 9 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | EQI IQR Vel | meas | m_spleen_EQI_10 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-103\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-103, ShearWave Elastography Velocity sample 10 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | EQI IQR/Med | meas | m_spleen_EQI_4 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-107\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-107, ShearWave Elasto Stiffness sample 4 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | EQI IQR/Med | meas | m_spleen_EQI_5 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-108\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-108, ShearWave Elasto Stiffness sample 5 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | EQI IQR/Med | meas | m_spleen_EQI_6 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-109\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-109, ShearWave Elasto Stiffness sample 6 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | EQI IQR/Med | meas | m_spleen_EQI_7 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-110\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-110, ShearWave Elasto Stiffness sample 7 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | EQI IQR/Med | meas | m_spleen_EQI_8 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-111\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-111, ShearWave Elasto Stiffness sample 8 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | EQI IQR/Med | meas | m_spleen_EQI_9 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-112\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-112, ShearWave Elasto Stiffness sample 9 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | EQI IQR/Med | meas | m_spleen_EQI_10 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-113\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-113, ShearWave Elasto Stiffness sample 10 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | EQI IQR/Med Vel | meas | m_spleen_EQI_1 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-114\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-114, ShearWave Elasto Velocity sample 1 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | EQI IQR/Med Vel | meas | m_spleen_EQI_2 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-115\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-115, ShearWave Elasto Velocity sample 2 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | EQI IQR/Med Vel | meas | m_spleen_EQI_3 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-116\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-116, ShearWave Elasto Velocity sample 3 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | EQI IQR/Med Vel | meas | m_spleen_EQI_4 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-117\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-117, ShearWave Elasto Velocity sample 4 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | EQI IQR/Med Vel | meas | m_spleen_EQI_5 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-118\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-118, ShearWave Elasto Velocity sample 5 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | EQI IQR/Med Vel | meas | m_spleen_EQI_6 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-119\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-119, ShearWave Elasto Velocity sample 6 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | EQI IQR/Med Vel | meas | m_spleen_EQI_7 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-120\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-120, ShearWave Elasto Velocity sample 7 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | EQI IQR/Med Vel | meas | m_spleen_EQI_8 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-121\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-121, ShearWave Elasto Velocity sample 8 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | EQI IQR/Med Vel | meas | m_spleen_EQI_9 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-122\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-122, ShearWave Elasto Velocity sample 9 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | EQI IQR/Med Vel | meas | m_spleen_EQI_10 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-123\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-123, ShearWave Elasto Velocity sample 10 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | Conf. Threshold | meas | m_spleen_EQI_4 | r_CFStiffness | unilateral | 125100\ C99-5100-127\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-127, ShearWave Elastography Stiffness sample 4 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | Conf. Threshold | meas | m_spleen_EQI_5 | r_CFStiffness | unilateral | 125100\ C99-5100-128\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-128, ShearWave Elastography Stiffness sample 5 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | Conf. Threshold | meas | m_spleen_EQI_6 | r_CFStiffness | unilateral | 125100\ C99-5100-129\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-129, ShearWave Elastography Stiffness sample 6 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | Conf. Threshold | meas | m_spleen_EQI_7 | r_CFStiffness | unilateral | 125100\ C99-5100-130\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-130, ShearWave Elastography Stiffness sample 7 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | Conf. Threshold | meas | m_spleen_EQI_8 | r_CFStiffness | unilateral | 125100\ C99-5100-131\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-131, ShearWave Elastography Stiffness sample 8 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | Conf. Threshold | meas | m_spleen_EQI_9 | r_CFStiffness | unilateral | 125100\ C99-5100-132\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-132, ShearWave Elastography Stiffness sample 9 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | Conf. Threshold | meas | m_spleen_EQI_10 | r_CFStiffness | unilateral | 125100\ C99-5100-133\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-133, ShearWave Elastography Stiffness sample 10 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | Circ | meas | m_spleen_EQI_4 | r_Circumference | unilateral | 125100\ C99-5100-137\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-137, ShearWave Elastography sample 4 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | Circ | meas | m_spleen_EQI_5 | r_Circumference | unilateral | 125100\ C99-5100-138\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-138, ShearWave Elastography sample 5 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | Circ | meas | m_spleen_EQI_6 | r_Circumference | unilateral | 125100\ C99-5100-139\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-139, ShearWave Elastography sample 6 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | Circ | meas | m_spleen_EQI_7 | r_Circumference | unilateral | 125100\ C99-5100-140\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-140, ShearWave Elastography sample 7 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | Circ | meas | m_spleen_EQI_8 | r_Circumference | unilateral | 125100\ C99-5100-141\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-141, ShearWave Elastography sample 8 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | Circ | meas | m_spleen_EQI_9 | r_Circumference | unilateral | 125100\ C99-5100-142\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-142, ShearWave Elastography sample 9 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | Circ | meas | m_spleen_EQI_10 | r_Circumference | unilateral | 125100\ C99-5100-143\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-143, ShearWave Elastography sample 10 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | Area | meas | m_spleen_EQI_4 | r_Area | unilateral | 125100\ C99-5100-147\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-147, ShearWave Elastography sample 4 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | Area | meas | m_spleen_EQI_5 | r_Area | unilateral | 125100\ C99-5100-148\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-148, ShearWave Elastography sample 5 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | Area | meas | m_spleen_EQI_6 | r_Area | unilateral | 125100\ C99-5100-149\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-149, ShearWave Elastography sample 6 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | Area | meas | m_spleen_EQI_7 | r_Area | unilateral | 125100\ C99-5100-150\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-150, ShearWave Elastography sample 7 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | Area | meas | m_spleen_EQI_8 | r_Area | unilateral | 125100\ C99-5100-151\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-151, ShearWave Elastography sample 8 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | Area | meas | m_spleen_EQI_9 | r_Area | unilateral | 125100\ C99-5100-152\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-152, ShearWave Elastography sample 9 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | Area | meas | m_spleen_EQI_10 | r_Area | unilateral | 125100\ C99-5100-153\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-153, ShearWave Elastography sample 10 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen H | Dist | meas | m_splnHt | r_Distance | unilateral | 125100\C7470-01\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C7470-01, Height |  | UCUM, mm, mm | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | EQI Avg | meas | m_spleen_EQI_1 | r_AvgStiffness_p | unilateral | 125100\C99-5100-001\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-001, ShearWave Elastography Stiffness sample 1 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | EQI Avg | meas | m_spleen_EQI_2 | r_AvgStiffness_p | unilateral | 125100\C99-5100-002\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-002, ShearWave Elastography Stiffness sample 2 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | EQI Avg | meas | m_spleen_EQI_3 | r_AvgStiffness_p | unilateral | 125100\C99-5100-003\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-003, ShearWave Elastography Stiffness sample 3 Average |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | EQI Std | meas | m_spleen_EQI_1 | r_StdStiffness_p | unilateral | 125100\C99-5100-044\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-044, ShearWave Elastography Stiffness sample 1 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | EQI Std | meas | m_spleen_EQI_2 | r_StdStiffness_p | unilateral | 125100\C99-5100-045\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-045, ShearWave Elastography Stiffness sample 2 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | EQI Std | meas | m_spleen_EQI_3 | r_StdStiffness_p | unilateral | 125100\C99-5100-046\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-046, ShearWave Elastography Stiffness sample 3 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | EQI Med | meas | m_spleen_EQI_1 | r_MedStiffness_p | unilateral | 125100\C99-5100-064\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-064, ShearWave Elastography Stiffness sample 1 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | EQI Med | meas | m_spleen_EQI_2 | r_MedStiffness_p | unilateral | 125100\C99-5100-065\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-065, ShearWave Elastography Stiffness sample 2 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | EQI Med | meas | m_spleen_EQI_3 | r_MedStiffness_p | unilateral | 125100\C99-5100-066\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-066, ShearWave Elastography Stiffness sample 3 Median |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | EQI IQR | meas | m_spleen_EQI_1 | r_IqrStiffness_p | unilateral | 125100\C99-5100-084\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-084, ShearWave Elastography Stiffness sample 1 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | EQI IQR | meas | m_spleen_EQI_2 | r_IqrStiffness_p | unilateral | 125100\C99-5100-085\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-085, ShearWave Elastography Stiffness sample 2 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | EQI IQR | meas | m_spleen_EQI_3 | r_IqrStiffness_p | unilateral | 125100\C99-5100-086\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-086, ShearWave Elastography Stiffness sample 3 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | EQI IQR/Med | meas | m_spleen_EQI_1 | r_IqrMedStiffness_p | unilateral | 125100\C99-5100-104\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-104, ShearWave Elasto Stiffness sample 1 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | EQI IQR/Med | meas | m_spleen_EQI_2 | r_IqrMedStiffness_p | unilateral | 125100\C99-5100-105\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-105, ShearWave Elasto Stiffness sample 2 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | EQI IQR/Med | meas | m_spleen_EQI_3 | r_IqrMedStiffness_p | unilateral | 125100\C99-5100-106\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-106, ShearWave Elasto Stiffness sample 3 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | Conf. Threshold | meas | m_spleen_EQI_1 | r_CFStiffness | unilateral | 125100\C99-5100-124\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-124, ShearWave Elastography Stiffness sample 1 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | Conf. Threshold | meas | m_spleen_EQI_2 | r_CFStiffness | unilateral | 125100\C99-5100-125\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-125, ShearWave Elastography Stiffness sample 2 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | Conf. Threshold | meas | m_spleen_EQI_3 | r_CFStiffness | unilateral | 125100\C99-5100-126\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-126, ShearWave Elastography Stiffness sample 3 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | Circ | meas | m_spleen_EQI_1 | r_Circumference | unilateral | 125100\C99-5100-134\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-134, ShearWave Elastography sample 1 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | Circ | meas | m_spleen_EQI_2 | r_Circumference | unilateral | 125100\C99-5100-135\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-135, ShearWave Elastography sample 2 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | Circ | meas | m_spleen_EQI_3 | r_Circumference | unilateral | 125100\C99-5100-136\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-136, ShearWave Elastography sample 3 Circumference |  | UCUM, mm, mm | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | Area | meas | m_spleen_EQI_1 | r_Area | unilateral | 125100\C99-5100-144\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-144, ShearWave Elastography sample 1 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | Area | meas | m_spleen_EQI_2 | r_Area | unilateral | 125100\C99-5100-145\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-145, ShearWave Elastography sample 2 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | Area | meas | m_spleen_EQI_3 | r_Area | unilateral | 125100\C99-5100-146\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-146, ShearWave Elastography sample 3 Area |  | UCUM, mm2, mm2 | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | EQI Min | meas | m_spleen_EQI_1 | r_MinStiffness_p | unilateral | 125100\C99-5100-154\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-154, ShearWave Elastography Stiffness sample 1 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | EQI Min | meas | m_spleen_EQI_2 | r_MinStiffness_p | unilateral | 125100\C99-5100-155\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-155, ShearWave Elastography Stiffness sample 2 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | EQI Min | meas | m_spleen_EQI_3 | r_MinStiffness_p | unilateral | 125100\C99-5100-156\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-156, ShearWave Elastography Stiffness sample 3 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | EQI Min | meas | m_spleen_EQI_4 | r_MinStiffness_p | unilateral | 125100\C99-5100-157\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-157, ShearWave Elastography Stiffness sample 4 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | EQI Min | meas | m_spleen_EQI_5 | r_MinStiffness_p | unilateral | 125100\C99-5100-158\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-158, ShearWave Elastography Stiffness sample 5 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | EQI Min | meas | m_spleen_EQI_6 | r_MinStiffness_p | unilateral | 125100\C99-5100-159\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-159, ShearWave Elastography Stiffness sample 6 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | EQI Min | meas | m_spleen_EQI_7 | r_MinStiffness_p | unilateral | 125100\C99-5100-160\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-160, ShearWave Elastography Stiffness sample 7 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | EQI Min | meas | m_spleen_EQI_8 | r_MinStiffness_p | unilateral | 125100\C99-5100-161\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-161, ShearWave Elastography Stiffness sample 8 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | EQI Min | meas | m_spleen_EQI_9 | r_MinStiffness_p | unilateral | 125100\C99-5100-162\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-162, ShearWave Elastography Stiffness sample 9 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | EQI Min | meas | m_spleen_EQI_10 | r_MinStiffness_p | unilateral | 125100\C99-5100-163\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-163, ShearWave Elastography Stiffness sample 10 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | EQI Min Vel | meas | m_spleen_EQI_1 | r_MinStiffness_v | unilateral | 125100\C99-5100-164\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-164, ShearWave Elastography Velocity sample 1 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | EQI Min Vel | meas | m_spleen_EQI_2 | r_MinStiffness_v | unilateral | 125100\C99-5100-165\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-165, ShearWave Elastography Velocity sample 2 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | EQI Min Vel | meas | m_spleen_EQI_3 | r_MinStiffness_v | unilateral | 125100\C99-5100-166\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-166, ShearWave Elastography Velocity sample 3 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | EQI Min Vel | meas | m_spleen_EQI_4 | r_MinStiffness_v | unilateral | 125100\C99-5100-167\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-167, ShearWave Elastography Velocity sample 4 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | EQI Min Vel | meas | m_spleen_EQI_5 | r_MinStiffness_v | unilateral | 125100\C99-5100-168\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-168, ShearWave Elastography Velocity sample 5 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | EQI Min Vel | meas | m_spleen_EQI_6 | r_MinStiffness_v | unilateral | 125100\C99-5100-169\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-169, ShearWave Elastography Velocity sample 6 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | EQI Min Vel | meas | m_spleen_EQI_7 | r_MinStiffness_v | unilateral | 125100\C99-5100-170\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-170, ShearWave Elastography Velocity sample 7 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | EQI Min Vel | meas | m_spleen_EQI_8 | r_MinStiffness_v | unilateral | 125100\C99-5100-171 \T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-171 , ShearWave Elastography Velocity sample 8 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | EQI Min Vel | meas | m_spleen_EQI_9 | r_MinStiffness_v | unilateral | 125100\C99-5100-172\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-172, ShearWave Elastography Velocity sample 9 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | EQI Min Vel | meas | m_spleen_EQI_10 | r_MinStiffness_v | unilateral | 125100\C99-5100-173\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-173, ShearWave Elastography Velocity sample 10 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | EQI Max | meas | m_spleen_EQI_1 | r_MaxStiffness_p | unilateral | 125100\C99-5100-174\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-174, ShearWave Elastography Stiffness sample 1 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | EQI Max | meas | m_spleen_EQI_2 | r_MaxStiffness_p | unilateral | 125100\C99-5100-175\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-175, ShearWave Elastography Stiffness sample 2 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | EQI Max | meas | m_spleen_EQI_3 | r_MaxStiffness_p | unilateral | 125100\C99-5100-176\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-176, ShearWave Elastography Stiffness sample 3 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | EQI Max | meas | m_spleen_EQI_4 | r_MaxStiffness_p | unilateral | 125100\C99-5100-177\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-177, ShearWave Elastography Stiffness sample 4 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | EQI Max | meas | m_spleen_EQI_5 | r_MaxStiffness_p | unilateral | 125100\C99-5100-178\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-178, ShearWave Elastography Stiffness sample 5 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | EQI Max | meas | m_spleen_EQI_6 | r_MaxStiffness_p | unilateral | 125100\C99-5100-179\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-179, ShearWave Elastography Stiffness sample 6 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | EQI Max | meas | m_spleen_EQI_7 | r_MaxStiffness_p | unilateral | 125100\C99-5100-180\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-180, ShearWave Elastography Stiffness sample 7 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | EQI Max | meas | m_spleen_EQI_8 | r_MaxStiffness_p | unilateral | 125100\C99-5100-181\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-181, ShearWave Elastography Stiffness sample 8 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | EQI Max | meas | m_spleen_EQI_9 | r_MaxStiffness_p | unilateral | 125100\C99-5100-182\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-182, ShearWave Elastography Stiffness sample 9 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | EQI Max | meas | m_spleen_EQI_10 | r_MaxStiffness_p | unilateral | 125100\C99-5100-183\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-183, ShearWave Elastography Stiffness sample 10 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 1 | EQI Max Vel | meas | m_spleen_EQI_1 | r_MaxStiffness_v | unilateral | 125100\C99-5100-184\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-184, ShearWave Elastography Velocity sample 1 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 2 | EQI Max Vel | meas | m_spleen_EQI_2 | r_MaxStiffness_v | unilateral | 125100\C99-5100-185\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-185, ShearWave Elastography Velocity sample 2 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 3 | EQI Max Vel | meas | m_spleen_EQI_3 | r_MaxStiffness_v | unilateral | 125100\C99-5100-186\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-186, ShearWave Elastography Velocity sample 3 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 4 | EQI Max Vel | meas | m_spleen_EQI_4 | r_MaxStiffness_v | unilateral | 125100\C99-5100-187\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-187, ShearWave Elastography Velocity sample 4 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 5 | EQI Max Vel | meas | m_spleen_EQI_5 | r_MaxStiffness_v | unilateral | 125100\C99-5100-188\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-188, ShearWave Elastography Velocity sample 5 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 6 | EQI Max Vel | meas | m_spleen_EQI_6 | r_MaxStiffness_v | unilateral | 125100\C99-5100-189\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-189, ShearWave Elastography Velocity sample 6 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 7 | EQI Max Vel | meas | m_spleen_EQI_7 | r_MaxStiffness_v | unilateral | 125100\C99-5100-190\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-190, ShearWave Elastography Velocity sample 7 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 8 | EQI Max Vel | meas | m_spleen_EQI_8 | r_MaxStiffness_v | unilateral | 125100\C99-5100-191 \T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-191 , ShearWave Elastography Velocity sample 8 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 9 | EQI Max Vel | meas | m_spleen_EQI_9 | r_MaxStiffness_v | unilateral | 125100\C99-5100-192\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-192, ShearWave Elastography Velocity sample 9 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI 10 | EQI Max Vel | meas | m_spleen_EQI_10 | r_MaxStiffness_v | unilateral | 125100\C99-5100-193\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-193, ShearWave Elastography Velocity sample 10 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen W | Dist | meas | m_splnWid | r_Distance | unilateral | 125100\G-A220\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | SNM3, G-A220, Width |  | UCUM, mm, mm | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen L | Dist | meas | m_splnLen | r_Distance | unilateral | 125100\G-A22A\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | SRT, G-A22A, Length |  | UCUM, mm, mm | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | EQI Avg | meas | m_pan_EQI_4 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-004\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-004, ShearWave Elastography Stiffness sample 4 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | EQI Avg | meas | m_pan_EQI_5 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-005\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-005, ShearWave Elastography Stiffness sample 5 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | EQI Avg | meas | m_pan_EQI_6 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-006\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-006, ShearWave Elastography Stiffness sample 6 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | EQI Avg | meas | m_pan_EQI_7 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-007\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-007, ShearWave Elastography Stiffness sample 7 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | EQI Avg | meas | m_pan_EQI_8 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-008\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-008, ShearWave Elastography Stiffness sample 8 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | EQI Avg | meas | m_pan_EQI_9 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-009\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-009, ShearWave Elastography Stiffness sample 9 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | EQI Avg | meas | m_pan_EQI_10 | r_AvgStiffness_p | unilateral | 125100\ C99-5100-010\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-010, ShearWave Elastography Stiffness sample 10 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | EQI Avg Vel | meas | m_pan_EQI_1 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-011\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-011, ShearWave Elastography Velocity sample 1 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | EQI Avg Vel | meas | m_pan_EQI_2 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-012\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-012, ShearWave Elastography Velocity sample 2 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | EQI Avg Vel | meas | m_pan_EQI_3 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-013\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-013, ShearWave Elastography Velocity sample 3 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | EQI Avg Vel | meas | m_pan_EQI_4 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-014\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-014, ShearWave Elastography Velocity sample 4 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | EQI Avg Vel | meas | m_pan_EQI_5 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-015\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-015, ShearWave Elastography Velocity sample 5 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | EQI Avg Vel | meas | m_pan_EQI_6 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-016\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-016, ShearWave Elastography Velocity sample 6 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | EQI Avg Vel | meas | m_pan_EQI_7 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-017\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-017, ShearWave Elastography Velocity sample 7 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | EQI Avg Vel | meas | m_pan_EQI_8 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-018 \T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-018 , ShearWave Elastography Velocity sample 8 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | EQI Avg Vel | meas | m_pan_EQI_9 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-019\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-019, ShearWave Elastography Velocity sample 9 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | EQI Avg Vel | meas | m_pan_EQI_10 | r_AvgStiffness_v | unilateral | 125100\ C99-5100-020\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-020, ShearWave Elastography Velocity sample 10 Average |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | EQI Std | meas | m_pan_EQI_4 | r_StdStiffness_p | unilateral | 125100\ C99-5100-047\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-047, ShearWave Elastography Stiffness sample 4 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | EQI Std | meas | m_pan_EQI_5 | r_StdStiffness_p | unilateral | 125100\ C99-5100-048\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-048, ShearWave Elastography Stiffness sample 5 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | EQI Std | meas | m_pan_EQI_6 | r_StdStiffness_p | unilateral | 125100\ C99-5100-049\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-049, ShearWave Elastography Stiffness sample 6 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | EQI Std | meas | m_pan_EQI_7 | r_StdStiffness_p | unilateral | 125100\ C99-5100-050\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-050, ShearWave Elastography Stiffness sample 7 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | EQI Std | meas | m_pan_EQI_8 | r_StdStiffness_p | unilateral | 125100\ C99-5100-051\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-051, ShearWave Elastography Stiffness sample 8 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | EQI Std | meas | m_pan_EQI_9 | r_StdStiffness_p | unilateral | 125100\ C99-5100-052\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-052, ShearWave Elastography Stiffness sample 9 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | EQI Std | meas | m_pan_EQI_10 | r_StdStiffness_p | unilateral | 125100\ C99-5100-053\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-053, ShearWave Elastography Stiffness sample 10 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | EQI Std Vel | meas | m_pan_EQI_1 | r_StdStiffness_v | unilateral | 125100\ C99-5100-054\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-054, ShearWave Elastography Velocity sample 1 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | EQI Std Vel | meas | m_pan_EQI_2 | r_StdStiffness_v | unilateral | 125100\ C99-5100-055\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-055, ShearWave Elastography Velocity sample 2 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | EQI Std Vel | meas | m_pan_EQI_3 | r_StdStiffness_v | unilateral | 125100\ C99-5100-056\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-056, ShearWave Elastography Velocity sample 3 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | EQI Std Vel | meas | m_pan_EQI_4 | r_StdStiffness_v | unilateral | 125100\ C99-5100-057\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-057, ShearWave Elastography Velocity sample 4 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | EQI Std Vel | meas | m_pan_EQI_5 | r_StdStiffness_v | unilateral | 125100\ C99-5100-058\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-058, ShearWave Elastography Velocity sample 5 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | EQI Std Vel | meas | m_pan_EQI_6 | r_StdStiffness_v | unilateral | 125100\ C99-5100-059\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-059, ShearWave Elastography Velocity sample 6 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | EQI Std Vel | meas | m_pan_EQI_7 | r_StdStiffness_v | unilateral | 125100\ C99-5100-060\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-060, ShearWave Elastography Velocity sample 7 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | EQI Std Vel | meas | m_pan_EQI_8 | r_StdStiffness_v | unilateral | 125100\ C99-5100-061\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-061, ShearWave Elastography Velocity sample 8 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | EQI Std Vel | meas | m_pan_EQI_9 | r_StdStiffness_v | unilateral | 125100\ C99-5100-062\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-062, ShearWave Elastography Velocity sample 9 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | EQI Std Vel | meas | m_pan_EQI_10 | r_StdStiffness_v | unilateral | 125100\ C99-5100-063\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-063, ShearWave Elastography Velocity sample 10 Standard Deviation |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | EQI Med | meas | m_pan_EQI_4 | r_MedStiffness_p | unilateral | 125100\ C99-5100-067\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-067, ShearWave Elastography Stiffness sample 4 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | EQI Med | meas | m_pan_EQI_5 | r_MedStiffness_p | unilateral | 125100\ C99-5100-068\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-068, ShearWave Elastography Stiffness sample 5 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | EQI Med | meas | m_pan_EQI_6 | r_MedStiffness_p | unilateral | 125100\ C99-5100-069\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-069, ShearWave Elastography Stiffness sample 6 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | EQI Med | meas | m_pan_EQI_7 | r_MedStiffness_p | unilateral | 125100\ C99-5100-070\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-070, ShearWave Elastography Stiffness sample 7 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | EQI Med | meas | m_pan_EQI_8 | r_MedStiffness_p | unilateral | 125100\ C99-5100-071\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-071, ShearWave Elastography Stiffness sample 8 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | EQI Med | meas | m_pan_EQI_9 | r_MedStiffness_p | unilateral | 125100\ C99-5100-072\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-072, ShearWave Elastography Stiffness sample 9 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | EQI Med | meas | m_pan_EQI_10 | r_MedStiffness_p | unilateral | 125100\ C99-5100-073\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-073, ShearWave Elastography Stiffness sample 10 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | EQI Med Vel | meas | m_pan_EQI_1 | r_MedStiffness_v | unilateral | 125100\ C99-5100-074\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-074, ShearWave Elastography Velocity sample 1 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | EQI Med Vel | meas | m_pan_EQI_2 | r_MedStiffness_v | unilateral | 125100\ C99-5100-075\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-075, ShearWave Elastography Velocity sample 2 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | EQI Med Vel | meas | m_pan_EQI_3 | r_MedStiffness_v | unilateral | 125100\ C99-5100-076\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-076, ShearWave Elastography Velocity sample 3 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | EQI Med Vel | meas | m_pan_EQI_4 | r_MedStiffness_v | unilateral | 125100\ C99-5100-077\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-077, ShearWave Elastography Velocity sample 4 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | EQI Med Vel | meas | m_pan_EQI_5 | r_MedStiffness_v | unilateral | 125100\ C99-5100-078\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-078, ShearWave Elastography Velocity sample 5 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | EQI Med Vel | meas | m_pan_EQI_6 | r_MedStiffness_v | unilateral | 125100\ C99-5100-079\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-079, ShearWave Elastography Velocity sample 6 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | EQI Med Vel | meas | m_pan_EQI_7 | r_MedStiffness_v | unilateral | 125100\ C99-5100-080\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-080, ShearWave Elastography Velocity sample 7 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | EQI Med Vel | meas | m_pan_EQI_8 | r_MedStiffness_v | unilateral | 125100\ C99-5100-081\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-081, ShearWave Elastography Velocity sample 8 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | EQI Med Vel | meas | m_pan_EQI_9 | r_MedStiffness_v | unilateral | 125100\ C99-5100-082\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-082, ShearWave Elastography Velocity sample 9 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | EQI Med Vel | meas | m_pan_EQI_10 | r_MedStiffness_v | unilateral | 125100\ C99-5100-083\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-083, ShearWave Elastography Velocity sample 10 Median |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | EQI IQR | meas | m_pan_EQI_4 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-087\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-087, ShearWave Elastography Stiffness sample 4 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | EQI IQR | meas | m_pan_EQI_5 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-088\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-088, ShearWave Elastography Stiffness sample 5 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | EQI IQR | meas | m_pan_EQI_6 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-089\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-089, ShearWave Elastography Stiffness sample 6 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | EQI IQR | meas | m_pan_EQI_7 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-090\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-090, ShearWave Elastography Stiffness sample 7 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | EQI IQR | meas | m_pan_EQI_8 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-091\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-091, ShearWave Elastography Stiffness sample 8 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | EQI IQR | meas | m_pan_EQI_9 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-092\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-092, ShearWave Elastography Stiffness sample 9 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | EQI IQR | meas | m_pan_EQI_10 | r_IqrStiffness_p | unilateral | 125100\ C99-5100-093\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-093, ShearWave Elastography Stiffness sample 10 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | EQI IQR Vel | meas | m_pan_EQI_1 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-094\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-094, ShearWave Elastography Velocity sample 1 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | EQI IQR Vel | meas | m_pan_EQI_2 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-095\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-095, ShearWave Elastography Velocity sample 2 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | EQI IQR Vel | meas | m_pan_EQI_3 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-096\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-096, ShearWave Elastography Velocity sample 3 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | EQI IQR Vel | meas | m_pan_EQI_4 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-097\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-097, ShearWave Elastography Velocity sample 4 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | EQI IQR Vel | meas | m_pan_EQI_5 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-098\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-098, ShearWave Elastography Velocity sample 5 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | EQI IQR Vel | meas | m_pan_EQI_6 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-099\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-099, ShearWave Elastography Velocity sample 6 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | EQI IQR Vel | meas | m_pan_EQI_7 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-100\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-100, ShearWave Elastography Velocity sample 7 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | EQI IQR Vel | meas | m_pan_EQI_8 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-101\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-101, ShearWave Elastography Velocity sample 8 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | EQI IQR Vel | meas | m_pan_EQI_9 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-102\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-102, ShearWave Elastography Velocity sample 9 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | EQI IQR Vel | meas | m_pan_EQI_10 | r_IqrStiffness_v | unilateral | 125100\ C99-5100-103\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-103, ShearWave Elastography Velocity sample 10 Interquartile Range |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | EQI IQR/Med | meas | m_pan_EQI_4 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-107\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-107, ShearWave Elasto Stiffness sample 4 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | EQI IQR/Med | meas | m_pan_EQI_5 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-108\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-108, ShearWave Elasto Stiffness sample 5 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | EQI IQR/Med | meas | m_pan_EQI_6 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-109\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-109, ShearWave Elasto Stiffness sample 6 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | EQI IQR/Med | meas | m_pan_EQI_7 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-110\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-110, ShearWave Elasto Stiffness sample 7 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | EQI IQR/Med | meas | m_pan_EQI_8 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-111\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-111, ShearWave Elasto Stiffness sample 8 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | EQI IQR/Med | meas | m_pan_EQI_9 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-112\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-112, ShearWave Elasto Stiffness sample 9 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | EQI IQR/Med | meas | m_pan_EQI_10 | r_IqrMedStiffness_p | unilateral | 125100\ C99-5100-113\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-113, ShearWave Elasto Stiffness sample 10 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | EQI IQR/Med Vel | meas | m_pan_EQI_1 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-114\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-114, ShearWave Elasto Velocity sample 1 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | EQI IQR/Med Vel | meas | m_pan_EQI_2 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-115\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-115, ShearWave Elasto Velocity sample 2 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | EQI IQR/Med Vel | meas | m_pan_EQI_3 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-116\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-116, ShearWave Elasto Velocity sample 3 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | EQI IQR/Med Vel | meas | m_pan_EQI_4 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-117\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-117, ShearWave Elasto Velocity sample 4 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | EQI IQR/Med Vel | meas | m_pan_EQI_5 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-118\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-118, ShearWave Elasto Velocity sample 5 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | EQI IQR/Med Vel | meas | m_pan_EQI_6 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-119\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-119, ShearWave Elasto Velocity sample 6 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | EQI IQR/Med Vel | meas | m_pan_EQI_7 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-120\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-120, ShearWave Elasto Velocity sample 7 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | EQI IQR/Med Vel | meas | m_pan_EQI_8 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-121\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-121, ShearWave Elasto Velocity sample 8 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | EQI IQR/Med Vel | meas | m_pan_EQI_9 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-122\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-122, ShearWave Elasto Velocity sample 9 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | EQI IQR/Med Vel | meas | m_pan_EQI_10 | r_IqrMedStiffness_v | unilateral | 125100\ C99-5100-123\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-123, ShearWave Elasto Velocity sample 10 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | Conf. Threshold | meas | m_pan_EQI_4 | r_CFStiffness | unilateral | 125100\ C99-5100-127\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-127, ShearWave Elastography Stiffness sample 4 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | Conf. Threshold | meas | m_pan_EQI_5 | r_CFStiffness | unilateral | 125100\ C99-5100-128\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-128, ShearWave Elastography Stiffness sample 5 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | Conf. Threshold | meas | m_pan_EQI_6 | r_CFStiffness | unilateral | 125100\ C99-5100-129\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-129, ShearWave Elastography Stiffness sample 6 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | Conf. Threshold | meas | m_pan_EQI_7 | r_CFStiffness | unilateral | 125100\ C99-5100-130\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-130, ShearWave Elastography Stiffness sample 7 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | Conf. Threshold | meas | m_pan_EQI_8 | r_CFStiffness | unilateral | 125100\ C99-5100-131\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-131, ShearWave Elastography Stiffness sample 8 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | Conf. Threshold | meas | m_pan_EQI_9 | r_CFStiffness | unilateral | 125100\ C99-5100-132\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-132, ShearWave Elastography Stiffness sample 9 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | Conf. Threshold | meas | m_pan_EQI_10 | r_CFStiffness | unilateral | 125100\ C99-5100-133\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-133, ShearWave Elastography Stiffness sample 10 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | Circ | meas | m_pan_EQI_4 | r_Circumference | unilateral | 125100\ C99-5100-137\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-137, ShearWave Elastography sample 4 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | Circ | meas | m_pan_EQI_5 | r_Circumference | unilateral | 125100\ C99-5100-138\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-138, ShearWave Elastography sample 5 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | Circ | meas | m_pan_EQI_6 | r_Circumference | unilateral | 125100\ C99-5100-139\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-139, ShearWave Elastography sample 6 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | Circ | meas | m_pan_EQI_7 | r_Circumference | unilateral | 125100\ C99-5100-140\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-140, ShearWave Elastography sample 7 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | Circ | meas | m_pan_EQI_8 | r_Circumference | unilateral | 125100\ C99-5100-141\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-141, ShearWave Elastography sample 8 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | Circ | meas | m_pan_EQI_9 | r_Circumference | unilateral | 125100\ C99-5100-142\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-142, ShearWave Elastography sample 9 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | Circ | meas | m_pan_EQI_10 | r_Circumference | unilateral | 125100\ C99-5100-143\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-143, ShearWave Elastography sample 10 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | Area | meas | m_pan_EQI_4 | r_Area | unilateral | 125100\ C99-5100-147\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-147, ShearWave Elastography sample 4 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | Area | meas | m_pan_EQI_5 | r_Area | unilateral | 125100\ C99-5100-148\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-148, ShearWave Elastography sample 5 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | Area | meas | m_pan_EQI_6 | r_Area | unilateral | 125100\ C99-5100-149\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-149, ShearWave Elastography sample 6 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | Area | meas | m_pan_EQI_7 | r_Area | unilateral | 125100\ C99-5100-150\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-150, ShearWave Elastography sample 7 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | Area | meas | m_pan_EQI_8 | r_Area | unilateral | 125100\ C99-5100-151\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-151, ShearWave Elastography sample 8 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | Area | meas | m_pan_EQI_9 | r_Area | unilateral | 125100\ C99-5100-152\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-152, ShearWave Elastography sample 9 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | Area | meas | m_pan_EQI_10 | r_Area | unilateral | 125100\ C99-5100-153\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-153, ShearWave Elastography sample 10 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | EQI Avg | meas | m_pan_EQI_1 | r_AvgStiffness_p | unilateral | 125100\C99-5100-001\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-001, ShearWave Elastography Stiffness sample 1 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | EQI Avg | meas | m_pan_EQI_2 | r_AvgStiffness_p | unilateral | 125100\C99-5100-002\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-002, ShearWave Elastography Stiffness sample 2 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | EQI Avg | meas | m_pan_EQI_3 | r_AvgStiffness_p | unilateral | 125100\C99-5100-003\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-003, ShearWave Elastography Stiffness sample 3 Average |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | EQI Std | meas | m_pan_EQI_1 | r_StdStiffness_p | unilateral | 125100\C99-5100-044\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-044, ShearWave Elastography Stiffness sample 1 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | EQI Std | meas | m_pan_EQI_2 | r_StdStiffness_p | unilateral | 125100\C99-5100-045\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-045, ShearWave Elastography Stiffness sample 2 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | EQI Std | meas | m_pan_EQI_3 | r_StdStiffness_p | unilateral | 125100\C99-5100-046\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-046, ShearWave Elastography Stiffness sample 3 Standard Deviation |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | EQI Med | meas | m_pan_EQI_1 | r_MedStiffness_p | unilateral | 125100\C99-5100-064\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-064, ShearWave Elastography Stiffness sample 1 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | EQI Med | meas | m_pan_EQI_2 | r_MedStiffness_p | unilateral | 125100\C99-5100-065\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-065, ShearWave Elastography Stiffness sample 2 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | EQI Med | meas | m_pan_EQI_3 | r_MedStiffness_p | unilateral | 125100\C99-5100-066\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-066, ShearWave Elastography Stiffness sample 3 Median |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | EQI IQR | meas | m_pan_EQI_1 | r_IqrStiffness_p | unilateral | 125100\C99-5100-084\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-084, ShearWave Elastography Stiffness sample 1 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | EQI IQR | meas | m_pan_EQI_2 | r_IqrStiffness_p | unilateral | 125100\C99-5100-085\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-085, ShearWave Elastography Stiffness sample 2 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | EQI IQR | meas | m_pan_EQI_3 | r_IqrStiffness_p | unilateral | 125100\C99-5100-086\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-086, ShearWave Elastography Stiffness sample 3 Interquartile Range |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | EQI IQR/Med | meas | m_pan_EQI_1 | r_IqrMedStiffness_p | unilateral | 125100\C99-5100-104\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-104, ShearWave Elasto Stiffness sample 1 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | EQI IQR/Med | meas | m_pan_EQI_2 | r_IqrMedStiffness_p | unilateral | 125100\C99-5100-105\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-105, ShearWave Elasto Stiffness sample 2 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | EQI IQR/Med | meas | m_pan_EQI_3 | r_IqrMedStiffness_p | unilateral | 125100\C99-5100-106\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-106, ShearWave Elasto Stiffness sample 3 Interquartile Range/Median |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | Conf. Threshold | meas | m_pan_EQI_1 | r_CFStiffness | unilateral | 125100\C99-5100-124\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-124, ShearWave Elastography Stiffness sample 1 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | Conf. Threshold | meas | m_pan_EQI_2 | r_CFStiffness | unilateral | 125100\C99-5100-125\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-125, ShearWave Elastography Stiffness sample 2 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | Conf. Threshold | meas | m_pan_EQI_3 | r_CFStiffness | unilateral | 125100\C99-5100-126\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-126, ShearWave Elastography Stiffness sample 3 Confidence Threshold |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | Circ | meas | m_pan_EQI_1 | r_Circumference | unilateral | 125100\C99-5100-134\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-134, ShearWave Elastography sample 1 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | Circ | meas | m_pan_EQI_2 | r_Circumference | unilateral | 125100\C99-5100-135\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-135, ShearWave Elastography sample 2 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | Circ | meas | m_pan_EQI_3 | r_Circumference | unilateral | 125100\C99-5100-136\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-136, ShearWave Elastography sample 3 Circumference |  | UCUM, mm, mm | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | Area | meas | m_pan_EQI_1 | r_Area | unilateral | 125100\C99-5100-144\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-144, ShearWave Elastography sample 1 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | Area | meas | m_pan_EQI_2 | r_Area | unilateral | 125100\C99-5100-145\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-145, ShearWave Elastography sample 2 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | Area | meas | m_pan_EQI_3 | r_Area | unilateral | 125100\C99-5100-146\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-146, ShearWave Elastography sample 3 Area |  | UCUM, mm2, mm2 | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | EQI Min | meas | m_pan_EQI_1 | r_MinStiffness_p | unilateral | 125100\C99-5100-154\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-154, ShearWave Elastography Stiffness sample 1 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | EQI Min | meas | m_pan_EQI_2 | r_MinStiffness_p | unilateral | 125100\C99-5100-155\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-155, ShearWave Elastography Stiffness sample 2 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | EQI Min | meas | m_pan_EQI_3 | r_MinStiffness_p | unilateral | 125100\C99-5100-156\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-156, ShearWave Elastography Stiffness sample 3 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | EQI Min | meas | m_pan_EQI_4 | r_MinStiffness_p | unilateral | 125100\C99-5100-157\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-157, ShearWave Elastography Stiffness sample 4 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | EQI Min | meas | m_pan_EQI_5 | r_MinStiffness_p | unilateral | 125100\C99-5100-158\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-158, ShearWave Elastography Stiffness sample 5 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | EQI Min | meas | m_pan_EQI_6 | r_MinStiffness_p | unilateral | 125100\C99-5100-159\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-159, ShearWave Elastography Stiffness sample 6 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | EQI Min | meas | m_pan_EQI_7 | r_MinStiffness_p | unilateral | 125100\C99-5100-160\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-160, ShearWave Elastography Stiffness sample 7 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | EQI Min | meas | m_pan_EQI_8 | r_MinStiffness_p | unilateral | 125100\C99-5100-161\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-161, ShearWave Elastography Stiffness sample 8 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | EQI Min | meas | m_pan_EQI_9 | r_MinStiffness_p | unilateral | 125100\C99-5100-162\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-162, ShearWave Elastography Stiffness sample 9 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | EQI Min | meas | m_pan_EQI_10 | r_MinStiffness_p | unilateral | 125100\C99-5100-163\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-163, ShearWave Elastography Stiffness sample 10 Minimum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | EQI Min Vel | meas | m_pan_EQI_1 | r_MinStiffness_v | unilateral | 125100\C99-5100-164\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-164, ShearWave Elastography Velocity sample 1 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | EQI Min Vel | meas | m_pan_EQI_2 | r_MinStiffness_v | unilateral | 125100\C99-5100-165\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-165, ShearWave Elastography Velocity sample 2 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | EQI Min Vel | meas | m_pan_EQI_3 | r_MinStiffness_v | unilateral | 125100\C99-5100-166\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-166, ShearWave Elastography Velocity sample 3 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | EQI Min Vel | meas | m_pan_EQI_4 | r_MinStiffness_v | unilateral | 125100\C99-5100-167\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-167, ShearWave Elastography Velocity sample 4 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | EQI Min Vel | meas | m_pan_EQI_5 | r_MinStiffness_v | unilateral | 125100\C99-5100-168\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-168, ShearWave Elastography Velocity sample 5 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | EQI Min Vel | meas | m_pan_EQI_6 | r_MinStiffness_v | unilateral | 125100\C99-5100-169\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-169, ShearWave Elastography Velocity sample 6 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | EQI Min Vel | meas | m_pan_EQI_7 | r_MinStiffness_v | unilateral | 125100\C99-5100-170\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-170, ShearWave Elastography Velocity sample 7 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | EQI Min Vel | meas | m_pan_EQI_8 | r_MinStiffness_v | unilateral | 125100\C99-5100-171 \T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-171 , ShearWave Elastography Velocity sample 8 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | EQI Min Vel | meas | m_pan_EQI_9 | r_MinStiffness_v | unilateral | 125100\C99-5100-172\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-172, ShearWave Elastography Velocity sample 9 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | EQI Min Vel | meas | m_pan_EQI_10 | r_MinStiffness_v | unilateral | 125100\C99-5100-173\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-173, ShearWave Elastography Velocity sample 10 Minimum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | EQI Max | meas | m_pan_EQI_1 | r_MaxStiffness_p | unilateral | 125100\C99-5100-174\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-174, ShearWave Elastography Stiffness sample 1 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | EQI Max | meas | m_pan_EQI_2 | r_MaxStiffness_p | unilateral | 125100\C99-5100-175\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-175, ShearWave Elastography Stiffness sample 2 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | EQI Max | meas | m_pan_EQI_3 | r_MaxStiffness_p | unilateral | 125100\C99-5100-176\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-176, ShearWave Elastography Stiffness sample 3 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | EQI Max | meas | m_pan_EQI_4 | r_MaxStiffness_p | unilateral | 125100\C99-5100-177\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-177, ShearWave Elastography Stiffness sample 4 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | EQI Max | meas | m_pan_EQI_5 | r_MaxStiffness_p | unilateral | 125100\C99-5100-178\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-178, ShearWave Elastography Stiffness sample 5 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | EQI Max | meas | m_pan_EQI_6 | r_MaxStiffness_p | unilateral | 125100\C99-5100-179\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-179, ShearWave Elastography Stiffness sample 6 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | EQI Max | meas | m_pan_EQI_7 | r_MaxStiffness_p | unilateral | 125100\C99-5100-180\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-180, ShearWave Elastography Stiffness sample 7 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | EQI Max | meas | m_pan_EQI_8 | r_MaxStiffness_p | unilateral | 125100\C99-5100-181\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-181, ShearWave Elastography Stiffness sample 8 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | EQI Max | meas | m_pan_EQI_9 | r_MaxStiffness_p | unilateral | 125100\C99-5100-182\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-182, ShearWave Elastography Stiffness sample 9 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | EQI Max | meas | m_pan_EQI_10 | r_MaxStiffness_p | unilateral | 125100\C99-5100-183\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-183, ShearWave Elastography Stiffness sample 10 Maximum |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 1 | EQI Max Vel | meas | m_pan_EQI_1 | r_MaxStiffness_v | unilateral | 125100\C99-5100-184\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-184, ShearWave Elastography Velocity sample 1 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 2 | EQI Max Vel | meas | m_pan_EQI_2 | r_MaxStiffness_v | unilateral | 125100\C99-5100-185\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-185, ShearWave Elastography Velocity sample 2 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 3 | EQI Max Vel | meas | m_pan_EQI_3 | r_MaxStiffness_v | unilateral | 125100\C99-5100-186\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-186, ShearWave Elastography Velocity sample 3 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 4 | EQI Max Vel | meas | m_pan_EQI_4 | r_MaxStiffness_v | unilateral | 125100\C99-5100-187\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-187, ShearWave Elastography Velocity sample 4 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 5 | EQI Max Vel | meas | m_pan_EQI_5 | r_MaxStiffness_v | unilateral | 125100\C99-5100-188\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-188, ShearWave Elastography Velocity sample 5 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 6 | EQI Max Vel | meas | m_pan_EQI_6 | r_MaxStiffness_v | unilateral | 125100\C99-5100-189\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-189, ShearWave Elastography Velocity sample 6 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 7 | EQI Max Vel | meas | m_pan_EQI_7 | r_MaxStiffness_v | unilateral | 125100\C99-5100-190\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-190, ShearWave Elastography Velocity sample 7 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 8 | EQI Max Vel | meas | m_pan_EQI_8 | r_MaxStiffness_v | unilateral | 125100\C99-5100-191 \T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-191 , ShearWave Elastography Velocity sample 8 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 9 | EQI Max Vel | meas | m_pan_EQI_9 | r_MaxStiffness_v | unilateral | 125100\C99-5100-192\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-192, ShearWave Elastography Velocity sample 9 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI 10 | EQI Max Vel | meas | m_pan_EQI_10 | r_MaxStiffness_v | unilateral | 125100\C99-5100-193\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-193, ShearWave Elastography Velocity sample 10 Maximum |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | TCD | MCA | Vel | meas | m_tcd_MCA | r_Vel | bilateral | 125100\C12105-01\C12221-03\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12105-01, MCA/Dist ICA Ratio Numerator
Vascular Ultrasound Procedure Report | TCD | ICA | Vel | meas | m_tcd_ICA | r_Vel | bilateral | 125100\C12105-02\C12221-03\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12105-02, MCA/Dist ICA Ratio Denominator
Vascular Ultrasound Procedure Report | TCD | Vertebral A | ? | meas | m_tcd_Vertebral_A | r_DopAngle | bilateral | 125100\C12105-04\C12122-01\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | SV Depth | meas | m_tcd_Vertebral_A | r_Svd | bilateral | 125100\C12105-04\C12221-02\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | Mean (TCD) | meas | m_tcd_Vertebral_A | r_TcdMean | bilateral | 125100\C12105-04\C12221-03\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | PI (TCD) | meas | m_tcd_Vertebral_A | r_TcdPI | bilateral | 125100\C12105-04\C12221-04\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | EDV | meas | m_tcd_Vertebral_A | r_EDV | bilateral | 125100\11653-3\C12105-04\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | MDV | meas | m_tcd_Vertebral_A | r_MDV | bilateral | 125100\11665-7\C12105-04\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | TAPV | meas | m_tcd_Vertebral_A | r_TAPV | bilateral | 125100\11692-1\C12105-04\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | PSV | meas | m_tcd_Vertebral_A | r_PSV | bilateral | 125100\11726-7\C12105-04\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | PI | meas | m_tcd_Vertebral_A | r_PI | bilateral | 125100\12008-9\C12105-04\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | RI | meas | m_tcd_Vertebral_A | r_RI | bilateral | 125100\12023-8\C12105-04\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | S/D | meas | m_tcd_Vertebral_A | r_SD | bilateral | 125100\12144-2\C12105-04\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | AI | meas | m_tcd_Vertebral_A | r_AccelIndex | bilateral | 125100\20167-3\C12105-04\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | AT | meas | m_tcd_Vertebral_A | r_AccelTime | bilateral | 125100\20168-1\C12105-04\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | DT | meas | m_tcd_Vertebral_A | r_DecelTime | bilateral | 125100\20217-6\C12105-04\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | PG | meas | m_tcd_Vertebral_A | r_PG | bilateral | 125100\20247-3\C12105-04\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | TAMV | meas | m_tcd_Vertebral_A | r_TAMV | bilateral | 125100\20352-1\C12105-04\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | VTI | meas | m_tcd_Vertebral_A | r_VTI | bilateral | 125100\20354-7\C12105-04\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | HR | meas | m_tcd_Vertebral_A | r_HeartRate | bilateral | 125100\8867-4\C12105-04\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Vertebral A | Dist | meas | m_tcd_Vertebral_A | r_Distance | bilateral | 125100\C12105-04\G-0364\T-40501 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | ? | meas | m_tcd_Bi_MCA | r_DopAngle | bilateral | 125100\C12122-01\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | SV Depth | meas | m_tcd_Bi_MCA | r_Svd | bilateral | 125100\C12221-02\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | Mean (TCD) | meas | m_tcd_Bi_MCA | r_TcdMean | bilateral | 125100\C12221-03\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | PI (TCD) | meas | m_tcd_Bi_MCA | r_TcdPI | bilateral | 125100\C12221-04\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | EDV | meas | m_tcd_Bi_MCA | r_EDV | bilateral | 125100\11653-3\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | MDV | meas | m_tcd_Bi_MCA | r_MDV | bilateral | 125100\11665-7\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | TAPV | meas | m_tcd_Bi_MCA | r_TAPV | bilateral | 125100\11692-1\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | PSV | meas | m_tcd_Bi_MCA | r_PSV | bilateral | 125100\11726-7\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | PI | meas | m_tcd_Bi_MCA | r_PI | bilateral | 125100\12008-9\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | RI | meas | m_tcd_Bi_MCA | r_RI | bilateral | 125100\12023-8\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | S/D | meas | m_tcd_Bi_MCA | r_SD | bilateral | 125100\12144-2\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | AI | meas | m_tcd_Bi_MCA | r_AccelIndex | bilateral | 125100\20167-3\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | AT | meas | m_tcd_Bi_MCA | r_AccelTime | bilateral | 125100\20168-1\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | DT | meas | m_tcd_Bi_MCA | r_DecelTime | bilateral | 125100\20217-6\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | PG | meas | m_tcd_Bi_MCA | r_PG | bilateral | 125100\20247-3\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | TAMV | meas | m_tcd_Bi_MCA | r_TAMV | bilateral | 125100\20352-1\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | VTI | meas | m_tcd_Bi_MCA | r_VTI | bilateral | 125100\20354-7\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | HR | meas | m_tcd_Bi_MCA | r_HeartRate | bilateral | 125100\8867-4\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur MCA | Dist | meas | m_tcd_Bi_MCA | r_Distance | bilateral | 125100\G-0364\G-0368\T-40501 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | ? | meas | m_tcd_Bi_ACA | r_DopAngle | bilateral | 125100\C12122-01\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | SV Depth | meas | m_tcd_Bi_ACA | r_Svd | bilateral | 125100\C12221-02\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | Mean (TCD) | meas | m_tcd_Bi_ACA | r_TcdMean | bilateral | 125100\C12221-03\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | PI (TCD) | meas | m_tcd_Bi_ACA | r_TcdPI | bilateral | 125100\C12221-04\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | EDV | meas | m_tcd_Bi_ACA | r_EDV | bilateral | 125100\11653-3\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | MDV | meas | m_tcd_Bi_ACA | r_MDV | bilateral | 125100\11665-7\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | TAPV | meas | m_tcd_Bi_ACA | r_TAPV | bilateral | 125100\11692-1\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | PSV | meas | m_tcd_Bi_ACA | r_PSV | bilateral | 125100\11726-7\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | PI | meas | m_tcd_Bi_ACA | r_PI | bilateral | 125100\12008-9\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | RI | meas | m_tcd_Bi_ACA | r_RI | bilateral | 125100\12023-8\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | S/D | meas | m_tcd_Bi_ACA | r_SD | bilateral | 125100\12144-2\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | AI | meas | m_tcd_Bi_ACA | r_AccelIndex | bilateral | 125100\20167-3\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | AT | meas | m_tcd_Bi_ACA | r_AccelTime | bilateral | 125100\20168-1\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | DT | meas | m_tcd_Bi_ACA | r_DecelTime | bilateral | 125100\20217-6\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | PG | meas | m_tcd_Bi_ACA | r_PG | bilateral | 125100\20247-3\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | TAMV | meas | m_tcd_Bi_ACA | r_TAMV | bilateral | 125100\20352-1\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | VTI | meas | m_tcd_Bi_ACA | r_VTI | bilateral | 125100\20354-7\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | HR | meas | m_tcd_Bi_ACA | r_HeartRate | bilateral | 125100\8867-4\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Bifur ACA | Dist | meas | m_tcd_Bi_ACA | r_Distance | bilateral | 125100\G-0364\G-0369\T-40501 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation
Vascular Ultrasound Procedure Report | TCD | Prox M1 | ? | meas | m_tcd_Prox_MCA_M1 | r_DopAngle | bilateral | 125100\C12122-01\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | SV Depth | meas | m_tcd_Prox_MCA_M1 | r_Svd | bilateral | 125100\C12221-02\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | Mean (TCD) | meas | m_tcd_Prox_MCA_M1 | r_TcdMean | bilateral | 125100\C12221-03\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | PI (TCD) | meas | m_tcd_Prox_MCA_M1 | r_TcdPI | bilateral | 125100\C12221-04\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | EDV | meas | m_tcd_Prox_MCA_M1 | r_EDV | bilateral | 125100\11653-3\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | MDV | meas | m_tcd_Prox_MCA_M1 | r_MDV | bilateral | 125100\11665-7\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | TAPV | meas | m_tcd_Prox_MCA_M1 | r_TAPV | bilateral | 125100\11692-1\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | PSV | meas | m_tcd_Prox_MCA_M1 | r_PSV | bilateral | 125100\11726-7\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | PI | meas | m_tcd_Prox_MCA_M1 | r_PI | bilateral | 125100\12008-9\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | RI | meas | m_tcd_Prox_MCA_M1 | r_RI | bilateral | 125100\12023-8\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | S/D | meas | m_tcd_Prox_MCA_M1 | r_SD | bilateral | 125100\12144-2\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | AI | meas | m_tcd_Prox_MCA_M1 | r_AccelIndex | bilateral | 125100\20167-3\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | AT | meas | m_tcd_Prox_MCA_M1 | r_AccelTime | bilateral | 125100\20168-1\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | DT | meas | m_tcd_Prox_MCA_M1 | r_DecelTime | bilateral | 125100\20217-6\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | PG | meas | m_tcd_Prox_MCA_M1 | r_PG | bilateral | 125100\20247-3\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | TAMV | meas | m_tcd_Prox_MCA_M1 | r_TAMV | bilateral | 125100\20352-1\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | VTI | meas | m_tcd_Prox_MCA_M1 | r_VTI | bilateral | 125100\20354-7\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | HR | meas | m_tcd_Prox_MCA_M1 | r_HeartRate | bilateral | 125100\8867-4\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox M1 | Dist | meas | m_tcd_Prox_MCA_M1 | r_Distance | bilateral | 125100\G-0364\G-A118\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | ? | meas | m_tcd_Dist_MCA_M1 | r_DopAngle | bilateral | 125100\C12122-01\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | SV Depth | meas | m_tcd_Dist_MCA_M1 | r_Svd | bilateral | 125100\C12221-02\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | Mean (TCD) | meas | m_tcd_Dist_MCA_M1 | r_TcdMean | bilateral | 125100\C12221-03\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | PI (TCD) | meas | m_tcd_Dist_MCA_M1 | r_TcdPI | bilateral | 125100\C12221-04\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | EDV | meas | m_tcd_Dist_MCA_M1 | r_EDV | bilateral | 125100\11653-3\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | MDV | meas | m_tcd_Dist_MCA_M1 | r_MDV | bilateral | 125100\11665-7\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | TAPV | meas | m_tcd_Dist_MCA_M1 | r_TAPV | bilateral | 125100\11692-1\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | PSV | meas | m_tcd_Dist_MCA_M1 | r_PSV | bilateral | 125100\11726-7\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | PI | meas | m_tcd_Dist_MCA_M1 | r_PI | bilateral | 125100\12008-9\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | RI | meas | m_tcd_Dist_MCA_M1 | r_RI | bilateral | 125100\12023-8\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | S/D | meas | m_tcd_Dist_MCA_M1 | r_SD | bilateral | 125100\12144-2\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | AI | meas | m_tcd_Dist_MCA_M1 | r_AccelIndex | bilateral | 125100\20167-3\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | AT | meas | m_tcd_Dist_MCA_M1 | r_AccelTime | bilateral | 125100\20168-1\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | DT | meas | m_tcd_Dist_MCA_M1 | r_DecelTime | bilateral | 125100\20217-6\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | PG | meas | m_tcd_Dist_MCA_M1 | r_PG | bilateral | 125100\20247-3\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | TAMV | meas | m_tcd_Dist_MCA_M1 | r_TAMV | bilateral | 125100\20352-1\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | VTI | meas | m_tcd_Dist_MCA_M1 | r_VTI | bilateral | 125100\20354-7\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | HR | meas | m_tcd_Dist_MCA_M1 | r_HeartRate | bilateral | 125100\8867-4\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist M1 | Dist | meas | m_tcd_Dist_MCA_M1 | r_Distance | bilateral | 125100\G-0364\G-A119\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | ? | meas | m_tcd_Mid_MCA_M1 | r_DopAngle | bilateral | 125100\C12122-01\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | SV Depth | meas | m_tcd_Mid_MCA_M1 | r_Svd | bilateral | 125100\C12221-02\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | Mean (TCD) | meas | m_tcd_Mid_MCA_M1 | r_TcdMean | bilateral | 125100\C12221-03\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | PI (TCD) | meas | m_tcd_Mid_MCA_M1 | r_TcdPI | bilateral | 125100\C12221-04\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | EDV | meas | m_tcd_Mid_MCA_M1 | r_EDV | bilateral | 125100\11653-3\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | MDV | meas | m_tcd_Mid_MCA_M1 | r_MDV | bilateral | 125100\11665-7\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | TAPV | meas | m_tcd_Mid_MCA_M1 | r_TAPV | bilateral | 125100\11692-1\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | PSV | meas | m_tcd_Mid_MCA_M1 | r_PSV | bilateral | 125100\11726-7\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | PI | meas | m_tcd_Mid_MCA_M1 | r_PI | bilateral | 125100\12008-9\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | RI | meas | m_tcd_Mid_MCA_M1 | r_RI | bilateral | 125100\12023-8\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | S/D | meas | m_tcd_Mid_MCA_M1 | r_SD | bilateral | 125100\12144-2\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | AI | meas | m_tcd_Mid_MCA_M1 | r_AccelIndex | bilateral | 125100\20167-3\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | AT | meas | m_tcd_Mid_MCA_M1 | r_AccelTime | bilateral | 125100\20168-1\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | DT | meas | m_tcd_Mid_MCA_M1 | r_DecelTime | bilateral | 125100\20217-6\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | PG | meas | m_tcd_Mid_MCA_M1 | r_PG | bilateral | 125100\20247-3\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | TAMV | meas | m_tcd_Mid_MCA_M1 | r_TAMV | bilateral | 125100\20352-1\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | VTI | meas | m_tcd_Mid_MCA_M1 | r_VTI | bilateral | 125100\20354-7\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | HR | meas | m_tcd_Mid_MCA_M1 | r_HeartRate | bilateral | 125100\8867-4\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid M1 | Dist | meas | m_tcd_Mid_MCA_M1 | r_Distance | bilateral | 125100\G-0364\G-A188\R-1024F\T-40501 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | M2 | ? | meas | m_tcd_MCA_M2 | r_DopAngle | bilateral | 125100\C12122-01\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | SV Depth | meas | m_tcd_MCA_M2 | r_Svd | bilateral | 125100\C12221-02\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | Mean (TCD) | meas | m_tcd_MCA_M2 | r_TcdMean | bilateral | 125100\C12221-03\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | PI (TCD) | meas | m_tcd_MCA_M2 | r_TcdPI | bilateral | 125100\C12221-04\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | EDV | meas | m_tcd_MCA_M2 | r_EDV | bilateral | 125100\11653-3\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | MDV | meas | m_tcd_MCA_M2 | r_MDV | bilateral | 125100\11665-7\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | TAPV | meas | m_tcd_MCA_M2 | r_TAPV | bilateral | 125100\11692-1\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | PSV | meas | m_tcd_MCA_M2 | r_PSV | bilateral | 125100\11726-7\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | PI | meas | m_tcd_MCA_M2 | r_PI | bilateral | 125100\12008-9\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | RI | meas | m_tcd_MCA_M2 | r_RI | bilateral | 125100\12023-8\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | S/D | meas | m_tcd_MCA_M2 | r_SD | bilateral | 125100\12144-2\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | AI | meas | m_tcd_MCA_M2 | r_AccelIndex | bilateral | 125100\20167-3\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | AT | meas | m_tcd_MCA_M2 | r_AccelTime | bilateral | 125100\20168-1\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | DT | meas | m_tcd_MCA_M2 | r_DecelTime | bilateral | 125100\20217-6\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | PG | meas | m_tcd_MCA_M2 | r_PG | bilateral | 125100\20247-3\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | TAMV | meas | m_tcd_MCA_M2 | r_TAMV | bilateral | 125100\20352-1\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | VTI | meas | m_tcd_MCA_M2 | r_VTI | bilateral | 125100\20354-7\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | HR | meas | m_tcd_MCA_M2 | r_HeartRate | bilateral | 125100\8867-4\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | M2 | Dist | meas | m_tcd_MCA_M2 | r_Distance | bilateral | 125100\G-0364\R-10251\T-40501 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | ? | meas | m_tcd_PCA_P1 | r_DopAngle | bilateral | 125100\C12122-01\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | SV Depth | meas | m_tcd_PCA_P1 | r_Svd | bilateral | 125100\C12221-02\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | Mean (TCD) | meas | m_tcd_PCA_P1 | r_TcdMean | bilateral | 125100\C12221-03\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | PI (TCD) | meas | m_tcd_PCA_P1 | r_TcdPI | bilateral | 125100\C12221-04\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | EDV | meas | m_tcd_PCA_P1 | r_EDV | bilateral | 125100\11653-3\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | MDV | meas | m_tcd_PCA_P1 | r_MDV | bilateral | 125100\11665-7\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | TAPV | meas | m_tcd_PCA_P1 | r_TAPV | bilateral | 125100\11692-1\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | PSV | meas | m_tcd_PCA_P1 | r_PSV | bilateral | 125100\11726-7\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | PI | meas | m_tcd_PCA_P1 | r_PI | bilateral | 125100\12008-9\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | RI | meas | m_tcd_PCA_P1 | r_RI | bilateral | 125100\12023-8\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | S/D | meas | m_tcd_PCA_P1 | r_SD | bilateral | 125100\12144-2\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | AI | meas | m_tcd_PCA_P1 | r_AccelIndex | bilateral | 125100\20167-3\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | AT | meas | m_tcd_PCA_P1 | r_AccelTime | bilateral | 125100\20168-1\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | DT | meas | m_tcd_PCA_P1 | r_DecelTime | bilateral | 125100\20217-6\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | PG | meas | m_tcd_PCA_P1 | r_PG | bilateral | 125100\20247-3\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | TAMV | meas | m_tcd_PCA_P1 | r_TAMV | bilateral | 125100\20352-1\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | VTI | meas | m_tcd_PCA_P1 | r_VTI | bilateral | 125100\20354-7\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | HR | meas | m_tcd_PCA_P1 | r_HeartRate | bilateral | 125100\8867-4\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P1 | Dist | meas | m_tcd_PCA_P1 | r_Distance | bilateral | 125100\G-0364\R-10253\T-40501 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | ? | meas | m_tcd_PCA_P2 | r_DopAngle | bilateral | 125100\C12122-01\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | SV Depth | meas | m_tcd_PCA_P2 | r_Svd | bilateral | 125100\C12221-02\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | Mean (TCD) | meas | m_tcd_PCA_P2 | r_TcdMean | bilateral | 125100\C12221-03\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | PI (TCD) | meas | m_tcd_PCA_P2 | r_TcdPI | bilateral | 125100\C12221-04\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | EDV | meas | m_tcd_PCA_P2 | r_EDV | bilateral | 125100\11653-3\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | MDV | meas | m_tcd_PCA_P2 | r_MDV | bilateral | 125100\11665-7\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | TAPV | meas | m_tcd_PCA_P2 | r_TAPV | bilateral | 125100\11692-1\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | PSV | meas | m_tcd_PCA_P2 | r_PSV | bilateral | 125100\11726-7\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | PI | meas | m_tcd_PCA_P2 | r_PI | bilateral | 125100\12008-9\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | RI | meas | m_tcd_PCA_P2 | r_RI | bilateral | 125100\12023-8\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | S/D | meas | m_tcd_PCA_P2 | r_SD | bilateral | 125100\12144-2\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | AI | meas | m_tcd_PCA_P2 | r_AccelIndex | bilateral | 125100\20167-3\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | AT | meas | m_tcd_PCA_P2 | r_AccelTime | bilateral | 125100\20168-1\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | DT | meas | m_tcd_PCA_P2 | r_DecelTime | bilateral | 125100\20217-6\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | PG | meas | m_tcd_PCA_P2 | r_PG | bilateral | 125100\20247-3\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | TAMV | meas | m_tcd_PCA_P2 | r_TAMV | bilateral | 125100\20352-1\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | VTI | meas | m_tcd_PCA_P2 | r_VTI | bilateral | 125100\20354-7\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | HR | meas | m_tcd_PCA_P2 | r_HeartRate | bilateral | 125100\8867-4\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | PCA P2 | Dist | meas | m_tcd_PCA_P2 | r_Distance | bilateral | 125100\G-0364\R-10255\T-40501 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment
Vascular Ultrasound Procedure Report | TCD | Term ICA | ? | meas | m_tcd_Term_ICA | r_DopAngle | bilateral | 125100\C12122-01\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | SV Depth | meas | m_tcd_Term_ICA | r_Svd | bilateral | 125100\C12221-02\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | Mean (TCD) | meas | m_tcd_Term_ICA | r_TcdMean | bilateral | 125100\C12221-03\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | PI (TCD) | meas | m_tcd_Term_ICA | r_TcdPI | bilateral | 125100\C12221-04\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | EDV | meas | m_tcd_Term_ICA | r_EDV | bilateral | 125100\11653-3\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | MDV | meas | m_tcd_Term_ICA | r_MDV | bilateral | 125100\11665-7\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | TAPV | meas | m_tcd_Term_ICA | r_TAPV | bilateral | 125100\11692-1\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | PSV | meas | m_tcd_Term_ICA | r_PSV | bilateral | 125100\11726-7\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | PI | meas | m_tcd_Term_ICA | r_PI | bilateral | 125100\12008-9\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | RI | meas | m_tcd_Term_ICA | r_RI | bilateral | 125100\12023-8\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | S/D | meas | m_tcd_Term_ICA | r_SD | bilateral | 125100\12144-2\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | AI | meas | m_tcd_Term_ICA | r_AccelIndex | bilateral | 125100\20167-3\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | AT | meas | m_tcd_Term_ICA | r_AccelTime | bilateral | 125100\20168-1\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | DT | meas | m_tcd_Term_ICA | r_DecelTime | bilateral | 125100\20217-6\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | PG | meas | m_tcd_Term_ICA | r_PG | bilateral | 125100\20247-3\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | TAMV | meas | m_tcd_Term_ICA | r_TAMV | bilateral | 125100\20352-1\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | VTI | meas | m_tcd_Term_ICA | r_VTI | bilateral | 125100\20354-7\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | HR | meas | m_tcd_Term_ICA | r_HeartRate | bilateral | 125100\8867-4\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Term ICA | Dist | meas | m_tcd_Term_ICA | r_Distance | bilateral | 125100\G-0364\R-102BD\T-40501 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, R-102BD, Terminal internal carotid artery
Vascular Ultrasound Procedure Report | TCD | Dist ICA | ? | meas | m_tcd_Dist_ICA | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | SV Depth | meas | m_tcd_Dist_ICA | r_Svd | bilateral | 125100\C12221-02\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | Mean (TCD) | meas | m_tcd_Dist_ICA | r_TcdMean | bilateral | 125100\C12221-03\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | PI (TCD) | meas | m_tcd_Dist_ICA | r_TcdPI | bilateral | 125100\C12221-04\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | EDV | meas | m_tcd_Dist_ICA | r_EDV | bilateral | 125100\11653-3\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | MDV | meas | m_tcd_Dist_ICA | r_MDV | bilateral | 125100\11665-7\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | TAPV | meas | m_tcd_Dist_ICA | r_TAPV | bilateral | 125100\11692-1\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | PSV | meas | m_tcd_Dist_ICA | r_PSV | bilateral | 125100\11726-7\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | PI | meas | m_tcd_Dist_ICA | r_PI | bilateral | 125100\12008-9\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | RI | meas | m_tcd_Dist_ICA | r_RI | bilateral | 125100\12023-8\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | S/D | meas | m_tcd_Dist_ICA | r_SD | bilateral | 125100\12144-2\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | AI | meas | m_tcd_Dist_ICA | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | AT | meas | m_tcd_Dist_ICA | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | DT | meas | m_tcd_Dist_ICA | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | PG | meas | m_tcd_Dist_ICA | r_PG | bilateral | 125100\20247-3\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | TAMV | meas | m_tcd_Dist_ICA | r_TAMV | bilateral | 125100\20352-1\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | VTI | meas | m_tcd_Dist_ICA | r_VTI | bilateral | 125100\20354-7\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | HR | meas | m_tcd_Dist_ICA | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist ICA | Dist | meas | m_tcd_Dist_ICA | r_Distance | bilateral | 125100\G-0364\G-A119\T-40501\T-45300 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | ? | meas | m_tcd_Car_Siph | r_DopAngle | bilateral | 125100\C12122-01\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | SV Depth | meas | m_tcd_Car_Siph | r_Svd | bilateral | 125100\C12221-02\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | Mean (TCD) | meas | m_tcd_Car_Siph | r_TcdMean | bilateral | 125100\C12221-03\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | PI (TCD) | meas | m_tcd_Car_Siph | r_TcdPI | bilateral | 125100\C12221-04\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | EDV | meas | m_tcd_Car_Siph | r_EDV | bilateral | 125100\11653-3\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | MDV | meas | m_tcd_Car_Siph | r_MDV | bilateral | 125100\11665-7\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | TAPV | meas | m_tcd_Car_Siph | r_TAPV | bilateral | 125100\11692-1\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | PSV | meas | m_tcd_Car_Siph | r_PSV | bilateral | 125100\11726-7\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | PI | meas | m_tcd_Car_Siph | r_PI | bilateral | 125100\12008-9\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | RI | meas | m_tcd_Car_Siph | r_RI | bilateral | 125100\12023-8\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | S/D | meas | m_tcd_Car_Siph | r_SD | bilateral | 125100\12144-2\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | AI | meas | m_tcd_Car_Siph | r_AccelIndex | bilateral | 125100\20167-3\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | AT | meas | m_tcd_Car_Siph | r_AccelTime | bilateral | 125100\20168-1\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | DT | meas | m_tcd_Car_Siph | r_DecelTime | bilateral | 125100\20217-6\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | PG | meas | m_tcd_Car_Siph | r_PG | bilateral | 125100\20247-3\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | TAMV | meas | m_tcd_Car_Siph | r_TAMV | bilateral | 125100\20352-1\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | VTI | meas | m_tcd_Car_Siph | r_VTI | bilateral | 125100\20354-7\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | HR | meas | m_tcd_Car_Siph | r_HeartRate | bilateral | 125100\8867-4\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | Carotid Siphon | Dist | meas | m_tcd_Car_Siph | r_Distance | bilateral | 125100\G-0364\T-40501\T-45308 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45308, Carotid Siphon
Vascular Ultrasound Procedure Report | TCD | PCoA | ? | meas | m_tcd_PComA | r_DopAngle | bilateral | 125100\C12122-01\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | SV Depth | meas | m_tcd_PComA | r_Svd | bilateral | 125100\C12221-02\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | Mean (TCD) | meas | m_tcd_PComA | r_TcdMean | bilateral | 125100\C12221-03\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | PI (TCD) | meas | m_tcd_PComA | r_TcdPI | bilateral | 125100\C12221-04\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | EDV | meas | m_tcd_PComA | r_EDV | bilateral | 125100\11653-3\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | MDV | meas | m_tcd_PComA | r_MDV | bilateral | 125100\11665-7\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | TAPV | meas | m_tcd_PComA | r_TAPV | bilateral | 125100\11692-1\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | PSV | meas | m_tcd_PComA | r_PSV | bilateral | 125100\11726-7\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | PI | meas | m_tcd_PComA | r_PI | bilateral | 125100\12008-9\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | RI | meas | m_tcd_PComA | r_RI | bilateral | 125100\12023-8\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | S/D | meas | m_tcd_PComA | r_SD | bilateral | 125100\12144-2\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | AI | meas | m_tcd_PComA | r_AccelIndex | bilateral | 125100\20167-3\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | AT | meas | m_tcd_PComA | r_AccelTime | bilateral | 125100\20168-1\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | DT | meas | m_tcd_PComA | r_DecelTime | bilateral | 125100\20217-6\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | PG | meas | m_tcd_PComA | r_PG | bilateral | 125100\20247-3\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | TAMV | meas | m_tcd_PComA | r_TAMV | bilateral | 125100\20352-1\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | VTI | meas | m_tcd_PComA | r_VTI | bilateral | 125100\20354-7\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | HR | meas | m_tcd_PComA | r_HeartRate | bilateral | 125100\8867-4\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | PCoA | Dist | meas | m_tcd_PComA | r_Distance | bilateral | 125100\G-0364\T-40501\T-45320 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45320, Posterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | ? | meas | m_tcd_Oph_A | r_DopAngle | bilateral | 125100\C12122-01\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | SV Depth | meas | m_tcd_Oph_A | r_Svd | bilateral | 125100\C12221-02\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | Mean (TCD) | meas | m_tcd_Oph_A | r_TcdMean | bilateral | 125100\C12221-03\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | PI (TCD) | meas | m_tcd_Oph_A | r_TcdPI | bilateral | 125100\C12221-04\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | EDV | meas | m_tcd_Oph_A | r_EDV | bilateral | 125100\11653-3\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | MDV | meas | m_tcd_Oph_A | r_MDV | bilateral | 125100\11665-7\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | TAPV | meas | m_tcd_Oph_A | r_TAPV | bilateral | 125100\11692-1\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | PSV | meas | m_tcd_Oph_A | r_PSV | bilateral | 125100\11726-7\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | PI | meas | m_tcd_Oph_A | r_PI | bilateral | 125100\12008-9\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | RI | meas | m_tcd_Oph_A | r_RI | bilateral | 125100\12023-8\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | S/D | meas | m_tcd_Oph_A | r_SD | bilateral | 125100\12144-2\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | AI | meas | m_tcd_Oph_A | r_AccelIndex | bilateral | 125100\20167-3\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | AT | meas | m_tcd_Oph_A | r_AccelTime | bilateral | 125100\20168-1\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | DT | meas | m_tcd_Oph_A | r_DecelTime | bilateral | 125100\20217-6\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | PG | meas | m_tcd_Oph_A | r_PG | bilateral | 125100\20247-3\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | TAMV | meas | m_tcd_Oph_A | r_TAMV | bilateral | 125100\20352-1\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | VTI | meas | m_tcd_Oph_A | r_VTI | bilateral | 125100\20354-7\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | HR | meas | m_tcd_Oph_A | r_HeartRate | bilateral | 125100\8867-4\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | Ophthalmic A | Dist | meas | m_tcd_Oph_A | r_Distance | bilateral | 125100\G-0364\T-40501\T-45400 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45400, Ophthalmic Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | ? | meas | m_tcd_AComA | r_DopAngle | bilateral | 125100\C12122-01\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | SV Depth | meas | m_tcd_AComA | r_Svd | bilateral | 125100\C12221-02\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | Mean (TCD) | meas | m_tcd_AComA | r_TcdMean | bilateral | 125100\C12221-03\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | PI (TCD) | meas | m_tcd_AComA | r_TcdPI | bilateral | 125100\C12221-04\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | EDV | meas | m_tcd_AComA | r_EDV | bilateral | 125100\11653-3\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | MDV | meas | m_tcd_AComA | r_MDV | bilateral | 125100\11665-7\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | TAPV | meas | m_tcd_AComA | r_TAPV | bilateral | 125100\11692-1\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | PSV | meas | m_tcd_AComA | r_PSV | bilateral | 125100\11726-7\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | PI | meas | m_tcd_AComA | r_PI | bilateral | 125100\12008-9\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | RI | meas | m_tcd_AComA | r_RI | bilateral | 125100\12023-8\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | S/D | meas | m_tcd_AComA | r_SD | bilateral | 125100\12144-2\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | AI | meas | m_tcd_AComA | r_AccelIndex | bilateral | 125100\20167-3\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | AT | meas | m_tcd_AComA | r_AccelTime | bilateral | 125100\20168-1\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | DT | meas | m_tcd_AComA | r_DecelTime | bilateral | 125100\20217-6\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | PG | meas | m_tcd_AComA | r_PG | bilateral | 125100\20247-3\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | TAMV | meas | m_tcd_AComA | r_TAMV | bilateral | 125100\20352-1\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | VTI | meas | m_tcd_AComA | r_VTI | bilateral | 125100\20354-7\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | HR | meas | m_tcd_AComA | r_HeartRate | bilateral | 125100\8867-4\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACoA | Dist | meas | m_tcd_AComA | r_Distance | bilateral | 125100\G-0364\T-40501\T-45530 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45530, Anterior Communicating Artery
Vascular Ultrasound Procedure Report | TCD | ACA | ? | meas | m_tcd_ACA_A1 | r_DopAngle | bilateral | 125100\C12122-01\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | SV Depth | meas | m_tcd_ACA_A1 | r_Svd | bilateral | 125100\C12221-02\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | Mean (TCD) | meas | m_tcd_ACA_A1 | r_TcdMean | bilateral | 125100\C12221-03\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | PI (TCD) | meas | m_tcd_ACA_A1 | r_TcdPI | bilateral | 125100\C12221-04\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | EDV | meas | m_tcd_ACA_A1 | r_EDV | bilateral | 125100\11653-3\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | MDV | meas | m_tcd_ACA_A1 | r_MDV | bilateral | 125100\11665-7\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | TAPV | meas | m_tcd_ACA_A1 | r_TAPV | bilateral | 125100\11692-1\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | PSV | meas | m_tcd_ACA_A1 | r_PSV | bilateral | 125100\11726-7\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | PI | meas | m_tcd_ACA_A1 | r_PI | bilateral | 125100\12008-9\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | RI | meas | m_tcd_ACA_A1 | r_RI | bilateral | 125100\12023-8\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | S/D | meas | m_tcd_ACA_A1 | r_SD | bilateral | 125100\12144-2\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | AI | meas | m_tcd_ACA_A1 | r_AccelIndex | bilateral | 125100\20167-3\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | AT | meas | m_tcd_ACA_A1 | r_AccelTime | bilateral | 125100\20168-1\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | DT | meas | m_tcd_ACA_A1 | r_DecelTime | bilateral | 125100\20217-6\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | PG | meas | m_tcd_ACA_A1 | r_PG | bilateral | 125100\20247-3\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | TAMV | meas | m_tcd_ACA_A1 | r_TAMV | bilateral | 125100\20352-1\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | VTI | meas | m_tcd_ACA_A1 | r_VTI | bilateral | 125100\20354-7\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | HR | meas | m_tcd_ACA_A1 | r_HeartRate | bilateral | 125100\8867-4\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | ACA | Dist | meas | m_tcd_ACA_A1 | r_Distance | bilateral | 125100\G-0364\T-40501\T-45540 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45540, Anterior Cerebral Artery
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | ? | meas | m_tcd_Prox_Bas_A | r_DopAngle | unilateral | 125100\C12122-01\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | SV Depth | meas | m_tcd_Prox_Bas_A | r_Svd | unilateral | 125100\C12221-02\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | Mean (TCD) | meas | m_tcd_Prox_Bas_A | r_TcdMean | unilateral | 125100\C12221-03\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | PI (TCD) | meas | m_tcd_Prox_Bas_A | r_TcdPI | unilateral | 125100\C12221-04\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | EDV | meas | m_tcd_Prox_Bas_A | r_EDV | unilateral | 125100\11653-3\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | MDV | meas | m_tcd_Prox_Bas_A | r_MDV | unilateral | 125100\11665-7\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | TAPV | meas | m_tcd_Prox_Bas_A | r_TAPV | unilateral | 125100\11692-1\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | PSV | meas | m_tcd_Prox_Bas_A | r_PSV | unilateral | 125100\11726-7\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | PI | meas | m_tcd_Prox_Bas_A | r_PI | unilateral | 125100\12008-9\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | RI | meas | m_tcd_Prox_Bas_A | r_RI | unilateral | 125100\12023-8\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | S/D | meas | m_tcd_Prox_Bas_A | r_SD | unilateral | 125100\12144-2\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | AI | meas | m_tcd_Prox_Bas_A | r_AccelIndex | unilateral | 125100\20167-3\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | AT | meas | m_tcd_Prox_Bas_A | r_AccelTime | unilateral | 125100\20168-1\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | DT | meas | m_tcd_Prox_Bas_A | r_DecelTime | unilateral | 125100\20217-6\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | PG | meas | m_tcd_Prox_Bas_A | r_PG | unilateral | 125100\20247-3\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | TAMV | meas | m_tcd_Prox_Bas_A | r_TAMV | unilateral | 125100\20352-1\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | VTI | meas | m_tcd_Prox_Bas_A | r_VTI | unilateral | 125100\20354-7\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | HR | meas | m_tcd_Prox_Bas_A | r_HeartRate | unilateral | 125100\8867-4\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Prox Basilar A | Dist | meas | m_tcd_Prox_Bas_A | r_Distance | unilateral | 125100\G-0364\G-A118\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | ? | meas | m_tcd_Dist_Bas_A | r_DopAngle | unilateral | 125100\C12122-01\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | SV Depth | meas | m_tcd_Dist_Bas_A | r_Svd | unilateral | 125100\C12221-02\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | Mean (TCD) | meas | m_tcd_Dist_Bas_A | r_TcdMean | unilateral | 125100\C12221-03\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | PI (TCD) | meas | m_tcd_Dist_Bas_A | r_TcdPI | unilateral | 125100\C12221-04\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | EDV | meas | m_tcd_Dist_Bas_A | r_EDV | unilateral | 125100\11653-3\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | MDV | meas | m_tcd_Dist_Bas_A | r_MDV | unilateral | 125100\11665-7\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | TAPV | meas | m_tcd_Dist_Bas_A | r_TAPV | unilateral | 125100\11692-1\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | PSV | meas | m_tcd_Dist_Bas_A | r_PSV | unilateral | 125100\11726-7\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | PI | meas | m_tcd_Dist_Bas_A | r_PI | unilateral | 125100\12008-9\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | RI | meas | m_tcd_Dist_Bas_A | r_RI | unilateral | 125100\12023-8\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | S/D | meas | m_tcd_Dist_Bas_A | r_SD | unilateral | 125100\12144-2\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | AI | meas | m_tcd_Dist_Bas_A | r_AccelIndex | unilateral | 125100\20167-3\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | AT | meas | m_tcd_Dist_Bas_A | r_AccelTime | unilateral | 125100\20168-1\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | DT | meas | m_tcd_Dist_Bas_A | r_DecelTime | unilateral | 125100\20217-6\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | PG | meas | m_tcd_Dist_Bas_A | r_PG | unilateral | 125100\20247-3\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | TAMV | meas | m_tcd_Dist_Bas_A | r_TAMV | unilateral | 125100\20352-1\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | VTI | meas | m_tcd_Dist_Bas_A | r_VTI | unilateral | 125100\20354-7\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | HR | meas | m_tcd_Dist_Bas_A | r_HeartRate | unilateral | 125100\8867-4\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Dist Basilar A | Dist | meas | m_tcd_Dist_Bas_A | r_Distance | unilateral | 125100\G-0364\G-A119\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | ? | meas | m_tcd_Mid_Bas_A | r_DopAngle | unilateral | 125100\C12122-01\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | SV Depth | meas | m_tcd_Mid_Bas_A | r_Svd | unilateral | 125100\C12221-02\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-02, Sample Volume Depth |  | UCUM, mm, mm | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | Mean (TCD) | meas | m_tcd_Mid_Bas_A | r_TcdMean | unilateral | 125100\C12221-03\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | PI (TCD) | meas | m_tcd_Mid_Bas_A | r_TcdPI | unilateral | 125100\C12221-04\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | EDV | meas | m_tcd_Mid_Bas_A | r_EDV | unilateral | 125100\11653-3\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | MDV | meas | m_tcd_Mid_Bas_A | r_MDV | unilateral | 125100\11665-7\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | TAPV | meas | m_tcd_Mid_Bas_A | r_TAPV | unilateral | 125100\11692-1\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | PSV | meas | m_tcd_Mid_Bas_A | r_PSV | unilateral | 125100\11726-7\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | PI | meas | m_tcd_Mid_Bas_A | r_PI | unilateral | 125100\12008-9\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | RI | meas | m_tcd_Mid_Bas_A | r_RI | unilateral | 125100\12023-8\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | S/D | meas | m_tcd_Mid_Bas_A | r_SD | unilateral | 125100\12144-2\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | AI | meas | m_tcd_Mid_Bas_A | r_AccelIndex | unilateral | 125100\20167-3\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | AT | meas | m_tcd_Mid_Bas_A | r_AccelTime | unilateral | 125100\20168-1\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | DT | meas | m_tcd_Mid_Bas_A | r_DecelTime | unilateral | 125100\20217-6\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | PG | meas | m_tcd_Mid_Bas_A | r_PG | unilateral | 125100\20247-3\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | TAMV | meas | m_tcd_Mid_Bas_A | r_TAMV | unilateral | 125100\20352-1\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | VTI | meas | m_tcd_Mid_Bas_A | r_VTI | unilateral | 125100\20354-7\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | HR | meas | m_tcd_Mid_Bas_A | r_HeartRate | unilateral | 125100\8867-4\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | TCD | Mid Basilar A | Dist | meas | m_tcd_Mid_Bas_A | r_Distance | unilateral | 125100\G-0364\G-A188\T-40501\T-45800 | SRT, T-40501, Blood Vessel of Head |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45800, Basilar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | ICA | ? | meas | m_ICA | r_DopAngle | bilateral | 125100\C12104-01\C12122-01\T-45005 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | ICA | EDV | meas | m_ICA | r_EDV | bilateral | 125100\11653-3\C12104-01\T-45005 | SRT, T-45005, Artery of neck |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | ICA | MDV | meas | m_ICA | r_MDV | bilateral | 125100\11665-7\C12104-01\T-45005 | SRT, T-45005, Artery of neck |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | ICA | TAPV | meas | m_ICA | r_TAPV | bilateral | 125100\11692-1\C12104-01\T-45005 | SRT, T-45005, Artery of neck |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | ICA | PSV | meas | m_ICA | r_PSV | bilateral | 125100\11726-7\C12104-01\T-45005 | SRT, T-45005, Artery of neck |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | ICA | PI | meas | m_ICA | r_PI | bilateral | 125100\12008-9\C12104-01\T-45005 | SRT, T-45005, Artery of neck |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | ICA | RI | meas | m_ICA | r_RI | bilateral | 125100\12023-8\C12104-01\T-45005 | SRT, T-45005, Artery of neck |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | ICA | S/D | meas | m_ICA | r_SD | bilateral | 125100\12144-2\C12104-01\T-45005 | SRT, T-45005, Artery of neck |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | ICA | AI | meas | m_ICA | r_AccelIndex | bilateral | 125100\20167-3\C12104-01\T-45005 | SRT, T-45005, Artery of neck |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | ICA | AT | meas | m_ICA | r_AccelTime | bilateral | 125100\20168-1\C12104-01\T-45005 | SRT, T-45005, Artery of neck |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | ICA | DT | meas | m_ICA | r_DecelTime | bilateral | 125100\20217-6\C12104-01\T-45005 | SRT, T-45005, Artery of neck |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | ICA | PG | meas | m_ICA | r_PG | bilateral | 125100\20247-3\C12104-01\T-45005 | SRT, T-45005, Artery of neck |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | ICA | MG | meas | m_ICA | r_PGmean | bilateral | 125100\20256-4\C12104-01\T-45005 | SRT, T-45005, Artery of neck |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | ICA | TAMV | meas | m_ICA | r_TAMV | bilateral | 125100\20352-1\C12104-01\T-45005 | SRT, T-45005, Artery of neck |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | ICA | VTI | meas | m_ICA | r_VTI | bilateral | 125100\20354-7\C12104-01\T-45005 | SRT, T-45005, Artery of neck |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | ICA | HR | meas | m_ICA | r_HeartRate | bilateral | 125100\8867-4\C12104-01\T-45005 | SRT, T-45005, Artery of neck |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | ICA | Dist | meas | m_ICA | r_Distance | bilateral | 125100\C12104-01\G-0364\T-45005 | SRT, T-45005, Artery of neck |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator
Vascular Ultrasound Procedure Report | Vascular | CCA | ? | meas | m_CCA | r_DopAngle | bilateral | 125100\C12104-02\C12122-01\T-45005 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA | EDV | meas | m_CCA | r_EDV | bilateral | 125100\11653-3\C12104-02\T-45005 | SRT, T-45005, Artery of neck |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA | MDV | meas | m_CCA | r_MDV | bilateral | 125100\11665-7\C12104-02\T-45005 | SRT, T-45005, Artery of neck |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA | TAPV | meas | m_CCA | r_TAPV | bilateral | 125100\11692-1\C12104-02\T-45005 | SRT, T-45005, Artery of neck |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA | PSV | meas | m_CCA | r_PSV | bilateral | 125100\11726-7\C12104-02\T-45005 | SRT, T-45005, Artery of neck |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA | PI | meas | m_CCA | r_PI | bilateral | 125100\12008-9\C12104-02\T-45005 | SRT, T-45005, Artery of neck |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA | RI | meas | m_CCA | r_RI | bilateral | 125100\12023-8\C12104-02\T-45005 | SRT, T-45005, Artery of neck |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA | S/D | meas | m_CCA | r_SD | bilateral | 125100\12144-2\C12104-02\T-45005 | SRT, T-45005, Artery of neck |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA | AI | meas | m_CCA | r_AccelIndex | bilateral | 125100\20167-3\C12104-02\T-45005 | SRT, T-45005, Artery of neck |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA | AT | meas | m_CCA | r_AccelTime | bilateral | 125100\20168-1\C12104-02\T-45005 | SRT, T-45005, Artery of neck |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA | DT | meas | m_CCA | r_DecelTime | bilateral | 125100\20217-6\C12104-02\T-45005 | SRT, T-45005, Artery of neck |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA | PG | meas | m_CCA | r_PG | bilateral | 125100\20247-3\C12104-02\T-45005 | SRT, T-45005, Artery of neck |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA | MG | meas | m_CCA | r_PGmean | bilateral | 125100\20256-4\C12104-02\T-45005 | SRT, T-45005, Artery of neck |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA | TAMV | meas | m_CCA | r_TAMV | bilateral | 125100\20352-1\C12104-02\T-45005 | SRT, T-45005, Artery of neck |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA | VTI | meas | m_CCA | r_VTI | bilateral | 125100\20354-7\C12104-02\T-45005 | SRT, T-45005, Artery of neck |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA | HR | meas | m_CCA | r_HeartRate | bilateral | 125100\8867-4\C12104-02\T-45005 | SRT, T-45005, Artery of neck |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA | Dist | meas | m_CCA | r_Distance | bilateral | 125100\C12104-02\G-0364\T-45005 | SRT, T-45005, Artery of neck |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator
Vascular Ultrasound Procedure Report | Vascular | CCA Total Plaque Vol | Volume | meas | m_CCA_TotalPlqVol | r_Volume | bilateral | 125100\122376\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | DCM, 122376, Total Plaque Volume |  | UCUM, mm3, mm3 | [7 empty] | SRT, T-45100, Common Carotid Artery
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | ? | meas | m_Prox_CCA | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA IMT | Mean (IMT) | meas | m_Prox_CCA_IMT | r_Mean_Dist | bilateral | 125100\C12122-08\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, mm, mm | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA IMT | SD (IMT) | meas | m_Prox_CCA_IMT | r_SD_Dist | bilateral | 125100\C12122-09\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-09, Standard Deviation of Intima Media Thickness |  | UCUM, mm, mm | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | EDV | meas | m_Prox_CCA | r_EDV | bilateral | 125100\11653-3\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | MDV | meas | m_Prox_CCA | r_MDV | bilateral | 125100\11665-7\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | TAPV | meas | m_Prox_CCA | r_TAPV | bilateral | 125100\11692-1\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | PSV | meas | m_Prox_CCA | r_PSV | bilateral | 125100\11726-7\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | PI | meas | m_Prox_CCA | r_PI | bilateral | 125100\12008-9\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | RI | meas | m_Prox_CCA | r_RI | bilateral | 125100\12023-8\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | S/D | meas | m_Prox_CCA | r_SD | bilateral | 125100\12144-2\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | AI | meas | m_Prox_CCA | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | AT | meas | m_Prox_CCA | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | DT | meas | m_Prox_CCA | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | PG | meas | m_Prox_CCA | r_PG | bilateral | 125100\20247-3\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | MG | meas | m_Prox_CCA | r_PGmean | bilateral | 125100\20256-4\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | TAMV | meas | m_Prox_CCA | r_TAMV | bilateral | 125100\20352-1\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | VTI | meas | m_Prox_CCA | r_VTI | bilateral | 125100\20354-7\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | HR | meas | m_Prox_CCA | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA | Dist | meas | m_Prox_CCA | r_Distance | bilateral | 125100\G-0364\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA Max Area % Reduction | Area Reduc | meas | m_Prox_CCA_MxAre%Red | r_AreaReduc_Percent | bilateral | 125100\G-0371\G-A118\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | SRT, G-0371, Percent Area Reduction |  | UCUM, %, Percent | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | ? | meas | m_Dist_CCA | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA IMT | Mean (IMT) | meas | m_Dist_CCA_IMT | r_Mean_Dist | bilateral | 125100\C12122-08\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, mm, mm | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA IMT | SD (IMT) | meas | m_Dist_CCA_IMT | r_SD_Dist | bilateral | 125100\C12122-09\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-09, Standard Deviation of Intima Media Thickness |  | UCUM, mm, mm | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | EDV | meas | m_Dist_CCA | r_EDV | bilateral | 125100\11653-3\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | MDV | meas | m_Dist_CCA | r_MDV | bilateral | 125100\11665-7\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | TAPV | meas | m_Dist_CCA | r_TAPV | bilateral | 125100\11692-1\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | PSV | meas | m_Dist_CCA | r_PSV | bilateral | 125100\11726-7\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | PI | meas | m_Dist_CCA | r_PI | bilateral | 125100\12008-9\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | RI | meas | m_Dist_CCA | r_RI | bilateral | 125100\12023-8\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | S/D | meas | m_Dist_CCA | r_SD | bilateral | 125100\12144-2\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | AI | meas | m_Dist_CCA | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | AT | meas | m_Dist_CCA | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | DT | meas | m_Dist_CCA | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | PG | meas | m_Dist_CCA | r_PG | bilateral | 125100\20247-3\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | MG | meas | m_Dist_CCA | r_PGmean | bilateral | 125100\20256-4\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | TAMV | meas | m_Dist_CCA | r_TAMV | bilateral | 125100\20352-1\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | VTI | meas | m_Dist_CCA | r_VTI | bilateral | 125100\20354-7\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | HR | meas | m_Dist_CCA | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA | Dist | meas | m_Dist_CCA | r_Distance | bilateral | 125100\G-0364\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA Max Area % Reduction | Area Reduc | meas | m_Dist_CCA_MxAre%Red | r_AreaReduc_Percent | bilateral | 125100\G-0371\G-A119\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | SRT, G-0371, Percent Area Reduction |  | UCUM, %, Percent | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | ? | meas | m_Mid_CCA | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA IMT | Mean (IMT) | meas | m_Mid_CCA_IMT | r_Mean_Dist | bilateral | 125100\C12122-08\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, mm, mm | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA IMT | SD (IMT) | meas | m_Mid_CCA_IMT | r_SD_Dist | bilateral | 125100\C12122-09\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-09, Standard Deviation of Intima Media Thickness |  | UCUM, mm, mm | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | EDV | meas | m_Mid_CCA | r_EDV | bilateral | 125100\11653-3\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | MDV | meas | m_Mid_CCA | r_MDV | bilateral | 125100\11665-7\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | TAPV | meas | m_Mid_CCA | r_TAPV | bilateral | 125100\11692-1\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | PSV | meas | m_Mid_CCA | r_PSV | bilateral | 125100\11726-7\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | PI | meas | m_Mid_CCA | r_PI | bilateral | 125100\12008-9\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | RI | meas | m_Mid_CCA | r_RI | bilateral | 125100\12023-8\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | S/D | meas | m_Mid_CCA | r_SD | bilateral | 125100\12144-2\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | AI | meas | m_Mid_CCA | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | AT | meas | m_Mid_CCA | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | DT | meas | m_Mid_CCA | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | PG | meas | m_Mid_CCA | r_PG | bilateral | 125100\20247-3\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | MG | meas | m_Mid_CCA | r_PGmean | bilateral | 125100\20256-4\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | TAMV | meas | m_Mid_CCA | r_TAMV | bilateral | 125100\20352-1\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | VTI | meas | m_Mid_CCA | r_VTI | bilateral | 125100\20354-7\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | HR | meas | m_Mid_CCA | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA | Dist | meas | m_Mid_CCA | r_Distance | bilateral | 125100\G-0364\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA Max Area % Reduction | Area Reduc | meas | m_Mid_CCA_MxAre%Red | r_AreaReduc_Percent | bilateral | 125100\G-0371\G-A188\T-45005\T-45100 | SRT, T-45005, Artery of neck |  | SRT, G-0371, Percent Area Reduction |  | UCUM, %, Percent | [7 empty] | SRT, T-45100, Common Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Bifurcation Total Plaque Vol | Volume | meas | m_Bifurc_TotalPlqVol | r_Volume | bilateral | 125100\122376\T-45005\T-45160 | SRT, T-45005, Artery of neck |  | DCM, 122376, Total Plaque Volume |  | UCUM, mm3, mm3 | [7 empty] | SRT, T-45160, Carotid Bifurcation
Vascular Ultrasound Procedure Report | Vascular | Bifurcation Max Area % Reduction | Area Reduc | meas | m_Bifurc_MxAre%Red | r_AreaReduc_Percent | bilateral | 125100\G-0371\T-45005\T-45160 | SRT, T-45005, Artery of neck |  | SRT, G-0371, Percent Area Reduction |  | UCUM, %, Percent | [7 empty] | SRT, T-45160, Carotid Bifurcation
Vascular Ultrasound Procedure Report | Vascular | Bulb | ? | meas | m_Bulb | r_DopAngle | bilateral | 125100\C12122-01\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb IMT | Mean (IMT) | meas | m_Bulb_IMT | r_Mean_Dist | bilateral | 125100\C12122-08\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, mm, mm | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb IMT | SD (IMT) | meas | m_Bulb_IMT | r_SD_Dist | bilateral | 125100\C12122-09\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-09, Standard Deviation of Intima Media Thickness |  | UCUM, mm, mm | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb Total Plaque Vol | Volume | meas | m_Bulb_TotalPlqVol | r_Volume | bilateral | 125100\122376\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | DCM, 122376, Total Plaque Volume |  | UCUM, mm3, mm3 | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb | EDV | meas | m_Bulb | r_EDV | bilateral | 125100\11653-3\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb | MDV | meas | m_Bulb | r_MDV | bilateral | 125100\11665-7\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb | TAPV | meas | m_Bulb | r_TAPV | bilateral | 125100\11692-1\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb | PSV | meas | m_Bulb | r_PSV | bilateral | 125100\11726-7\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb | PI | meas | m_Bulb | r_PI | bilateral | 125100\12008-9\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb | RI | meas | m_Bulb | r_RI | bilateral | 125100\12023-8\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb | S/D | meas | m_Bulb | r_SD | bilateral | 125100\12144-2\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb | AI | meas | m_Bulb | r_AccelIndex | bilateral | 125100\20167-3\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb | AT | meas | m_Bulb | r_AccelTime | bilateral | 125100\20168-1\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb | DT | meas | m_Bulb | r_DecelTime | bilateral | 125100\20217-6\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb | PG | meas | m_Bulb | r_PG | bilateral | 125100\20247-3\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb | MG | meas | m_Bulb | r_PGmean | bilateral | 125100\20256-4\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb | TAMV | meas | m_Bulb | r_TAMV | bilateral | 125100\20352-1\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb | VTI | meas | m_Bulb | r_VTI | bilateral | 125100\20354-7\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb | HR | meas | m_Bulb | r_HeartRate | bilateral | 125100\8867-4\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb | Dist | meas | m_Bulb | r_Distance | bilateral | 125100\G-0364\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | Bulb Max Area % Reduction | Area Reduc | meas | m_Bulb_MxAre%Red | r_AreaReduc_Percent | bilateral | 125100\G-0371\T-45005\T-45170 | SRT, T-45005, Artery of neck |  | SRT, G-0371, Percent Area Reduction |  | UCUM, %, Percent | [7 empty] | SRT, T-45170, Carotid Bulb
Vascular Ultrasound Procedure Report | Vascular | ECA Total Plaque Vol | Volume | meas | m_ECA_TotalPlqVol | r_Volume | bilateral | 125100\122376\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | DCM, 122376, Total Plaque Volume |  | UCUM, mm3, mm3 | [7 empty] | SRT, T-45200, External Carotid Artery
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | ? | meas | m_Prox_ECA | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | EDV | meas | m_Prox_ECA | r_EDV | bilateral | 125100\11653-3\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | MDV | meas | m_Prox_ECA | r_MDV | bilateral | 125100\11665-7\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | TAPV | meas | m_Prox_ECA | r_TAPV | bilateral | 125100\11692-1\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | PSV | meas | m_Prox_ECA | r_PSV | bilateral | 125100\11726-7\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | PI | meas | m_Prox_ECA | r_PI | bilateral | 125100\12008-9\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | RI | meas | m_Prox_ECA | r_RI | bilateral | 125100\12023-8\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | S/D | meas | m_Prox_ECA | r_SD | bilateral | 125100\12144-2\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | AI | meas | m_Prox_ECA | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | AT | meas | m_Prox_ECA | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | DT | meas | m_Prox_ECA | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | PG | meas | m_Prox_ECA | r_PG | bilateral | 125100\20247-3\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | MG | meas | m_Prox_ECA | r_PGmean | bilateral | 125100\20256-4\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | TAMV | meas | m_Prox_ECA | r_TAMV | bilateral | 125100\20352-1\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | VTI | meas | m_Prox_ECA | r_VTI | bilateral | 125100\20354-7\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | HR | meas | m_Prox_ECA | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA | Dist | meas | m_Prox_ECA | r_Distance | bilateral | 125100\G-0364\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA Max Area % Reduction | Area Reduc | meas | m_Prox_ECA_MxAre%Red | r_AreaReduc_Percent | bilateral | 125100\G-0371\G-A118\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | SRT, G-0371, Percent Area Reduction |  | UCUM, %, Percent | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist ECA Max Area % Reduction | Area Reduc | meas | m_Dist_ECA_MxAre%Red | r_AreaReduc_Percent | bilateral | 125100\G-0371\G-A119\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | SRT, G-0371, Percent Area Reduction |  | UCUM, %, Percent | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid ECA Max Area % Reduction | Area Reduc | meas | m_Mid_ECA_MxAre%Red | r_AreaReduc_Percent | bilateral | 125100\G-0371\G-A188\T-45005\T-45200 | SRT, T-45005, Artery of neck |  | SRT, G-0371, Percent Area Reduction |  | UCUM, %, Percent | [7 empty] | SRT, T-45200, External Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | ICA Total Plaque Vol | Volume | meas | m_ICA_TotalPlqVol | r_Volume | bilateral | 125100\122376\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | DCM, 122376, Total Plaque Volume |  | UCUM, mm3, mm3 | [7 empty] | SRT, T-45300, Internal Carotid Artery
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | ? | meas | m_Prox_ICA | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA IMT | Mean (IMT) | meas | m_Prox_ICA_IMT | r_Mean_Dist | bilateral | 125100\C12122-08\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, mm, mm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA IMT | SD (IMT) | meas | m_Prox_ICA_IMT | r_SD_Dist | bilateral | 125100\C12122-09\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-09, Standard Deviation of Intima Media Thickness |  | UCUM, mm, mm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | EDV | meas | m_Prox_ICA | r_EDV | bilateral | 125100\11653-3\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | MDV | meas | m_Prox_ICA | r_MDV | bilateral | 125100\11665-7\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | TAPV | meas | m_Prox_ICA | r_TAPV | bilateral | 125100\11692-1\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | PSV | meas | m_Prox_ICA | r_PSV | bilateral | 125100\11726-7\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | PI | meas | m_Prox_ICA | r_PI | bilateral | 125100\12008-9\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | RI | meas | m_Prox_ICA | r_RI | bilateral | 125100\12023-8\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | S/D | meas | m_Prox_ICA | r_SD | bilateral | 125100\12144-2\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | AI | meas | m_Prox_ICA | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | AT | meas | m_Prox_ICA | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | DT | meas | m_Prox_ICA | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | PG | meas | m_Prox_ICA | r_PG | bilateral | 125100\20247-3\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | MG | meas | m_Prox_ICA | r_PGmean | bilateral | 125100\20256-4\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | TAMV | meas | m_Prox_ICA | r_TAMV | bilateral | 125100\20352-1\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | VTI | meas | m_Prox_ICA | r_VTI | bilateral | 125100\20354-7\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | HR | meas | m_Prox_ICA | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA | Dist | meas | m_Prox_ICA | r_Distance | bilateral | 125100\G-0364\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA Max Area % Reduction | Area Reduc | meas | m_Prox_ICA_MxAre%Red | r_AreaReduc_Percent | bilateral | 125100\G-0371\G-A118\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | SRT, G-0371, Percent Area Reduction |  | UCUM, %, Percent | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | ? | meas | m_Dist_ICA | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA IMT | Mean (IMT) | meas | m_Dist_ICA_IMT | r_Mean_Dist | bilateral | 125100\C12122-08\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, mm, mm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA IMT | SD (IMT) | meas | m_Dist_ICA_IMT | r_SD_Dist | bilateral | 125100\C12122-09\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-09, Standard Deviation of Intima Media Thickness |  | UCUM, mm, mm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | EDV | meas | m_Dist_ICA | r_EDV | bilateral | 125100\11653-3\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | MDV | meas | m_Dist_ICA | r_MDV | bilateral | 125100\11665-7\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | TAPV | meas | m_Dist_ICA | r_TAPV | bilateral | 125100\11692-1\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | PSV | meas | m_Dist_ICA | r_PSV | bilateral | 125100\11726-7\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | PI | meas | m_Dist_ICA | r_PI | bilateral | 125100\12008-9\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | RI | meas | m_Dist_ICA | r_RI | bilateral | 125100\12023-8\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | S/D | meas | m_Dist_ICA | r_SD | bilateral | 125100\12144-2\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | AI | meas | m_Dist_ICA | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | AT | meas | m_Dist_ICA | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | DT | meas | m_Dist_ICA | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | PG | meas | m_Dist_ICA | r_PG | bilateral | 125100\20247-3\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | MG | meas | m_Dist_ICA | r_PGmean | bilateral | 125100\20256-4\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | TAMV | meas | m_Dist_ICA | r_TAMV | bilateral | 125100\20352-1\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | VTI | meas | m_Dist_ICA | r_VTI | bilateral | 125100\20354-7\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | HR | meas | m_Dist_ICA | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA | Dist | meas | m_Dist_ICA | r_Distance | bilateral | 125100\G-0364\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA Max Area % Reduction | Area Reduc | meas | m_Dist_ICA_MxAre%Red | r_AreaReduc_Percent | bilateral | 125100\G-0371\G-A119\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | SRT, G-0371, Percent Area Reduction |  | UCUM, %, Percent | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | ? | meas | m_Mid_ICA | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA IMT | Mean (IMT) | meas | m_Mid_ICA_IMT | r_Mean_Dist | bilateral | 125100\C12122-08\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, mm, mm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA IMT | SD (IMT) | meas | m_Mid_ICA_IMT | r_SD_Dist | bilateral | 125100\C12122-09\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-09, Standard Deviation of Intima Media Thickness |  | UCUM, mm, mm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | EDV | meas | m_Mid_ICA | r_EDV | bilateral | 125100\11653-3\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | MDV | meas | m_Mid_ICA | r_MDV | bilateral | 125100\11665-7\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | TAPV | meas | m_Mid_ICA | r_TAPV | bilateral | 125100\11692-1\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | PSV | meas | m_Mid_ICA | r_PSV | bilateral | 125100\11726-7\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | PI | meas | m_Mid_ICA | r_PI | bilateral | 125100\12008-9\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | RI | meas | m_Mid_ICA | r_RI | bilateral | 125100\12023-8\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | S/D | meas | m_Mid_ICA | r_SD | bilateral | 125100\12144-2\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | AI | meas | m_Mid_ICA | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | AT | meas | m_Mid_ICA | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | DT | meas | m_Mid_ICA | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | PG | meas | m_Mid_ICA | r_PG | bilateral | 125100\20247-3\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | MG | meas | m_Mid_ICA | r_PGmean | bilateral | 125100\20256-4\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | TAMV | meas | m_Mid_ICA | r_TAMV | bilateral | 125100\20352-1\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | VTI | meas | m_Mid_ICA | r_VTI | bilateral | 125100\20354-7\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | HR | meas | m_Mid_ICA | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA | Dist | meas | m_Mid_ICA | r_Distance | bilateral | 125100\G-0364\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA Max Area % Reduction | Area Reduc | meas | m_Mid_ICA_MxAre%Red | r_AreaReduc_Percent | bilateral | 125100\G-0371\G-A188\T-45005\T-45300 | SRT, T-45005, Artery of neck |  | SRT, G-0371, Percent Area Reduction |  | UCUM, %, Percent | [7 empty] | SRT, T-45300, Internal Carotid Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | ? | meas | m_Vertebral_A | r_DopAngle | bilateral | 125100\C12122-01\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | EDV | meas | m_Vertebral_A | r_EDV | bilateral | 125100\11653-3\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | MDV | meas | m_Vertebral_A | r_MDV | bilateral | 125100\11665-7\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | TAPV | meas | m_Vertebral_A | r_TAPV | bilateral | 125100\11692-1\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | PSV | meas | m_Vertebral_A | r_PSV | bilateral | 125100\11726-7\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | PI | meas | m_Vertebral_A | r_PI | bilateral | 125100\12008-9\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | RI | meas | m_Vertebral_A | r_RI | bilateral | 125100\12023-8\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | S/D | meas | m_Vertebral_A | r_SD | bilateral | 125100\12144-2\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | AI | meas | m_Vertebral_A | r_AccelIndex | bilateral | 125100\20167-3\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | AT | meas | m_Vertebral_A | r_AccelTime | bilateral | 125100\20168-1\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | DT | meas | m_Vertebral_A | r_DecelTime | bilateral | 125100\20217-6\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | PG | meas | m_Vertebral_A | r_PG | bilateral | 125100\20247-3\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | MG | meas | m_Vertebral_A | r_PGmean | bilateral | 125100\20256-4\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | TAMV | meas | m_Vertebral_A | r_TAMV | bilateral | 125100\20352-1\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | VTI | meas | m_Vertebral_A | r_VTI | bilateral | 125100\20354-7\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | HR | meas | m_Vertebral_A | r_HeartRate | bilateral | 125100\8867-4\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Vertebral A | Dist | meas | m_Vertebral_A | r_Distance | bilateral | 125100\G-0364\T-45005\T-45700 | SRT, T-45005, Artery of neck |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-45700, Vertebral Artery
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | ? | meas | m_Prox_SCL_A | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | EDV | meas | m_Prox_SCL_A | r_EDV | bilateral | 125100\11653-3\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | MDV | meas | m_Prox_SCL_A | r_MDV | bilateral | 125100\11665-7\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | TAPV | meas | m_Prox_SCL_A | r_TAPV | bilateral | 125100\11692-1\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | PSV | meas | m_Prox_SCL_A | r_PSV | bilateral | 125100\11726-7\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | PI | meas | m_Prox_SCL_A | r_PI | bilateral | 125100\12008-9\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | RI | meas | m_Prox_SCL_A | r_RI | bilateral | 125100\12023-8\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | S/D | meas | m_Prox_SCL_A | r_SD | bilateral | 125100\12144-2\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | AI | meas | m_Prox_SCL_A | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | AT | meas | m_Prox_SCL_A | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | DT | meas | m_Prox_SCL_A | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | PG | meas | m_Prox_SCL_A | r_PG | bilateral | 125100\20247-3\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | MG | meas | m_Prox_SCL_A | r_PGmean | bilateral | 125100\20256-4\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | TAMV | meas | m_Prox_SCL_A | r_TAMV | bilateral | 125100\20352-1\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | VTI | meas | m_Prox_SCL_A | r_VTI | bilateral | 125100\20354-7\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | HR | meas | m_Prox_SCL_A | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A | Dist | meas | m_Prox_SCL_A | r_Distance | bilateral | 125100\G-0364\G-A118\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | ? | meas | m_Mid_SCL_A | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | EDV | meas | m_Mid_SCL_A | r_EDV | bilateral | 125100\11653-3\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | MDV | meas | m_Mid_SCL_A | r_MDV | bilateral | 125100\11665-7\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | TAPV | meas | m_Mid_SCL_A | r_TAPV | bilateral | 125100\11692-1\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | PSV | meas | m_Mid_SCL_A | r_PSV | bilateral | 125100\11726-7\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | PI | meas | m_Mid_SCL_A | r_PI | bilateral | 125100\12008-9\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | RI | meas | m_Mid_SCL_A | r_RI | bilateral | 125100\12023-8\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | S/D | meas | m_Mid_SCL_A | r_SD | bilateral | 125100\12144-2\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | AI | meas | m_Mid_SCL_A | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | AT | meas | m_Mid_SCL_A | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | DT | meas | m_Mid_SCL_A | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | PG | meas | m_Mid_SCL_A | r_PG | bilateral | 125100\20247-3\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | MG | meas | m_Mid_SCL_A | r_PGmean | bilateral | 125100\20256-4\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | TAMV | meas | m_Mid_SCL_A | r_TAMV | bilateral | 125100\20352-1\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | VTI | meas | m_Mid_SCL_A | r_VTI | bilateral | 125100\20354-7\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | HR | meas | m_Mid_SCL_A | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A | Dist | meas | m_Mid_SCL_A | r_Distance | bilateral | 125100\G-0364\G-A188\T-45005\T-46100 | SRT, T-45005, Artery of neck |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46100, Subclavian Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Aorta PSV | ? | meas | m_aorta_PSV | r_DopAngle | unilateral | 125100\C12122-01\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-42000, Aorta
Vascular Ultrasound Procedure Report | Abdominal | Aorta PSV | PSV | meas | m_aorta_PSV | r_PSV | unilateral | 125100\11726-7\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-42000, Aorta
Vascular Ultrasound Procedure Report | Abdominal | Ao Prox Diam | Dist | meas | m_aoProx | r_Distance | unilateral | 125100\G-0364\G-A118\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | ? | meas | m_distaorta | r_DopAngle | unilateral | 125100\C12122-01\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | EDV | meas | m_distaorta | r_EDV | unilateral | 125100\11653-3\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | MDV | meas | m_distaorta | r_MDV | unilateral | 125100\11665-7\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | TAPV | meas | m_distaorta | r_TAPV | unilateral | 125100\11692-1\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | PSV | meas | m_distaorta | r_PSV | unilateral | 125100\11726-7\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | PI | meas | m_distaorta | r_PI | unilateral | 125100\12008-9\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | RI | meas | m_distaorta | r_RI | unilateral | 125100\12023-8\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | S/D | meas | m_distaorta | r_SD | unilateral | 125100\12144-2\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | AI | meas | m_distaorta | r_AccelIndex | unilateral | 125100\20167-3\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | AT | meas | m_distaorta | r_AccelTime | unilateral | 125100\20168-1\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | DT | meas | m_distaorta | r_DecelTime | unilateral | 125100\20217-6\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | PG | meas | m_distaorta | r_PG | unilateral | 125100\20247-3\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | MG | meas | m_distaorta | r_PGmean | unilateral | 125100\20256-4\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | TAMV | meas | m_distaorta | r_TAMV | unilateral | 125100\20352-1\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | VTI | meas | m_distaorta | r_VTI | unilateral | 125100\20354-7\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | HR | meas | m_distaorta | r_HeartRate | unilateral | 125100\8867-4\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Ao Dist Diam | Dist | meas | m_aoDist | r_Distance | unilateral | 125100\G-0364\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ao | Dist | meas | m_distaorta | r_Distance | unilateral | 125100\G-0364\G-A119\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Ao Mid Diam | Dist | meas | m_aoMid | r_Distance | unilateral | 125100\G-0364\G-A188\T-42000\T-46002 | SRT, T-46002, Artery of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-42000, Aorta |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | ? | meas | m_suprenaorta | r_DopAngle | unilateral | 125100\C12122-01\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | EDV | meas | m_suprenaorta | r_EDV | unilateral | 125100\11653-3\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | MDV | meas | m_suprenaorta | r_MDV | unilateral | 125100\11665-7\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | TAPV | meas | m_suprenaorta | r_TAPV | unilateral | 125100\11692-1\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | PSV | meas | m_suprenaorta | r_PSV | unilateral | 125100\11726-7\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | PI | meas | m_suprenaorta | r_PI | unilateral | 125100\12008-9\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | RI | meas | m_suprenaorta | r_RI | unilateral | 125100\12023-8\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | S/D | meas | m_suprenaorta | r_SD | unilateral | 125100\12144-2\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | AI | meas | m_suprenaorta | r_AccelIndex | unilateral | 125100\20167-3\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | AT | meas | m_suprenaorta | r_AccelTime | unilateral | 125100\20168-1\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | DT | meas | m_suprenaorta | r_DecelTime | unilateral | 125100\20217-6\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | PG | meas | m_suprenaorta | r_PG | unilateral | 125100\20247-3\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | MG | meas | m_suprenaorta | r_PGmean | unilateral | 125100\20256-4\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | TAMV | meas | m_suprenaorta | r_TAMV | unilateral | 125100\20352-1\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | VTI | meas | m_suprenaorta | r_VTI | unilateral | 125100\20354-7\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | HR | meas | m_suprenaorta | r_HeartRate | unilateral | 125100\8867-4\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Sup Ren Ao | Dist | meas | m_suprenaorta | r_Distance | unilateral | 125100\G-0364\T-42510\T-46002 | SRT, T-46002, Artery of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-42510, Supra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | ? | meas | m_infrenaorta | r_DopAngle | unilateral | 125100\C12122-01\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | EDV | meas | m_infrenaorta | r_EDV | unilateral | 125100\11653-3\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | MDV | meas | m_infrenaorta | r_MDV | unilateral | 125100\11665-7\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | TAPV | meas | m_infrenaorta | r_TAPV | unilateral | 125100\11692-1\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | PSV | meas | m_infrenaorta | r_PSV | unilateral | 125100\11726-7\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | PI | meas | m_infrenaorta | r_PI | unilateral | 125100\12008-9\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | RI | meas | m_infrenaorta | r_RI | unilateral | 125100\12023-8\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | S/D | meas | m_infrenaorta | r_SD | unilateral | 125100\12144-2\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | AI | meas | m_infrenaorta | r_AccelIndex | unilateral | 125100\20167-3\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | AT | meas | m_infrenaorta | r_AccelTime | unilateral | 125100\20168-1\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | DT | meas | m_infrenaorta | r_DecelTime | unilateral | 125100\20217-6\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | PG | meas | m_infrenaorta | r_PG | unilateral | 125100\20247-3\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | MG | meas | m_infrenaorta | r_PGmean | unilateral | 125100\20256-4\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | TAMV | meas | m_infrenaorta | r_TAMV | unilateral | 125100\20352-1\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | VTI | meas | m_infrenaorta | r_VTI | unilateral | 125100\20354-7\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | HR | meas | m_infrenaorta | r_HeartRate | unilateral | 125100\8867-4\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Inf Ren Ao | Dist | meas | m_infrenaorta | r_Distance | unilateral | 125100\G-0364\T-42520\T-46002 | SRT, T-46002, Artery of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-42520, Infra-renal Aorta
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | ? | meas | m_celiacaxis | r_DopAngle | unilateral | 125100\C12122-01\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | EDV | meas | m_celiacaxis | r_EDV | unilateral | 125100\11653-3\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | MDV | meas | m_celiacaxis | r_MDV | unilateral | 125100\11665-7\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | TAPV | meas | m_celiacaxis | r_TAPV | unilateral | 125100\11692-1\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | PSV | meas | m_celiacaxis | r_PSV | unilateral | 125100\11726-7\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | PI | meas | m_celiacaxis | r_PI | unilateral | 125100\12008-9\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | RI | meas | m_celiacaxis | r_RI | unilateral | 125100\12023-8\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | S/D | meas | m_celiacaxis | r_SD | unilateral | 125100\12144-2\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | AI | meas | m_celiacaxis | r_AccelIndex | unilateral | 125100\20167-3\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | AT | meas | m_celiacaxis | r_AccelTime | unilateral | 125100\20168-1\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | DT | meas | m_celiacaxis | r_DecelTime | unilateral | 125100\20217-6\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | PG | meas | m_celiacaxis | r_PG | unilateral | 125100\20247-3\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | MG | meas | m_celiacaxis | r_PGmean | unilateral | 125100\20256-4\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | TAMV | meas | m_celiacaxis | r_TAMV | unilateral | 125100\20352-1\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | VTI | meas | m_celiacaxis | r_VTI | unilateral | 125100\20354-7\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | HR | meas | m_celiacaxis | r_HeartRate | unilateral | 125100\8867-4\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Celiac A | Dist | meas | m_celiacaxis | r_Distance | unilateral | 125100\G-0364\T-46002\T-46400 | SRT, T-46002, Artery of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46400, Celiac Axis
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | ? | meas | m_comhepart | r_DopAngle | unilateral | 125100\C12122-01\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | EDV | meas | m_comhepart | r_EDV | unilateral | 125100\11653-3\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | MDV | meas | m_comhepart | r_MDV | unilateral | 125100\11665-7\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | TAPV | meas | m_comhepart | r_TAPV | unilateral | 125100\11692-1\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | PSV | meas | m_comhepart | r_PSV | unilateral | 125100\11726-7\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | PI | meas | m_comhepart | r_PI | unilateral | 125100\12008-9\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | RI | meas | m_comhepart | r_RI | unilateral | 125100\12023-8\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | S/D | meas | m_comhepart | r_SD | unilateral | 125100\12144-2\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | AI | meas | m_comhepart | r_AccelIndex | unilateral | 125100\20167-3\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | AT | meas | m_comhepart | r_AccelTime | unilateral | 125100\20168-1\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | DT | meas | m_comhepart | r_DecelTime | unilateral | 125100\20217-6\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | PG | meas | m_comhepart | r_PG | unilateral | 125100\20247-3\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | MG | meas | m_comhepart | r_PGmean | unilateral | 125100\20256-4\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | TAMV | meas | m_comhepart | r_TAMV | unilateral | 125100\20352-1\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | VTI | meas | m_comhepart | r_VTI | unilateral | 125100\20354-7\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | HR | meas | m_comhepart | r_HeartRate | unilateral | 125100\8867-4\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | Com Hepatic A | Dist | meas | m_comhepart | r_Distance | unilateral | 125100\G-0364\T-46002\T-46421 | SRT, T-46002, Artery of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46421, Common Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | ? | meas | m_r_hepart | r_DopAngle | unilateral | 125100\C12122-01\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | EDV | meas | m_r_hepart | r_EDV | unilateral | 125100\11653-3\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | MDV | meas | m_r_hepart | r_MDV | unilateral | 125100\11665-7\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | TAPV | meas | m_r_hepart | r_TAPV | unilateral | 125100\11692-1\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | PSV | meas | m_r_hepart | r_PSV | unilateral | 125100\11726-7\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | PI | meas | m_r_hepart | r_PI | unilateral | 125100\12008-9\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | RI | meas | m_r_hepart | r_RI | unilateral | 125100\12023-8\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | S/D | meas | m_r_hepart | r_SD | unilateral | 125100\12144-2\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | AI | meas | m_r_hepart | r_AccelIndex | unilateral | 125100\20167-3\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | AT | meas | m_r_hepart | r_AccelTime | unilateral | 125100\20168-1\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | DT | meas | m_r_hepart | r_DecelTime | unilateral | 125100\20217-6\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | PG | meas | m_r_hepart | r_PG | unilateral | 125100\20247-3\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | MG | meas | m_r_hepart | r_PGmean | unilateral | 125100\20256-4\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | TAMV | meas | m_r_hepart | r_TAMV | unilateral | 125100\20352-1\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | VTI | meas | m_r_hepart | r_VTI | unilateral | 125100\20354-7\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | HR | meas | m_r_hepart | r_HeartRate | unilateral | 125100\8867-4\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic A | Dist | meas | m_r_hepart | r_Distance | unilateral | 125100\G-0364\T-46002\T-46423 | SRT, T-46002, Artery of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46423, Right Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | ? | meas | m_l_hepart | r_DopAngle | unilateral | 125100\C12122-01\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | EDV | meas | m_l_hepart | r_EDV | unilateral | 125100\11653-3\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | MDV | meas | m_l_hepart | r_MDV | unilateral | 125100\11665-7\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | TAPV | meas | m_l_hepart | r_TAPV | unilateral | 125100\11692-1\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | PSV | meas | m_l_hepart | r_PSV | unilateral | 125100\11726-7\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | PI | meas | m_l_hepart | r_PI | unilateral | 125100\12008-9\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | RI | meas | m_l_hepart | r_RI | unilateral | 125100\12023-8\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | S/D | meas | m_l_hepart | r_SD | unilateral | 125100\12144-2\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | AI | meas | m_l_hepart | r_AccelIndex | unilateral | 125100\20167-3\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | AT | meas | m_l_hepart | r_AccelTime | unilateral | 125100\20168-1\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | DT | meas | m_l_hepart | r_DecelTime | unilateral | 125100\20217-6\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | PG | meas | m_l_hepart | r_PG | unilateral | 125100\20247-3\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | MG | meas | m_l_hepart | r_PGmean | unilateral | 125100\20256-4\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | TAMV | meas | m_l_hepart | r_TAMV | unilateral | 125100\20352-1\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | VTI | meas | m_l_hepart | r_VTI | unilateral | 125100\20354-7\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | HR | meas | m_l_hepart | r_HeartRate | unilateral | 125100\8867-4\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic A | Dist | meas | m_l_hepart | r_Distance | unilateral | 125100\G-0364\T-46002\T-46427 | SRT, T-46002, Artery of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46427, Left Branch of Hepatic Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | ? | meas | m_gda | r_DopAngle | unilateral | 125100\C12122-01\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | EDV | meas | m_gda | r_EDV | unilateral | 125100\11653-3\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | MDV | meas | m_gda | r_MDV | unilateral | 125100\11665-7\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | TAPV | meas | m_gda | r_TAPV | unilateral | 125100\11692-1\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | PSV | meas | m_gda | r_PSV | unilateral | 125100\11726-7\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | PI | meas | m_gda | r_PI | unilateral | 125100\12008-9\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | RI | meas | m_gda | r_RI | unilateral | 125100\12023-8\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | S/D | meas | m_gda | r_SD | unilateral | 125100\12144-2\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | AI | meas | m_gda | r_AccelIndex | unilateral | 125100\20167-3\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | AT | meas | m_gda | r_AccelTime | unilateral | 125100\20168-1\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | DT | meas | m_gda | r_DecelTime | unilateral | 125100\20217-6\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | PG | meas | m_gda | r_PG | unilateral | 125100\20247-3\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | MG | meas | m_gda | r_PGmean | unilateral | 125100\20256-4\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | TAMV | meas | m_gda | r_TAMV | unilateral | 125100\20352-1\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | VTI | meas | m_gda | r_VTI | unilateral | 125100\20354-7\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | HR | meas | m_gda | r_HeartRate | unilateral | 125100\8867-4\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | GDA | Dist | meas | m_gda | r_Distance | unilateral | 125100\G-0364\T-46002\T-46440 | SRT, T-46002, Artery of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46440, Gastroduodenal Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | ? | meas | m_splenicart | r_DopAngle | unilateral | 125100\C12122-01\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | EDV | meas | m_splenicart | r_EDV | unilateral | 125100\11653-3\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | MDV | meas | m_splenicart | r_MDV | unilateral | 125100\11665-7\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | TAPV | meas | m_splenicart | r_TAPV | unilateral | 125100\11692-1\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | PSV | meas | m_splenicart | r_PSV | unilateral | 125100\11726-7\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | PI | meas | m_splenicart | r_PI | unilateral | 125100\12008-9\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | RI | meas | m_splenicart | r_RI | unilateral | 125100\12023-8\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | S/D | meas | m_splenicart | r_SD | unilateral | 125100\12144-2\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | AI | meas | m_splenicart | r_AccelIndex | unilateral | 125100\20167-3\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | AT | meas | m_splenicart | r_AccelTime | unilateral | 125100\20168-1\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | DT | meas | m_splenicart | r_DecelTime | unilateral | 125100\20217-6\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | PG | meas | m_splenicart | r_PG | unilateral | 125100\20247-3\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | MG | meas | m_splenicart | r_PGmean | unilateral | 125100\20256-4\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | TAMV | meas | m_splenicart | r_TAMV | unilateral | 125100\20352-1\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | VTI | meas | m_splenicart | r_VTI | unilateral | 125100\20354-7\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | HR | meas | m_splenicart | r_HeartRate | unilateral | 125100\8867-4\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Splenic A | Dist | meas | m_splenicart | r_Distance | unilateral | 125100\G-0364\T-46002\T-46460 | SRT, T-46002, Artery of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46460, Splenic Artery
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | ? | meas | m_proxsma | r_DopAngle | unilateral | 125100\C12122-01\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | EDV | meas | m_proxsma | r_EDV | unilateral | 125100\11653-3\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | MDV | meas | m_proxsma | r_MDV | unilateral | 125100\11665-7\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | TAPV | meas | m_proxsma | r_TAPV | unilateral | 125100\11692-1\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | PSV | meas | m_proxsma | r_PSV | unilateral | 125100\11726-7\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | PI | meas | m_proxsma | r_PI | unilateral | 125100\12008-9\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | RI | meas | m_proxsma | r_RI | unilateral | 125100\12023-8\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | S/D | meas | m_proxsma | r_SD | unilateral | 125100\12144-2\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | AI | meas | m_proxsma | r_AccelIndex | unilateral | 125100\20167-3\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | AT | meas | m_proxsma | r_AccelTime | unilateral | 125100\20168-1\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | DT | meas | m_proxsma | r_DecelTime | unilateral | 125100\20217-6\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | PG | meas | m_proxsma | r_PG | unilateral | 125100\20247-3\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | MG | meas | m_proxsma | r_PGmean | unilateral | 125100\20256-4\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | TAMV | meas | m_proxsma | r_TAMV | unilateral | 125100\20352-1\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | VTI | meas | m_proxsma | r_VTI | unilateral | 125100\20354-7\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | HR | meas | m_proxsma | r_HeartRate | unilateral | 125100\8867-4\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox SMA | Dist | meas | m_proxsma | r_Distance | unilateral | 125100\G-0364\G-A118\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | ? | meas | m_distsma | r_DopAngle | unilateral | 125100\C12122-01\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | EDV | meas | m_distsma | r_EDV | unilateral | 125100\11653-3\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | MDV | meas | m_distsma | r_MDV | unilateral | 125100\11665-7\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | TAPV | meas | m_distsma | r_TAPV | unilateral | 125100\11692-1\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | PSV | meas | m_distsma | r_PSV | unilateral | 125100\11726-7\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | PI | meas | m_distsma | r_PI | unilateral | 125100\12008-9\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | RI | meas | m_distsma | r_RI | unilateral | 125100\12023-8\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | S/D | meas | m_distsma | r_SD | unilateral | 125100\12144-2\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | AI | meas | m_distsma | r_AccelIndex | unilateral | 125100\20167-3\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | AT | meas | m_distsma | r_AccelTime | unilateral | 125100\20168-1\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | DT | meas | m_distsma | r_DecelTime | unilateral | 125100\20217-6\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | PG | meas | m_distsma | r_PG | unilateral | 125100\20247-3\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | MG | meas | m_distsma | r_PGmean | unilateral | 125100\20256-4\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | TAMV | meas | m_distsma | r_TAMV | unilateral | 125100\20352-1\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | VTI | meas | m_distsma | r_VTI | unilateral | 125100\20354-7\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | HR | meas | m_distsma | r_HeartRate | unilateral | 125100\8867-4\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist SMA | Dist | meas | m_distsma | r_Distance | unilateral | 125100\G-0364\G-A119\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | ? | meas | m_midsma | r_DopAngle | unilateral | 125100\C12122-01\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | EDV | meas | m_midsma | r_EDV | unilateral | 125100\11653-3\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | MDV | meas | m_midsma | r_MDV | unilateral | 125100\11665-7\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | TAPV | meas | m_midsma | r_TAPV | unilateral | 125100\11692-1\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | PSV | meas | m_midsma | r_PSV | unilateral | 125100\11726-7\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | PI | meas | m_midsma | r_PI | unilateral | 125100\12008-9\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | RI | meas | m_midsma | r_RI | unilateral | 125100\12023-8\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | S/D | meas | m_midsma | r_SD | unilateral | 125100\12144-2\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | AI | meas | m_midsma | r_AccelIndex | unilateral | 125100\20167-3\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | AT | meas | m_midsma | r_AccelTime | unilateral | 125100\20168-1\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | DT | meas | m_midsma | r_DecelTime | unilateral | 125100\20217-6\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | PG | meas | m_midsma | r_PG | unilateral | 125100\20247-3\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | MG | meas | m_midsma | r_PGmean | unilateral | 125100\20256-4\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | TAMV | meas | m_midsma | r_TAMV | unilateral | 125100\20352-1\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | VTI | meas | m_midsma | r_VTI | unilateral | 125100\20354-7\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | HR | meas | m_midsma | r_HeartRate | unilateral | 125100\8867-4\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid SMA | Dist | meas | m_midsma | r_Distance | unilateral | 125100\G-0364\G-A188\T-46002\T-46510 | SRT, T-46002, Artery of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | IMA | ? | meas | m_ima | r_DopAngle | unilateral | 125100\C12122-01\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Abdominal | IMA | EDV | meas | m_ima | r_EDV | unilateral | 125100\11653-3\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Abdominal | IMA | MDV | meas | m_ima | r_MDV | unilateral | 125100\11665-7\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Abdominal | IMA | TAPV | meas | m_ima | r_TAPV | unilateral | 125100\11692-1\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Abdominal | IMA | PSV | meas | m_ima | r_PSV | unilateral | 125100\11726-7\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Abdominal | IMA | PI | meas | m_ima | r_PI | unilateral | 125100\12008-9\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Abdominal | IMA | RI | meas | m_ima | r_RI | unilateral | 125100\12023-8\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Abdominal | IMA | S/D | meas | m_ima | r_SD | unilateral | 125100\12144-2\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Abdominal | IMA | AI | meas | m_ima | r_AccelIndex | unilateral | 125100\20167-3\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Abdominal | IMA | AT | meas | m_ima | r_AccelTime | unilateral | 125100\20168-1\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Abdominal | IMA | DT | meas | m_ima | r_DecelTime | unilateral | 125100\20217-6\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Abdominal | IMA | PG | meas | m_ima | r_PG | unilateral | 125100\20247-3\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Abdominal | IMA | MG | meas | m_ima | r_PGmean | unilateral | 125100\20256-4\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Abdominal | IMA | TAMV | meas | m_ima | r_TAMV | unilateral | 125100\20352-1\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Abdominal | IMA | VTI | meas | m_ima | r_VTI | unilateral | 125100\20354-7\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Abdominal | IMA | HR | meas | m_ima | r_HeartRate | unilateral | 125100\8867-4\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Abdominal | IMA | Dist | meas | m_ima | r_Distance | unilateral | 125100\G-0364\T-46002\T-46520 | SRT, T-46002, Artery of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46520, Inferior Mesenteric Artery
Vascular Ultrasound Procedure Report | Vascular | Antecube | ? | meas | m_Antecube | r_DopAngle | bilateral | 125100\C12122-01\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Antecube | EDV | meas | m_Antecube | r_EDV | bilateral | 125100\11653-3\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Antecube | MDV | meas | m_Antecube | r_MDV | bilateral | 125100\11665-7\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Antecube | TAPV | meas | m_Antecube | r_TAPV | bilateral | 125100\11692-1\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Antecube | PSV | meas | m_Antecube | r_PSV | bilateral | 125100\11726-7\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Antecube | PI | meas | m_Antecube | r_PI | bilateral | 125100\12008-9\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Antecube | RI | meas | m_Antecube | r_RI | bilateral | 125100\12023-8\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Antecube | S/D | meas | m_Antecube | r_SD | bilateral | 125100\12144-2\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Antecube | AI | meas | m_Antecube | r_AccelIndex | bilateral | 125100\20167-3\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Antecube | AT | meas | m_Antecube | r_AccelTime | bilateral | 125100\20168-1\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Antecube | DT | meas | m_Antecube | r_DecelTime | bilateral | 125100\20217-6\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Antecube | PG | meas | m_Antecube | r_PG | bilateral | 125100\20247-3\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Antecube | MG | meas | m_Antecube | r_PGmean | bilateral | 125100\20256-4\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Antecube | TAMV | meas | m_Antecube | r_TAMV | bilateral | 125100\20352-1\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Antecube | VTI | meas | m_Antecube | r_VTI | bilateral | 125100\20354-7\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Antecube | HR | meas | m_Antecube | r_HeartRate | bilateral | 125100\8867-4\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Antecube | Dist | meas | m_Antecube | r_Distance | bilateral | 125100\G-0364\sup71_001\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | 99PMSBLUS, sup71_001, Antecube
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | ? | meas | m_Brach_Ceph_A | r_DopAngle | bilateral | 125100\C12122-01\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | EDV | meas | m_Brach_Ceph_A | r_EDV | bilateral | 125100\11653-3\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | MDV | meas | m_Brach_Ceph_A | r_MDV | bilateral | 125100\11665-7\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | TAPV | meas | m_Brach_Ceph_A | r_TAPV | bilateral | 125100\11692-1\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | PSV | meas | m_Brach_Ceph_A | r_PSV | bilateral | 125100\11726-7\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | PI | meas | m_Brach_Ceph_A | r_PI | bilateral | 125100\12008-9\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | RI | meas | m_Brach_Ceph_A | r_RI | bilateral | 125100\12023-8\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | S/D | meas | m_Brach_Ceph_A | r_SD | bilateral | 125100\12144-2\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | AI | meas | m_Brach_Ceph_A | r_AccelIndex | bilateral | 125100\20167-3\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | AT | meas | m_Brach_Ceph_A | r_AccelTime | bilateral | 125100\20168-1\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | DT | meas | m_Brach_Ceph_A | r_DecelTime | bilateral | 125100\20217-6\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | PG | meas | m_Brach_Ceph_A | r_PG | bilateral | 125100\20247-3\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | MG | meas | m_Brach_Ceph_A | r_PGmean | bilateral | 125100\20256-4\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | TAMV | meas | m_Brach_Ceph_A | r_TAMV | bilateral | 125100\20352-1\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | VTI | meas | m_Brach_Ceph_A | r_VTI | bilateral | 125100\20354-7\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | HR | meas | m_Brach_Ceph_A | r_HeartRate | bilateral | 125100\8867-4\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A | Dist | meas | m_Brach_Ceph_A | r_Distance | bilateral | 125100\G-0364\T-46010\T-47020 | SRT, T-47020, Artery Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46010, Innominate Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | ? | meas | m_Axillary_A | r_DopAngle | bilateral | 125100\C12122-01\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | EDV | meas | m_Axillary_A | r_EDV | bilateral | 125100\11653-3\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | MDV | meas | m_Axillary_A | r_MDV | bilateral | 125100\11665-7\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | TAPV | meas | m_Axillary_A | r_TAPV | bilateral | 125100\11692-1\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | PSV | meas | m_Axillary_A | r_PSV | bilateral | 125100\11726-7\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | PI | meas | m_Axillary_A | r_PI | bilateral | 125100\12008-9\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | RI | meas | m_Axillary_A | r_RI | bilateral | 125100\12023-8\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | S/D | meas | m_Axillary_A | r_SD | bilateral | 125100\12144-2\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | AI | meas | m_Axillary_A | r_AccelIndex | bilateral | 125100\20167-3\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | AT | meas | m_Axillary_A | r_AccelTime | bilateral | 125100\20168-1\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | DT | meas | m_Axillary_A | r_DecelTime | bilateral | 125100\20217-6\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | PG | meas | m_Axillary_A | r_PG | bilateral | 125100\20247-3\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | MG | meas | m_Axillary_A | r_PGmean | bilateral | 125100\20256-4\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | TAMV | meas | m_Axillary_A | r_TAMV | bilateral | 125100\20352-1\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | VTI | meas | m_Axillary_A | r_VTI | bilateral | 125100\20354-7\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | HR | meas | m_Axillary_A | r_HeartRate | bilateral | 125100\8867-4\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Ax A | Dist | meas | m_Axillary_A | r_Distance | bilateral | 125100\G-0364\T-47020\T-47100 | SRT, T-47020, Artery Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47100, Axillary Artery
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | ? | meas | m_Prox_Brachial_A | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | EDV | meas | m_Prox_Brachial_A | r_EDV | bilateral | 125100\11653-3\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | MDV | meas | m_Prox_Brachial_A | r_MDV | bilateral | 125100\11665-7\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | TAPV | meas | m_Prox_Brachial_A | r_TAPV | bilateral | 125100\11692-1\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | PSV | meas | m_Prox_Brachial_A | r_PSV | bilateral | 125100\11726-7\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | PI | meas | m_Prox_Brachial_A | r_PI | bilateral | 125100\12008-9\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | RI | meas | m_Prox_Brachial_A | r_RI | bilateral | 125100\12023-8\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | S/D | meas | m_Prox_Brachial_A | r_SD | bilateral | 125100\12144-2\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | AI | meas | m_Prox_Brachial_A | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | AT | meas | m_Prox_Brachial_A | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | DT | meas | m_Prox_Brachial_A | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | PG | meas | m_Prox_Brachial_A | r_PG | bilateral | 125100\20247-3\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | MG | meas | m_Prox_Brachial_A | r_PGmean | bilateral | 125100\20256-4\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | TAMV | meas | m_Prox_Brachial_A | r_TAMV | bilateral | 125100\20352-1\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | VTI | meas | m_Prox_Brachial_A | r_VTI | bilateral | 125100\20354-7\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | HR | meas | m_Prox_Brachial_A | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A | Dist | meas | m_Prox_Brachial_A | r_Distance | bilateral | 125100\G-0364\G-A118\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | ? | meas | m_Dist_Brachial_A | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | EDV | meas | m_Dist_Brachial_A | r_EDV | bilateral | 125100\11653-3\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | MDV | meas | m_Dist_Brachial_A | r_MDV | bilateral | 125100\11665-7\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | TAPV | meas | m_Dist_Brachial_A | r_TAPV | bilateral | 125100\11692-1\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | PSV | meas | m_Dist_Brachial_A | r_PSV | bilateral | 125100\11726-7\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | PI | meas | m_Dist_Brachial_A | r_PI | bilateral | 125100\12008-9\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | RI | meas | m_Dist_Brachial_A | r_RI | bilateral | 125100\12023-8\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | S/D | meas | m_Dist_Brachial_A | r_SD | bilateral | 125100\12144-2\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | AI | meas | m_Dist_Brachial_A | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | AT | meas | m_Dist_Brachial_A | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | DT | meas | m_Dist_Brachial_A | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | PG | meas | m_Dist_Brachial_A | r_PG | bilateral | 125100\20247-3\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | MG | meas | m_Dist_Brachial_A | r_PGmean | bilateral | 125100\20256-4\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | TAMV | meas | m_Dist_Brachial_A | r_TAMV | bilateral | 125100\20352-1\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | VTI | meas | m_Dist_Brachial_A | r_VTI | bilateral | 125100\20354-7\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | HR | meas | m_Dist_Brachial_A | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A | Dist | meas | m_Dist_Brachial_A | r_Distance | bilateral | 125100\G-0364\G-A119\T-47020\T-47160 | SRT, T-47020, Artery Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47160, Brachial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | ? | meas | m_Prox_Ulnar_A | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | EDV | meas | m_Prox_Ulnar_A | r_EDV | bilateral | 125100\11653-3\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | MDV | meas | m_Prox_Ulnar_A | r_MDV | bilateral | 125100\11665-7\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | TAPV | meas | m_Prox_Ulnar_A | r_TAPV | bilateral | 125100\11692-1\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | PSV | meas | m_Prox_Ulnar_A | r_PSV | bilateral | 125100\11726-7\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | PI | meas | m_Prox_Ulnar_A | r_PI | bilateral | 125100\12008-9\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | RI | meas | m_Prox_Ulnar_A | r_RI | bilateral | 125100\12023-8\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | S/D | meas | m_Prox_Ulnar_A | r_SD | bilateral | 125100\12144-2\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | AI | meas | m_Prox_Ulnar_A | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | AT | meas | m_Prox_Ulnar_A | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | DT | meas | m_Prox_Ulnar_A | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | PG | meas | m_Prox_Ulnar_A | r_PG | bilateral | 125100\20247-3\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | MG | meas | m_Prox_Ulnar_A | r_PGmean | bilateral | 125100\20256-4\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | TAMV | meas | m_Prox_Ulnar_A | r_TAMV | bilateral | 125100\20352-1\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | VTI | meas | m_Prox_Ulnar_A | r_VTI | bilateral | 125100\20354-7\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | HR | meas | m_Prox_Ulnar_A | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A | Dist | meas | m_Prox_Ulnar_A | r_Distance | bilateral | 125100\G-0364\G-A118\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | ? | meas | m_Dist_Ulnar_A | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | EDV | meas | m_Dist_Ulnar_A | r_EDV | bilateral | 125100\11653-3\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | MDV | meas | m_Dist_Ulnar_A | r_MDV | bilateral | 125100\11665-7\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | TAPV | meas | m_Dist_Ulnar_A | r_TAPV | bilateral | 125100\11692-1\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | PSV | meas | m_Dist_Ulnar_A | r_PSV | bilateral | 125100\11726-7\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | PI | meas | m_Dist_Ulnar_A | r_PI | bilateral | 125100\12008-9\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | RI | meas | m_Dist_Ulnar_A | r_RI | bilateral | 125100\12023-8\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | S/D | meas | m_Dist_Ulnar_A | r_SD | bilateral | 125100\12144-2\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | AI | meas | m_Dist_Ulnar_A | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | AT | meas | m_Dist_Ulnar_A | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | DT | meas | m_Dist_Ulnar_A | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | PG | meas | m_Dist_Ulnar_A | r_PG | bilateral | 125100\20247-3\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | MG | meas | m_Dist_Ulnar_A | r_PGmean | bilateral | 125100\20256-4\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | TAMV | meas | m_Dist_Ulnar_A | r_TAMV | bilateral | 125100\20352-1\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | VTI | meas | m_Dist_Ulnar_A | r_VTI | bilateral | 125100\20354-7\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | HR | meas | m_Dist_Ulnar_A | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A | Dist | meas | m_Dist_Ulnar_A | r_Distance | bilateral | 125100\G-0364\G-A119\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | ? | meas | m_Mid_Ulnar_A | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | EDV | meas | m_Mid_Ulnar_A | r_EDV | bilateral | 125100\11653-3\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | MDV | meas | m_Mid_Ulnar_A | r_MDV | bilateral | 125100\11665-7\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | TAPV | meas | m_Mid_Ulnar_A | r_TAPV | bilateral | 125100\11692-1\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | PSV | meas | m_Mid_Ulnar_A | r_PSV | bilateral | 125100\11726-7\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | PI | meas | m_Mid_Ulnar_A | r_PI | bilateral | 125100\12008-9\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | RI | meas | m_Mid_Ulnar_A | r_RI | bilateral | 125100\12023-8\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | S/D | meas | m_Mid_Ulnar_A | r_SD | bilateral | 125100\12144-2\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | AI | meas | m_Mid_Ulnar_A | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | AT | meas | m_Mid_Ulnar_A | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | DT | meas | m_Mid_Ulnar_A | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | PG | meas | m_Mid_Ulnar_A | r_PG | bilateral | 125100\20247-3\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | MG | meas | m_Mid_Ulnar_A | r_PGmean | bilateral | 125100\20256-4\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | TAMV | meas | m_Mid_Ulnar_A | r_TAMV | bilateral | 125100\20352-1\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | VTI | meas | m_Mid_Ulnar_A | r_VTI | bilateral | 125100\20354-7\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | HR | meas | m_Mid_Ulnar_A | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A | Dist | meas | m_Mid_Ulnar_A | r_Distance | bilateral | 125100\G-0364\G-A188\T-47020\T-47200 | SRT, T-47020, Artery Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47200, Ulnar Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | ? | meas | m_Prox_Radial_A | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | EDV | meas | m_Prox_Radial_A | r_EDV | bilateral | 125100\11653-3\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | MDV | meas | m_Prox_Radial_A | r_MDV | bilateral | 125100\11665-7\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | TAPV | meas | m_Prox_Radial_A | r_TAPV | bilateral | 125100\11692-1\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | PSV | meas | m_Prox_Radial_A | r_PSV | bilateral | 125100\11726-7\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | PI | meas | m_Prox_Radial_A | r_PI | bilateral | 125100\12008-9\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | RI | meas | m_Prox_Radial_A | r_RI | bilateral | 125100\12023-8\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | S/D | meas | m_Prox_Radial_A | r_SD | bilateral | 125100\12144-2\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | AI | meas | m_Prox_Radial_A | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | AT | meas | m_Prox_Radial_A | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | DT | meas | m_Prox_Radial_A | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | PG | meas | m_Prox_Radial_A | r_PG | bilateral | 125100\20247-3\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | MG | meas | m_Prox_Radial_A | r_PGmean | bilateral | 125100\20256-4\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | TAMV | meas | m_Prox_Radial_A | r_TAMV | bilateral | 125100\20352-1\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | VTI | meas | m_Prox_Radial_A | r_VTI | bilateral | 125100\20354-7\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | HR | meas | m_Prox_Radial_A | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A | Dist | meas | m_Prox_Radial_A | r_Distance | bilateral | 125100\G-0364\G-A118\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | ? | meas | m_Dist_Radial_A | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | EDV | meas | m_Dist_Radial_A | r_EDV | bilateral | 125100\11653-3\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | MDV | meas | m_Dist_Radial_A | r_MDV | bilateral | 125100\11665-7\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | TAPV | meas | m_Dist_Radial_A | r_TAPV | bilateral | 125100\11692-1\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | PSV | meas | m_Dist_Radial_A | r_PSV | bilateral | 125100\11726-7\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | PI | meas | m_Dist_Radial_A | r_PI | bilateral | 125100\12008-9\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | RI | meas | m_Dist_Radial_A | r_RI | bilateral | 125100\12023-8\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | S/D | meas | m_Dist_Radial_A | r_SD | bilateral | 125100\12144-2\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | AI | meas | m_Dist_Radial_A | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | AT | meas | m_Dist_Radial_A | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | DT | meas | m_Dist_Radial_A | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | PG | meas | m_Dist_Radial_A | r_PG | bilateral | 125100\20247-3\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | MG | meas | m_Dist_Radial_A | r_PGmean | bilateral | 125100\20256-4\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | TAMV | meas | m_Dist_Radial_A | r_TAMV | bilateral | 125100\20352-1\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | VTI | meas | m_Dist_Radial_A | r_VTI | bilateral | 125100\20354-7\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | HR | meas | m_Dist_Radial_A | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A | Dist | meas | m_Dist_Radial_A | r_Distance | bilateral | 125100\G-0364\G-A119\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | ? | meas | m_Mid_Radial_A | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | EDV | meas | m_Mid_Radial_A | r_EDV | bilateral | 125100\11653-3\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | MDV | meas | m_Mid_Radial_A | r_MDV | bilateral | 125100\11665-7\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | TAPV | meas | m_Mid_Radial_A | r_TAPV | bilateral | 125100\11692-1\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | PSV | meas | m_Mid_Radial_A | r_PSV | bilateral | 125100\11726-7\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | PI | meas | m_Mid_Radial_A | r_PI | bilateral | 125100\12008-9\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | RI | meas | m_Mid_Radial_A | r_RI | bilateral | 125100\12023-8\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | S/D | meas | m_Mid_Radial_A | r_SD | bilateral | 125100\12144-2\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | AI | meas | m_Mid_Radial_A | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | AT | meas | m_Mid_Radial_A | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | DT | meas | m_Mid_Radial_A | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | PG | meas | m_Mid_Radial_A | r_PG | bilateral | 125100\20247-3\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | MG | meas | m_Mid_Radial_A | r_PGmean | bilateral | 125100\20256-4\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | TAMV | meas | m_Mid_Radial_A | r_TAMV | bilateral | 125100\20352-1\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | VTI | meas | m_Mid_Radial_A | r_VTI | bilateral | 125100\20354-7\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | HR | meas | m_Mid_Radial_A | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A | Dist | meas | m_Mid_Radial_A | r_Distance | bilateral | 125100\G-0364\G-A188\T-47020\T-47300 | SRT, T-47020, Artery Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47300, Radial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Iliac Artery | ? | meas | m_iliac_artery | r_DopAngle | bilateral | 125100\C12122-01\T-46700\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46700, Iliac artery
Vascular Ultrasound Procedure Report | Abdominal | Iliac Artery | EDV | meas | m_iliac_artery | r_EDV | bilateral | 125100\11653-3\T-46700\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46700, Iliac artery
Vascular Ultrasound Procedure Report | Abdominal | Iliac Artery | MDV | meas | m_iliac_artery | r_MDV | bilateral | 125100\11665-7\T-46700\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46700, Iliac artery
Vascular Ultrasound Procedure Report | Abdominal | Iliac Artery | TAPV | meas | m_iliac_artery | r_TAPV | bilateral | 125100\11692-1\T-46700\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46700, Iliac artery
Vascular Ultrasound Procedure Report | Abdominal | Iliac Artery | PSV | meas | m_iliac_artery | r_PSV | bilateral | 125100\11726-7\T-46700\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46700, Iliac artery
Vascular Ultrasound Procedure Report | Abdominal | Iliac Artery | PI | meas | m_iliac_artery | r_PI | bilateral | 125100\12008-9\T-46700\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46700, Iliac artery
Vascular Ultrasound Procedure Report | Abdominal | Iliac Artery | RI | meas | m_iliac_artery | r_RI | bilateral | 125100\12023-8\T-46700\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46700, Iliac artery
Vascular Ultrasound Procedure Report | Abdominal | Iliac Artery | S/D | meas | m_iliac_artery | r_SD | bilateral | 125100\12144-2\T-46700\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46700, Iliac artery
Vascular Ultrasound Procedure Report | Abdominal | Iliac Artery | AI | meas | m_iliac_artery | r_AccelIndex | bilateral | 125100\20167-3\T-46700\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46700, Iliac artery
Vascular Ultrasound Procedure Report | Abdominal | Iliac Artery | AT | meas | m_iliac_artery | r_AccelTime | bilateral | 125100\20168-1\T-46700\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46700, Iliac artery
Vascular Ultrasound Procedure Report | Abdominal | Iliac Artery | DT | meas | m_iliac_artery | r_DecelTime | bilateral | 125100\20217-6\T-46700\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46700, Iliac artery
Vascular Ultrasound Procedure Report | Abdominal | Iliac Artery | PG | meas | m_iliac_artery | r_PG | bilateral | 125100\20247-3\T-46700\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46700, Iliac artery
Vascular Ultrasound Procedure Report | Abdominal | Iliac Artery | MG | meas | m_iliac_artery | r_PGmean | bilateral | 125100\20256-4\T-46700\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46700, Iliac artery
Vascular Ultrasound Procedure Report | Abdominal | Iliac Artery | TAMV | meas | m_iliac_artery | r_TAMV | bilateral | 125100\20352-1\T-46700\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46700, Iliac artery
Vascular Ultrasound Procedure Report | Abdominal | Iliac Artery | VTI | meas | m_iliac_artery | r_VTI | bilateral | 125100\20354-7\T-46700\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46700, Iliac artery
Vascular Ultrasound Procedure Report | Abdominal | Iliac Artery | HR | meas | m_iliac_artery | r_HeartRate | bilateral | 125100\8867-4\T-46700\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46700, Iliac artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | ? | meas | m_Com_Iliac_A | r_DopAngle | bilateral | 125100\C12122-01\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | EDV | meas | m_Com_Iliac_A | r_EDV | bilateral | 125100\11653-3\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | MDV | meas | m_Com_Iliac_A | r_MDV | bilateral | 125100\11665-7\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | TAPV | meas | m_Com_Iliac_A | r_TAPV | bilateral | 125100\11692-1\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | PSV | meas | m_Com_Iliac_A | r_PSV | bilateral | 125100\11726-7\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | PI | meas | m_Com_Iliac_A | r_PI | bilateral | 125100\12008-9\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | RI | meas | m_Com_Iliac_A | r_RI | bilateral | 125100\12023-8\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | S/D | meas | m_Com_Iliac_A | r_SD | bilateral | 125100\12144-2\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | AI | meas | m_Com_Iliac_A | r_AccelIndex | bilateral | 125100\20167-3\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | AT | meas | m_Com_Iliac_A | r_AccelTime | bilateral | 125100\20168-1\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | DT | meas | m_Com_Iliac_A | r_DecelTime | bilateral | 125100\20217-6\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | PG | meas | m_Com_Iliac_A | r_PG | bilateral | 125100\20247-3\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | MG | meas | m_Com_Iliac_A | r_PGmean | bilateral | 125100\20256-4\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | TAMV | meas | m_Com_Iliac_A | r_TAMV | bilateral | 125100\20352-1\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | VTI | meas | m_Com_Iliac_A | r_VTI | bilateral | 125100\20354-7\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | HR | meas | m_Com_Iliac_A | r_HeartRate | bilateral | 125100\8867-4\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A | Dist | meas | m_Com_Iliac_A | r_Distance | bilateral | 125100\G-0364\T-46710\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46710, Common Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | ? | meas | m_Int_Iliac_A | r_DopAngle | bilateral | 125100\C12122-01\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | EDV | meas | m_Int_Iliac_A | r_EDV | bilateral | 125100\11653-3\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | MDV | meas | m_Int_Iliac_A | r_MDV | bilateral | 125100\11665-7\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | TAPV | meas | m_Int_Iliac_A | r_TAPV | bilateral | 125100\11692-1\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | PSV | meas | m_Int_Iliac_A | r_PSV | bilateral | 125100\11726-7\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | PI | meas | m_Int_Iliac_A | r_PI | bilateral | 125100\12008-9\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | RI | meas | m_Int_Iliac_A | r_RI | bilateral | 125100\12023-8\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | S/D | meas | m_Int_Iliac_A | r_SD | bilateral | 125100\12144-2\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | AI | meas | m_Int_Iliac_A | r_AccelIndex | bilateral | 125100\20167-3\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | AT | meas | m_Int_Iliac_A | r_AccelTime | bilateral | 125100\20168-1\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | DT | meas | m_Int_Iliac_A | r_DecelTime | bilateral | 125100\20217-6\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | PG | meas | m_Int_Iliac_A | r_PG | bilateral | 125100\20247-3\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | MG | meas | m_Int_Iliac_A | r_PGmean | bilateral | 125100\20256-4\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | TAMV | meas | m_Int_Iliac_A | r_TAMV | bilateral | 125100\20352-1\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | VTI | meas | m_Int_Iliac_A | r_VTI | bilateral | 125100\20354-7\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | HR | meas | m_Int_Iliac_A | r_HeartRate | bilateral | 125100\8867-4\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A | Dist | meas | m_Int_Iliac_A | r_Distance | bilateral | 125100\G-0364\T-46740\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46740, Internal Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | ? | meas | m_Ext_Iliac_A | r_DopAngle | bilateral | 125100\C12122-01\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | EDV | meas | m_Ext_Iliac_A | r_EDV | bilateral | 125100\11653-3\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | MDV | meas | m_Ext_Iliac_A | r_MDV | bilateral | 125100\11665-7\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | TAPV | meas | m_Ext_Iliac_A | r_TAPV | bilateral | 125100\11692-1\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | PSV | meas | m_Ext_Iliac_A | r_PSV | bilateral | 125100\11726-7\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | PI | meas | m_Ext_Iliac_A | r_PI | bilateral | 125100\12008-9\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | RI | meas | m_Ext_Iliac_A | r_RI | bilateral | 125100\12023-8\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | S/D | meas | m_Ext_Iliac_A | r_SD | bilateral | 125100\12144-2\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | AI | meas | m_Ext_Iliac_A | r_AccelIndex | bilateral | 125100\20167-3\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | AT | meas | m_Ext_Iliac_A | r_AccelTime | bilateral | 125100\20168-1\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | DT | meas | m_Ext_Iliac_A | r_DecelTime | bilateral | 125100\20217-6\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | PG | meas | m_Ext_Iliac_A | r_PG | bilateral | 125100\20247-3\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | MG | meas | m_Ext_Iliac_A | r_PGmean | bilateral | 125100\20256-4\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | TAMV | meas | m_Ext_Iliac_A | r_TAMV | bilateral | 125100\20352-1\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | VTI | meas | m_Ext_Iliac_A | r_VTI | bilateral | 125100\20354-7\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | HR | meas | m_Ext_Iliac_A | r_HeartRate | bilateral | 125100\8867-4\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A | Dist | meas | m_Ext_Iliac_A | r_Distance | bilateral | 125100\G-0364\T-46910\T-47040 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46910, External Iliac Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | ? | meas | m_CFA | r_DopAngle | bilateral | 125100\C12122-01\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | EDV | meas | m_CFA | r_EDV | bilateral | 125100\11653-3\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | MDV | meas | m_CFA | r_MDV | bilateral | 125100\11665-7\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | TAPV | meas | m_CFA | r_TAPV | bilateral | 125100\11692-1\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | PSV | meas | m_CFA | r_PSV | bilateral | 125100\11726-7\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | PI | meas | m_CFA | r_PI | bilateral | 125100\12008-9\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | RI | meas | m_CFA | r_RI | bilateral | 125100\12023-8\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | S/D | meas | m_CFA | r_SD | bilateral | 125100\12144-2\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | AI | meas | m_CFA | r_AccelIndex | bilateral | 125100\20167-3\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | AT | meas | m_CFA | r_AccelTime | bilateral | 125100\20168-1\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | DT | meas | m_CFA | r_DecelTime | bilateral | 125100\20217-6\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | PG | meas | m_CFA | r_PG | bilateral | 125100\20247-3\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | MG | meas | m_CFA | r_PGmean | bilateral | 125100\20256-4\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | TAMV | meas | m_CFA | r_TAMV | bilateral | 125100\20352-1\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | VTI | meas | m_CFA | r_VTI | bilateral | 125100\20354-7\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | HR | meas | m_CFA | r_HeartRate | bilateral | 125100\8867-4\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | CFA | Dist | meas | m_CFA | r_Distance | bilateral | 125100\G-0364\T-47040\T-47400 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47400, Common Femoral Artery
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | ? | meas | m_Prox_SFA | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | EDV | meas | m_Prox_SFA | r_EDV | bilateral | 125100\11653-3\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | MDV | meas | m_Prox_SFA | r_MDV | bilateral | 125100\11665-7\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | TAPV | meas | m_Prox_SFA | r_TAPV | bilateral | 125100\11692-1\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | PSV | meas | m_Prox_SFA | r_PSV | bilateral | 125100\11726-7\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | PI | meas | m_Prox_SFA | r_PI | bilateral | 125100\12008-9\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | RI | meas | m_Prox_SFA | r_RI | bilateral | 125100\12023-8\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | S/D | meas | m_Prox_SFA | r_SD | bilateral | 125100\12144-2\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | AI | meas | m_Prox_SFA | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | AT | meas | m_Prox_SFA | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | DT | meas | m_Prox_SFA | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | PG | meas | m_Prox_SFA | r_PG | bilateral | 125100\20247-3\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | MG | meas | m_Prox_SFA | r_PGmean | bilateral | 125100\20256-4\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | TAMV | meas | m_Prox_SFA | r_TAMV | bilateral | 125100\20352-1\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | VTI | meas | m_Prox_SFA | r_VTI | bilateral | 125100\20354-7\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | HR | meas | m_Prox_SFA | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA | Dist | meas | m_Prox_SFA | r_Distance | bilateral | 125100\G-0364\G-A118\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | ? | meas | m_Dist_SFA | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | EDV | meas | m_Dist_SFA | r_EDV | bilateral | 125100\11653-3\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | MDV | meas | m_Dist_SFA | r_MDV | bilateral | 125100\11665-7\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | TAPV | meas | m_Dist_SFA | r_TAPV | bilateral | 125100\11692-1\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | PSV | meas | m_Dist_SFA | r_PSV | bilateral | 125100\11726-7\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | PI | meas | m_Dist_SFA | r_PI | bilateral | 125100\12008-9\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | RI | meas | m_Dist_SFA | r_RI | bilateral | 125100\12023-8\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | S/D | meas | m_Dist_SFA | r_SD | bilateral | 125100\12144-2\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | AI | meas | m_Dist_SFA | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | AT | meas | m_Dist_SFA | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | DT | meas | m_Dist_SFA | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | PG | meas | m_Dist_SFA | r_PG | bilateral | 125100\20247-3\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | MG | meas | m_Dist_SFA | r_PGmean | bilateral | 125100\20256-4\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | TAMV | meas | m_Dist_SFA | r_TAMV | bilateral | 125100\20352-1\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | VTI | meas | m_Dist_SFA | r_VTI | bilateral | 125100\20354-7\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | HR | meas | m_Dist_SFA | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA | Dist | meas | m_Dist_SFA | r_Distance | bilateral | 125100\G-0364\G-A119\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | ? | meas | m_Mid_SFA | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | EDV | meas | m_Mid_SFA | r_EDV | bilateral | 125100\11653-3\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | MDV | meas | m_Mid_SFA | r_MDV | bilateral | 125100\11665-7\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | TAPV | meas | m_Mid_SFA | r_TAPV | bilateral | 125100\11692-1\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | PSV | meas | m_Mid_SFA | r_PSV | bilateral | 125100\11726-7\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | PI | meas | m_Mid_SFA | r_PI | bilateral | 125100\12008-9\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | RI | meas | m_Mid_SFA | r_RI | bilateral | 125100\12023-8\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | S/D | meas | m_Mid_SFA | r_SD | bilateral | 125100\12144-2\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | AI | meas | m_Mid_SFA | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | AT | meas | m_Mid_SFA | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | DT | meas | m_Mid_SFA | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | PG | meas | m_Mid_SFA | r_PG | bilateral | 125100\20247-3\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | MG | meas | m_Mid_SFA | r_PGmean | bilateral | 125100\20256-4\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | TAMV | meas | m_Mid_SFA | r_TAMV | bilateral | 125100\20352-1\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | VTI | meas | m_Mid_SFA | r_VTI | bilateral | 125100\20354-7\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | HR | meas | m_Mid_SFA | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA | Dist | meas | m_Mid_SFA | r_Distance | bilateral | 125100\G-0364\G-A188\T-47040\T-47403 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47403, Superficial Femoral Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | ? | meas | m_PFA | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | EDV | meas | m_PFA | r_EDV | bilateral | 125100\11653-3\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | MDV | meas | m_PFA | r_MDV | bilateral | 125100\11665-7\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | TAPV | meas | m_PFA | r_TAPV | bilateral | 125100\11692-1\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | PSV | meas | m_PFA | r_PSV | bilateral | 125100\11726-7\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | PI | meas | m_PFA | r_PI | bilateral | 125100\12008-9\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | RI | meas | m_PFA | r_RI | bilateral | 125100\12023-8\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | S/D | meas | m_PFA | r_SD | bilateral | 125100\12144-2\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | AI | meas | m_PFA | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | AT | meas | m_PFA | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | DT | meas | m_PFA | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | PG | meas | m_PFA | r_PG | bilateral | 125100\20247-3\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | MG | meas | m_PFA | r_PGmean | bilateral | 125100\20256-4\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | TAMV | meas | m_PFA | r_TAMV | bilateral | 125100\20352-1\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | VTI | meas | m_PFA | r_VTI | bilateral | 125100\20354-7\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | HR | meas | m_PFA | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA | Dist | meas | m_PFA | r_Distance | bilateral | 125100\G-0364\G-A118\T-47040\T-47440 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47440, Profunda Femoris Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | ? | meas | m_Prox_POPA | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | EDV | meas | m_Prox_POPA | r_EDV | bilateral | 125100\11653-3\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | MDV | meas | m_Prox_POPA | r_MDV | bilateral | 125100\11665-7\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | TAPV | meas | m_Prox_POPA | r_TAPV | bilateral | 125100\11692-1\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | PSV | meas | m_Prox_POPA | r_PSV | bilateral | 125100\11726-7\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | PI | meas | m_Prox_POPA | r_PI | bilateral | 125100\12008-9\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | RI | meas | m_Prox_POPA | r_RI | bilateral | 125100\12023-8\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | S/D | meas | m_Prox_POPA | r_SD | bilateral | 125100\12144-2\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | AI | meas | m_Prox_POPA | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | AT | meas | m_Prox_POPA | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | DT | meas | m_Prox_POPA | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | PG | meas | m_Prox_POPA | r_PG | bilateral | 125100\20247-3\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | MG | meas | m_Prox_POPA | r_PGmean | bilateral | 125100\20256-4\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | TAMV | meas | m_Prox_POPA | r_TAMV | bilateral | 125100\20352-1\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | VTI | meas | m_Prox_POPA | r_VTI | bilateral | 125100\20354-7\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | HR | meas | m_Prox_POPA | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A | Dist | meas | m_Prox_POPA | r_Distance | bilateral | 125100\G-0364\G-A118\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | ? | meas | m_Dist_POPA | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | EDV | meas | m_Dist_POPA | r_EDV | bilateral | 125100\11653-3\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | MDV | meas | m_Dist_POPA | r_MDV | bilateral | 125100\11665-7\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | TAPV | meas | m_Dist_POPA | r_TAPV | bilateral | 125100\11692-1\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | PSV | meas | m_Dist_POPA | r_PSV | bilateral | 125100\11726-7\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | PI | meas | m_Dist_POPA | r_PI | bilateral | 125100\12008-9\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | RI | meas | m_Dist_POPA | r_RI | bilateral | 125100\12023-8\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | S/D | meas | m_Dist_POPA | r_SD | bilateral | 125100\12144-2\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | AI | meas | m_Dist_POPA | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | AT | meas | m_Dist_POPA | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | DT | meas | m_Dist_POPA | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | PG | meas | m_Dist_POPA | r_PG | bilateral | 125100\20247-3\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | MG | meas | m_Dist_POPA | r_PGmean | bilateral | 125100\20256-4\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | TAMV | meas | m_Dist_POPA | r_TAMV | bilateral | 125100\20352-1\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | VTI | meas | m_Dist_POPA | r_VTI | bilateral | 125100\20354-7\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | HR | meas | m_Dist_POPA | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A | Dist | meas | m_Dist_POPA | r_Distance | bilateral | 125100\G-0364\G-A119\T-47040\T-47500 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47500, Popliteal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | ? | meas | m_Prox_PTA | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | EDV | meas | m_Prox_PTA | r_EDV | bilateral | 125100\11653-3\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | MDV | meas | m_Prox_PTA | r_MDV | bilateral | 125100\11665-7\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | TAPV | meas | m_Prox_PTA | r_TAPV | bilateral | 125100\11692-1\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | PSV | meas | m_Prox_PTA | r_PSV | bilateral | 125100\11726-7\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | PI | meas | m_Prox_PTA | r_PI | bilateral | 125100\12008-9\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | RI | meas | m_Prox_PTA | r_RI | bilateral | 125100\12023-8\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | S/D | meas | m_Prox_PTA | r_SD | bilateral | 125100\12144-2\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | AI | meas | m_Prox_PTA | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | AT | meas | m_Prox_PTA | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | DT | meas | m_Prox_PTA | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | PG | meas | m_Prox_PTA | r_PG | bilateral | 125100\20247-3\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | MG | meas | m_Prox_PTA | r_PGmean | bilateral | 125100\20256-4\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | TAMV | meas | m_Prox_PTA | r_TAMV | bilateral | 125100\20352-1\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | VTI | meas | m_Prox_PTA | r_VTI | bilateral | 125100\20354-7\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | HR | meas | m_Prox_PTA | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA | Dist | meas | m_Prox_PTA | r_Distance | bilateral | 125100\G-0364\G-A118\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | ? | meas | m_Dist_PTA | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | EDV | meas | m_Dist_PTA | r_EDV | bilateral | 125100\11653-3\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | MDV | meas | m_Dist_PTA | r_MDV | bilateral | 125100\11665-7\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | TAPV | meas | m_Dist_PTA | r_TAPV | bilateral | 125100\11692-1\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | PSV | meas | m_Dist_PTA | r_PSV | bilateral | 125100\11726-7\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | PI | meas | m_Dist_PTA | r_PI | bilateral | 125100\12008-9\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | RI | meas | m_Dist_PTA | r_RI | bilateral | 125100\12023-8\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | S/D | meas | m_Dist_PTA | r_SD | bilateral | 125100\12144-2\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | AI | meas | m_Dist_PTA | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | AT | meas | m_Dist_PTA | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | DT | meas | m_Dist_PTA | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | PG | meas | m_Dist_PTA | r_PG | bilateral | 125100\20247-3\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | MG | meas | m_Dist_PTA | r_PGmean | bilateral | 125100\20256-4\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | TAMV | meas | m_Dist_PTA | r_TAMV | bilateral | 125100\20352-1\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | VTI | meas | m_Dist_PTA | r_VTI | bilateral | 125100\20354-7\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | HR | meas | m_Dist_PTA | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA | Dist | meas | m_Dist_PTA | r_Distance | bilateral | 125100\G-0364\G-A119\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | ? | meas | m_Mid_PTA | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | EDV | meas | m_Mid_PTA | r_EDV | bilateral | 125100\11653-3\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | MDV | meas | m_Mid_PTA | r_MDV | bilateral | 125100\11665-7\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | TAPV | meas | m_Mid_PTA | r_TAPV | bilateral | 125100\11692-1\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | PSV | meas | m_Mid_PTA | r_PSV | bilateral | 125100\11726-7\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | PI | meas | m_Mid_PTA | r_PI | bilateral | 125100\12008-9\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | RI | meas | m_Mid_PTA | r_RI | bilateral | 125100\12023-8\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | S/D | meas | m_Mid_PTA | r_SD | bilateral | 125100\12144-2\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | AI | meas | m_Mid_PTA | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | AT | meas | m_Mid_PTA | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | DT | meas | m_Mid_PTA | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | PG | meas | m_Mid_PTA | r_PG | bilateral | 125100\20247-3\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | MG | meas | m_Mid_PTA | r_PGmean | bilateral | 125100\20256-4\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | TAMV | meas | m_Mid_PTA | r_TAMV | bilateral | 125100\20352-1\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | VTI | meas | m_Mid_PTA | r_VTI | bilateral | 125100\20354-7\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | HR | meas | m_Mid_PTA | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA | Dist | meas | m_Mid_PTA | r_Distance | bilateral | 125100\G-0364\G-A188\T-47040\T-47600 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47600, Posterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | ? | meas | m_Prox_Pero_A | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | EDV | meas | m_Prox_Pero_A | r_EDV | bilateral | 125100\11653-3\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | MDV | meas | m_Prox_Pero_A | r_MDV | bilateral | 125100\11665-7\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | TAPV | meas | m_Prox_Pero_A | r_TAPV | bilateral | 125100\11692-1\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | PSV | meas | m_Prox_Pero_A | r_PSV | bilateral | 125100\11726-7\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | PI | meas | m_Prox_Pero_A | r_PI | bilateral | 125100\12008-9\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | RI | meas | m_Prox_Pero_A | r_RI | bilateral | 125100\12023-8\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | S/D | meas | m_Prox_Pero_A | r_SD | bilateral | 125100\12144-2\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | AI | meas | m_Prox_Pero_A | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | AT | meas | m_Prox_Pero_A | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | DT | meas | m_Prox_Pero_A | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | PG | meas | m_Prox_Pero_A | r_PG | bilateral | 125100\20247-3\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | MG | meas | m_Prox_Pero_A | r_PGmean | bilateral | 125100\20256-4\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | TAMV | meas | m_Prox_Pero_A | r_TAMV | bilateral | 125100\20352-1\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | VTI | meas | m_Prox_Pero_A | r_VTI | bilateral | 125100\20354-7\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | HR | meas | m_Prox_Pero_A | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A | Dist | meas | m_Prox_Pero_A | r_Distance | bilateral | 125100\G-0364\G-A118\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | ? | meas | m_Dist_Pero_A | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | EDV | meas | m_Dist_Pero_A | r_EDV | bilateral | 125100\11653-3\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | MDV | meas | m_Dist_Pero_A | r_MDV | bilateral | 125100\11665-7\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | TAPV | meas | m_Dist_Pero_A | r_TAPV | bilateral | 125100\11692-1\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | PSV | meas | m_Dist_Pero_A | r_PSV | bilateral | 125100\11726-7\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | PI | meas | m_Dist_Pero_A | r_PI | bilateral | 125100\12008-9\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | RI | meas | m_Dist_Pero_A | r_RI | bilateral | 125100\12023-8\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | S/D | meas | m_Dist_Pero_A | r_SD | bilateral | 125100\12144-2\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | AI | meas | m_Dist_Pero_A | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | AT | meas | m_Dist_Pero_A | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | DT | meas | m_Dist_Pero_A | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | PG | meas | m_Dist_Pero_A | r_PG | bilateral | 125100\20247-3\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | MG | meas | m_Dist_Pero_A | r_PGmean | bilateral | 125100\20256-4\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | TAMV | meas | m_Dist_Pero_A | r_TAMV | bilateral | 125100\20352-1\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | VTI | meas | m_Dist_Pero_A | r_VTI | bilateral | 125100\20354-7\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | HR | meas | m_Dist_Pero_A | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A | Dist | meas | m_Dist_Pero_A | r_Distance | bilateral | 125100\G-0364\G-A119\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | ? | meas | m_Mid_Pero_A | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | EDV | meas | m_Mid_Pero_A | r_EDV | bilateral | 125100\11653-3\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | MDV | meas | m_Mid_Pero_A | r_MDV | bilateral | 125100\11665-7\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | TAPV | meas | m_Mid_Pero_A | r_TAPV | bilateral | 125100\11692-1\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | PSV | meas | m_Mid_Pero_A | r_PSV | bilateral | 125100\11726-7\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | PI | meas | m_Mid_Pero_A | r_PI | bilateral | 125100\12008-9\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | RI | meas | m_Mid_Pero_A | r_RI | bilateral | 125100\12023-8\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | S/D | meas | m_Mid_Pero_A | r_SD | bilateral | 125100\12144-2\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | AI | meas | m_Mid_Pero_A | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | AT | meas | m_Mid_Pero_A | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | DT | meas | m_Mid_Pero_A | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | PG | meas | m_Mid_Pero_A | r_PG | bilateral | 125100\20247-3\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | MG | meas | m_Mid_Pero_A | r_PGmean | bilateral | 125100\20256-4\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | TAMV | meas | m_Mid_Pero_A | r_TAMV | bilateral | 125100\20352-1\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | VTI | meas | m_Mid_Pero_A | r_VTI | bilateral | 125100\20354-7\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | HR | meas | m_Mid_Pero_A | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A | Dist | meas | m_Mid_Pero_A | r_Distance | bilateral | 125100\G-0364\G-A188\T-47040\T-47630 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47630, Peroneal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | ? | meas | m_Prox_ATA | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | EDV | meas | m_Prox_ATA | r_EDV | bilateral | 125100\11653-3\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | MDV | meas | m_Prox_ATA | r_MDV | bilateral | 125100\11665-7\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | TAPV | meas | m_Prox_ATA | r_TAPV | bilateral | 125100\11692-1\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | PSV | meas | m_Prox_ATA | r_PSV | bilateral | 125100\11726-7\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | PI | meas | m_Prox_ATA | r_PI | bilateral | 125100\12008-9\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | RI | meas | m_Prox_ATA | r_RI | bilateral | 125100\12023-8\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | S/D | meas | m_Prox_ATA | r_SD | bilateral | 125100\12144-2\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | AI | meas | m_Prox_ATA | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | AT | meas | m_Prox_ATA | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | DT | meas | m_Prox_ATA | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | PG | meas | m_Prox_ATA | r_PG | bilateral | 125100\20247-3\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | MG | meas | m_Prox_ATA | r_PGmean | bilateral | 125100\20256-4\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | TAMV | meas | m_Prox_ATA | r_TAMV | bilateral | 125100\20352-1\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | VTI | meas | m_Prox_ATA | r_VTI | bilateral | 125100\20354-7\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | HR | meas | m_Prox_ATA | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA | Dist | meas | m_Prox_ATA | r_Distance | bilateral | 125100\G-0364\G-A118\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | ? | meas | m_Dist_ATA | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | EDV | meas | m_Dist_ATA | r_EDV | bilateral | 125100\11653-3\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | MDV | meas | m_Dist_ATA | r_MDV | bilateral | 125100\11665-7\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | TAPV | meas | m_Dist_ATA | r_TAPV | bilateral | 125100\11692-1\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | PSV | meas | m_Dist_ATA | r_PSV | bilateral | 125100\11726-7\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | PI | meas | m_Dist_ATA | r_PI | bilateral | 125100\12008-9\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | RI | meas | m_Dist_ATA | r_RI | bilateral | 125100\12023-8\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | S/D | meas | m_Dist_ATA | r_SD | bilateral | 125100\12144-2\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | AI | meas | m_Dist_ATA | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | AT | meas | m_Dist_ATA | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | DT | meas | m_Dist_ATA | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | PG | meas | m_Dist_ATA | r_PG | bilateral | 125100\20247-3\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | MG | meas | m_Dist_ATA | r_PGmean | bilateral | 125100\20256-4\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | TAMV | meas | m_Dist_ATA | r_TAMV | bilateral | 125100\20352-1\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | VTI | meas | m_Dist_ATA | r_VTI | bilateral | 125100\20354-7\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | HR | meas | m_Dist_ATA | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA | Dist | meas | m_Dist_ATA | r_Distance | bilateral | 125100\G-0364\G-A119\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | ? | meas | m_Mid_ATA | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | EDV | meas | m_Mid_ATA | r_EDV | bilateral | 125100\11653-3\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | MDV | meas | m_Mid_ATA | r_MDV | bilateral | 125100\11665-7\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | TAPV | meas | m_Mid_ATA | r_TAPV | bilateral | 125100\11692-1\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | PSV | meas | m_Mid_ATA | r_PSV | bilateral | 125100\11726-7\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | PI | meas | m_Mid_ATA | r_PI | bilateral | 125100\12008-9\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | RI | meas | m_Mid_ATA | r_RI | bilateral | 125100\12023-8\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | S/D | meas | m_Mid_ATA | r_SD | bilateral | 125100\12144-2\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | AI | meas | m_Mid_ATA | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | AT | meas | m_Mid_ATA | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | DT | meas | m_Mid_ATA | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | PG | meas | m_Mid_ATA | r_PG | bilateral | 125100\20247-3\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | MG | meas | m_Mid_ATA | r_PGmean | bilateral | 125100\20256-4\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | TAMV | meas | m_Mid_ATA | r_TAMV | bilateral | 125100\20352-1\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | VTI | meas | m_Mid_ATA | r_VTI | bilateral | 125100\20354-7\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | HR | meas | m_Mid_ATA | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA | Dist | meas | m_Mid_ATA | r_Distance | bilateral | 125100\G-0364\G-A188\T-47040\T-47700 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47700, Anterior Tibial Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | ? | meas | m_Dor_Ped_ATA | r_DopAngle | bilateral | 125100\C12122-01\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | EDV | meas | m_Dor_Ped_ATA | r_EDV | bilateral | 125100\11653-3\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | MDV | meas | m_Dor_Ped_ATA | r_MDV | bilateral | 125100\11665-7\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | TAPV | meas | m_Dor_Ped_ATA | r_TAPV | bilateral | 125100\11692-1\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | PSV | meas | m_Dor_Ped_ATA | r_PSV | bilateral | 125100\11726-7\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | PI | meas | m_Dor_Ped_ATA | r_PI | bilateral | 125100\12008-9\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | RI | meas | m_Dor_Ped_ATA | r_RI | bilateral | 125100\12023-8\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | S/D | meas | m_Dor_Ped_ATA | r_SD | bilateral | 125100\12144-2\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | AI | meas | m_Dor_Ped_ATA | r_AccelIndex | bilateral | 125100\20167-3\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | AT | meas | m_Dor_Ped_ATA | r_AccelTime | bilateral | 125100\20168-1\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | DT | meas | m_Dor_Ped_ATA | r_DecelTime | bilateral | 125100\20217-6\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | PG | meas | m_Dor_Ped_ATA | r_PG | bilateral | 125100\20247-3\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | MG | meas | m_Dor_Ped_ATA | r_PGmean | bilateral | 125100\20256-4\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | TAMV | meas | m_Dor_Ped_ATA | r_TAMV | bilateral | 125100\20352-1\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | VTI | meas | m_Dor_Ped_ATA | r_VTI | bilateral | 125100\20354-7\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | HR | meas | m_Dor_Ped_ATA | r_HeartRate | bilateral | 125100\8867-4\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis | Dist | meas | m_Dor_Ped_ATA | r_Distance | bilateral | 125100\G-0364\T-47040\T-47741 | SRT, T-47040, Artery of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-47741, Dorsalis Pedis Artery
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | ? | meas | m_proxivc | r_DopAngle | unilateral | 125100\C12122-01\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | EDV | meas | m_proxivc | r_EDV | unilateral | 125100\11653-3\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | MDV | meas | m_proxivc | r_MDV | unilateral | 125100\11665-7\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | TAPV | meas | m_proxivc | r_TAPV | unilateral | 125100\11692-1\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | PSV | meas | m_proxivc | r_PSV | unilateral | 125100\11726-7\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | PI | meas | m_proxivc | r_PI | unilateral | 125100\12008-9\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | RI | meas | m_proxivc | r_RI | unilateral | 125100\12023-8\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | S/D | meas | m_proxivc | r_SD | unilateral | 125100\12144-2\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | AI | meas | m_proxivc | r_AccelIndex | unilateral | 125100\20167-3\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | AT | meas | m_proxivc | r_AccelTime | unilateral | 125100\20168-1\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | DT | meas | m_proxivc | r_DecelTime | unilateral | 125100\20217-6\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | PG | meas | m_proxivc | r_PG | unilateral | 125100\20247-3\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | MG | meas | m_proxivc | r_PGmean | unilateral | 125100\20256-4\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | TAMV | meas | m_proxivc | r_TAMV | unilateral | 125100\20352-1\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | VTI | meas | m_proxivc | r_VTI | unilateral | 125100\20354-7\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | HR | meas | m_proxivc | r_HeartRate | unilateral | 125100\8867-4\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox IVC | Dist | meas | m_proxivc | r_Distance | unilateral | 125100\G-0364\G-A118\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | ? | meas | m_distivc | r_DopAngle | unilateral | 125100\C12122-01\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | EDV | meas | m_distivc | r_EDV | unilateral | 125100\11653-3\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | MDV | meas | m_distivc | r_MDV | unilateral | 125100\11665-7\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | TAPV | meas | m_distivc | r_TAPV | unilateral | 125100\11692-1\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | PSV | meas | m_distivc | r_PSV | unilateral | 125100\11726-7\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | PI | meas | m_distivc | r_PI | unilateral | 125100\12008-9\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | RI | meas | m_distivc | r_RI | unilateral | 125100\12023-8\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | S/D | meas | m_distivc | r_SD | unilateral | 125100\12144-2\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | AI | meas | m_distivc | r_AccelIndex | unilateral | 125100\20167-3\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | AT | meas | m_distivc | r_AccelTime | unilateral | 125100\20168-1\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | DT | meas | m_distivc | r_DecelTime | unilateral | 125100\20217-6\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | PG | meas | m_distivc | r_PG | unilateral | 125100\20247-3\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | MG | meas | m_distivc | r_PGmean | unilateral | 125100\20256-4\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | TAMV | meas | m_distivc | r_TAMV | unilateral | 125100\20352-1\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | VTI | meas | m_distivc | r_VTI | unilateral | 125100\20354-7\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | HR | meas | m_distivc | r_HeartRate | unilateral | 125100\8867-4\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist IVC | Dist | meas | m_distivc | r_Distance | unilateral | 125100\G-0364\G-A119\T-48710\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48710, Inferior Vena Cava |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | ? | meas | m_hepVeinRight | r_DopAngle | unilateral | 125100\C12122-01\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | EDV | meas | m_hepVeinRight | r_EDV | unilateral | 125100\11653-3\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | MDV | meas | m_hepVeinRight | r_MDV | unilateral | 125100\11665-7\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | TAPV | meas | m_hepVeinRight | r_TAPV | unilateral | 125100\11692-1\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | PSV | meas | m_hepVeinRight | r_PSV | unilateral | 125100\11726-7\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | PI | meas | m_hepVeinRight | r_PI | unilateral | 125100\12008-9\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | RI | meas | m_hepVeinRight | r_RI | unilateral | 125100\12023-8\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | S/D | meas | m_hepVeinRight | r_SD | unilateral | 125100\12144-2\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | AI | meas | m_hepVeinRight | r_AccelIndex | unilateral | 125100\20167-3\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | AT | meas | m_hepVeinRight | r_AccelTime | unilateral | 125100\20168-1\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | DT | meas | m_hepVeinRight | r_DecelTime | unilateral | 125100\20217-6\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | PG | meas | m_hepVeinRight | r_PG | unilateral | 125100\20247-3\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | MG | meas | m_hepVeinRight | r_PGmean | unilateral | 125100\20256-4\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | TAMV | meas | m_hepVeinRight | r_TAMV | unilateral | 125100\20352-1\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | VTI | meas | m_hepVeinRight | r_VTI | unilateral | 125100\20354-7\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | HR | meas | m_hepVeinRight | r_HeartRate | unilateral | 125100\8867-4\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | R Hepatic V | Dist | meas | m_hepVeinRight | r_Distance | unilateral | 125100\G-0364\T-48725\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48725, Right Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | ? | meas | m_midhepvein | r_DopAngle | unilateral | 125100\C12122-01\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | EDV | meas | m_midhepvein | r_EDV | unilateral | 125100\11653-3\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | MDV | meas | m_midhepvein | r_MDV | unilateral | 125100\11665-7\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | TAPV | meas | m_midhepvein | r_TAPV | unilateral | 125100\11692-1\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | PSV | meas | m_midhepvein | r_PSV | unilateral | 125100\11726-7\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | PI | meas | m_midhepvein | r_PI | unilateral | 125100\12008-9\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | RI | meas | m_midhepvein | r_RI | unilateral | 125100\12023-8\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | S/D | meas | m_midhepvein | r_SD | unilateral | 125100\12144-2\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | AI | meas | m_midhepvein | r_AccelIndex | unilateral | 125100\20167-3\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | AT | meas | m_midhepvein | r_AccelTime | unilateral | 125100\20168-1\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | DT | meas | m_midhepvein | r_DecelTime | unilateral | 125100\20217-6\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | PG | meas | m_midhepvein | r_PG | unilateral | 125100\20247-3\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | MG | meas | m_midhepvein | r_PGmean | unilateral | 125100\20256-4\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | TAMV | meas | m_midhepvein | r_TAMV | unilateral | 125100\20352-1\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | VTI | meas | m_midhepvein | r_VTI | unilateral | 125100\20354-7\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | HR | meas | m_midhepvein | r_HeartRate | unilateral | 125100\8867-4\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Hepatic V | Dist | meas | m_midhepvein | r_Distance | unilateral | 125100\G-0364\T-48726\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48726, Middle Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | ? | meas | m_hepVeinLeft | r_DopAngle | unilateral | 125100\C12122-01\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | EDV | meas | m_hepVeinLeft | r_EDV | unilateral | 125100\11653-3\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | MDV | meas | m_hepVeinLeft | r_MDV | unilateral | 125100\11665-7\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | TAPV | meas | m_hepVeinLeft | r_TAPV | unilateral | 125100\11692-1\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | PSV | meas | m_hepVeinLeft | r_PSV | unilateral | 125100\11726-7\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | PI | meas | m_hepVeinLeft | r_PI | unilateral | 125100\12008-9\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | RI | meas | m_hepVeinLeft | r_RI | unilateral | 125100\12023-8\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | S/D | meas | m_hepVeinLeft | r_SD | unilateral | 125100\12144-2\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | AI | meas | m_hepVeinLeft | r_AccelIndex | unilateral | 125100\20167-3\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | AT | meas | m_hepVeinLeft | r_AccelTime | unilateral | 125100\20168-1\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | DT | meas | m_hepVeinLeft | r_DecelTime | unilateral | 125100\20217-6\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | PG | meas | m_hepVeinLeft | r_PG | unilateral | 125100\20247-3\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | MG | meas | m_hepVeinLeft | r_PGmean | unilateral | 125100\20256-4\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | TAMV | meas | m_hepVeinLeft | r_TAMV | unilateral | 125100\20352-1\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | VTI | meas | m_hepVeinLeft | r_VTI | unilateral | 125100\20354-7\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | HR | meas | m_hepVeinLeft | r_HeartRate | unilateral | 125100\8867-4\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | L Hepatic V | Dist | meas | m_hepVeinLeft | r_Distance | unilateral | 125100\G-0364\T-48727\T-487A0 | SRT, T-487A0, Vein of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48727, Left Hepatic Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | ? | meas | m_midportvein | r_DopAngle | unilateral | 125100\C12122-01\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | EDV | meas | m_midportvein | r_EDV | unilateral | 125100\11653-3\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | MDV | meas | m_midportvein | r_MDV | unilateral | 125100\11665-7\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | TAPV | meas | m_midportvein | r_TAPV | unilateral | 125100\11692-1\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | PSV | meas | m_midportvein | r_PSV | unilateral | 125100\11726-7\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | PI | meas | m_midportvein | r_PI | unilateral | 125100\12008-9\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | RI | meas | m_midportvein | r_RI | unilateral | 125100\12023-8\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | S/D | meas | m_midportvein | r_SD | unilateral | 125100\12144-2\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | AI | meas | m_midportvein | r_AccelIndex | unilateral | 125100\20167-3\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | AT | meas | m_midportvein | r_AccelTime | unilateral | 125100\20168-1\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | DT | meas | m_midportvein | r_DecelTime | unilateral | 125100\20217-6\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | PG | meas | m_midportvein | r_PG | unilateral | 125100\20247-3\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | MG | meas | m_midportvein | r_PGmean | unilateral | 125100\20256-4\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | TAMV | meas | m_midportvein | r_TAMV | unilateral | 125100\20352-1\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | VTI | meas | m_midportvein | r_VTI | unilateral | 125100\20354-7\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | HR | meas | m_midportvein | r_HeartRate | unilateral | 125100\8867-4\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | M Port V | Dist | meas | m_midportvein | r_Distance | unilateral | 125100\G-0364\T-487A0\T-48810 | SRT, T-487A0, Vein of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48810, Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | ? | meas | m_portVeinLeft | r_DopAngle | unilateral | 125100\C12122-01\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | EDV | meas | m_portVeinLeft | r_EDV | unilateral | 125100\11653-3\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | MDV | meas | m_portVeinLeft | r_MDV | unilateral | 125100\11665-7\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | TAPV | meas | m_portVeinLeft | r_TAPV | unilateral | 125100\11692-1\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | PSV | meas | m_portVeinLeft | r_PSV | unilateral | 125100\11726-7\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | PI | meas | m_portVeinLeft | r_PI | unilateral | 125100\12008-9\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | RI | meas | m_portVeinLeft | r_RI | unilateral | 125100\12023-8\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | S/D | meas | m_portVeinLeft | r_SD | unilateral | 125100\12144-2\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | AI | meas | m_portVeinLeft | r_AccelIndex | unilateral | 125100\20167-3\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | AT | meas | m_portVeinLeft | r_AccelTime | unilateral | 125100\20168-1\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | DT | meas | m_portVeinLeft | r_DecelTime | unilateral | 125100\20217-6\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | PG | meas | m_portVeinLeft | r_PG | unilateral | 125100\20247-3\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | MG | meas | m_portVeinLeft | r_PGmean | unilateral | 125100\20256-4\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | TAMV | meas | m_portVeinLeft | r_TAMV | unilateral | 125100\20352-1\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | VTI | meas | m_portVeinLeft | r_VTI | unilateral | 125100\20354-7\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | HR | meas | m_portVeinLeft | r_HeartRate | unilateral | 125100\8867-4\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | L Port V | Dist | meas | m_portVeinLeft | r_Distance | unilateral | 125100\G-0364\T-487A0\T-4881F | SRT, T-487A0, Vein of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-4881F, Left Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | ? | meas | m_portVeinRight | r_DopAngle | unilateral | 125100\C12122-01\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | EDV | meas | m_portVeinRight | r_EDV | unilateral | 125100\11653-3\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | MDV | meas | m_portVeinRight | r_MDV | unilateral | 125100\11665-7\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | TAPV | meas | m_portVeinRight | r_TAPV | unilateral | 125100\11692-1\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | PSV | meas | m_portVeinRight | r_PSV | unilateral | 125100\11726-7\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | PI | meas | m_portVeinRight | r_PI | unilateral | 125100\12008-9\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | RI | meas | m_portVeinRight | r_RI | unilateral | 125100\12023-8\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | S/D | meas | m_portVeinRight | r_SD | unilateral | 125100\12144-2\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | AI | meas | m_portVeinRight | r_AccelIndex | unilateral | 125100\20167-3\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | AT | meas | m_portVeinRight | r_AccelTime | unilateral | 125100\20168-1\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | DT | meas | m_portVeinRight | r_DecelTime | unilateral | 125100\20217-6\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | PG | meas | m_portVeinRight | r_PG | unilateral | 125100\20247-3\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | MG | meas | m_portVeinRight | r_PGmean | unilateral | 125100\20256-4\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | TAMV | meas | m_portVeinRight | r_TAMV | unilateral | 125100\20352-1\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | VTI | meas | m_portVeinRight | r_VTI | unilateral | 125100\20354-7\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | HR | meas | m_portVeinRight | r_HeartRate | unilateral | 125100\8867-4\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | R Port V | Dist | meas | m_portVeinRight | r_Distance | unilateral | 125100\G-0364\T-487A0\T-4882A | SRT, T-487A0, Vein of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-4882A, Right Main Branch of Portal Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | ? | meas | m_smv | r_DopAngle | unilateral | 125100\C12122-01\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | EDV | meas | m_smv | r_EDV | unilateral | 125100\11653-3\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | MDV | meas | m_smv | r_MDV | unilateral | 125100\11665-7\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | TAPV | meas | m_smv | r_TAPV | unilateral | 125100\11692-1\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | PSV | meas | m_smv | r_PSV | unilateral | 125100\11726-7\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | PI | meas | m_smv | r_PI | unilateral | 125100\12008-9\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | RI | meas | m_smv | r_RI | unilateral | 125100\12023-8\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | S/D | meas | m_smv | r_SD | unilateral | 125100\12144-2\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | AI | meas | m_smv | r_AccelIndex | unilateral | 125100\20167-3\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | AT | meas | m_smv | r_AccelTime | unilateral | 125100\20168-1\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | DT | meas | m_smv | r_DecelTime | unilateral | 125100\20217-6\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | PG | meas | m_smv | r_PG | unilateral | 125100\20247-3\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | MG | meas | m_smv | r_PGmean | unilateral | 125100\20256-4\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | TAMV | meas | m_smv | r_TAMV | unilateral | 125100\20352-1\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | VTI | meas | m_smv | r_VTI | unilateral | 125100\20354-7\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | HR | meas | m_smv | r_HeartRate | unilateral | 125100\8867-4\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | SMV | Dist | meas | m_smv | r_Distance | unilateral | 125100\G-0364\T-487A0\T-48840 | SRT, T-487A0, Vein of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48840, Superior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | ? | meas | m_splenicvein | r_DopAngle | unilateral | 125100\C12122-01\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | EDV | meas | m_splenicvein | r_EDV | unilateral | 125100\11653-3\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | MDV | meas | m_splenicvein | r_MDV | unilateral | 125100\11665-7\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | TAPV | meas | m_splenicvein | r_TAPV | unilateral | 125100\11692-1\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | PSV | meas | m_splenicvein | r_PSV | unilateral | 125100\11726-7\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | PI | meas | m_splenicvein | r_PI | unilateral | 125100\12008-9\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | RI | meas | m_splenicvein | r_RI | unilateral | 125100\12023-8\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | S/D | meas | m_splenicvein | r_SD | unilateral | 125100\12144-2\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | AI | meas | m_splenicvein | r_AccelIndex | unilateral | 125100\20167-3\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | AT | meas | m_splenicvein | r_AccelTime | unilateral | 125100\20168-1\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | DT | meas | m_splenicvein | r_DecelTime | unilateral | 125100\20217-6\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | PG | meas | m_splenicvein | r_PG | unilateral | 125100\20247-3\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | MG | meas | m_splenicvein | r_PGmean | unilateral | 125100\20256-4\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | TAMV | meas | m_splenicvein | r_TAMV | unilateral | 125100\20352-1\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | VTI | meas | m_splenicvein | r_VTI | unilateral | 125100\20354-7\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | HR | meas | m_splenicvein | r_HeartRate | unilateral | 125100\8867-4\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | Splenic V | Dist | meas | m_splenicvein | r_Distance | unilateral | 125100\G-0364\T-487A0\T-48890 | SRT, T-487A0, Vein of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48890, Splenic Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | ? | meas | m_imv | r_DopAngle | unilateral | 125100\C12122-01\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | EDV | meas | m_imv | r_EDV | unilateral | 125100\11653-3\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | MDV | meas | m_imv | r_MDV | unilateral | 125100\11665-7\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | TAPV | meas | m_imv | r_TAPV | unilateral | 125100\11692-1\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | PSV | meas | m_imv | r_PSV | unilateral | 125100\11726-7\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | PI | meas | m_imv | r_PI | unilateral | 125100\12008-9\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | RI | meas | m_imv | r_RI | unilateral | 125100\12023-8\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | S/D | meas | m_imv | r_SD | unilateral | 125100\12144-2\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | AI | meas | m_imv | r_AccelIndex | unilateral | 125100\20167-3\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | AT | meas | m_imv | r_AccelTime | unilateral | 125100\20168-1\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | DT | meas | m_imv | r_DecelTime | unilateral | 125100\20217-6\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | PG | meas | m_imv | r_PG | unilateral | 125100\20247-3\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | MG | meas | m_imv | r_PGmean | unilateral | 125100\20256-4\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | TAMV | meas | m_imv | r_TAMV | unilateral | 125100\20352-1\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | VTI | meas | m_imv | r_VTI | unilateral | 125100\20354-7\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | HR | meas | m_imv | r_HeartRate | unilateral | 125100\8867-4\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Abdominal | IMV | Dist | meas | m_imv | r_Distance | unilateral | 125100\G-0364\T-487A0\T-48910 | SRT, T-487A0, Vein of Abdomen |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48910, Inferior Mesenteric Vein
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | ? | meas | m_Prox_Basilic_V | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | EDV | meas | m_Prox_Basilic_V | r_EDV | bilateral | 125100\11653-3\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | MDV | meas | m_Prox_Basilic_V | r_MDV | bilateral | 125100\11665-7\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | TAPV | meas | m_Prox_Basilic_V | r_TAPV | bilateral | 125100\11692-1\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | PSV | meas | m_Prox_Basilic_V | r_PSV | bilateral | 125100\11726-7\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | PI | meas | m_Prox_Basilic_V | r_PI | bilateral | 125100\12008-9\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | RI | meas | m_Prox_Basilic_V | r_RI | bilateral | 125100\12023-8\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | S/D | meas | m_Prox_Basilic_V | r_SD | bilateral | 125100\12144-2\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | AI | meas | m_Prox_Basilic_V | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | AT | meas | m_Prox_Basilic_V | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | DT | meas | m_Prox_Basilic_V | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | PG | meas | m_Prox_Basilic_V | r_PG | bilateral | 125100\20247-3\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | MG | meas | m_Prox_Basilic_V | r_PGmean | bilateral | 125100\20256-4\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | TAMV | meas | m_Prox_Basilic_V | r_TAMV | bilateral | 125100\20352-1\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | VTI | meas | m_Prox_Basilic_V | r_VTI | bilateral | 125100\20354-7\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | HR | meas | m_Prox_Basilic_V | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilic V | Dist | meas | m_Prox_Basilic_V | r_Distance | bilateral | 125100\G-0364\G-A118\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | ? | meas | m_Dist_Basilic_V | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | EDV | meas | m_Dist_Basilic_V | r_EDV | bilateral | 125100\11653-3\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | MDV | meas | m_Dist_Basilic_V | r_MDV | bilateral | 125100\11665-7\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | TAPV | meas | m_Dist_Basilic_V | r_TAPV | bilateral | 125100\11692-1\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | PSV | meas | m_Dist_Basilic_V | r_PSV | bilateral | 125100\11726-7\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | PI | meas | m_Dist_Basilic_V | r_PI | bilateral | 125100\12008-9\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | RI | meas | m_Dist_Basilic_V | r_RI | bilateral | 125100\12023-8\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | S/D | meas | m_Dist_Basilic_V | r_SD | bilateral | 125100\12144-2\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | AI | meas | m_Dist_Basilic_V | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | AT | meas | m_Dist_Basilic_V | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | DT | meas | m_Dist_Basilic_V | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | PG | meas | m_Dist_Basilic_V | r_PG | bilateral | 125100\20247-3\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | MG | meas | m_Dist_Basilic_V | r_PGmean | bilateral | 125100\20256-4\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | TAMV | meas | m_Dist_Basilic_V | r_TAMV | bilateral | 125100\20352-1\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | VTI | meas | m_Dist_Basilic_V | r_VTI | bilateral | 125100\20354-7\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | HR | meas | m_Dist_Basilic_V | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilic V | Dist | meas | m_Dist_Basilic_V | r_Distance | bilateral | 125100\G-0364\G-A119\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | ? | meas | m_Mid_Basilic_V | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | EDV | meas | m_Mid_Basilic_V | r_EDV | bilateral | 125100\11653-3\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | MDV | meas | m_Mid_Basilic_V | r_MDV | bilateral | 125100\11665-7\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | TAPV | meas | m_Mid_Basilic_V | r_TAPV | bilateral | 125100\11692-1\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | PSV | meas | m_Mid_Basilic_V | r_PSV | bilateral | 125100\11726-7\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | PI | meas | m_Mid_Basilic_V | r_PI | bilateral | 125100\12008-9\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | RI | meas | m_Mid_Basilic_V | r_RI | bilateral | 125100\12023-8\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | S/D | meas | m_Mid_Basilic_V | r_SD | bilateral | 125100\12144-2\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | AI | meas | m_Mid_Basilic_V | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | AT | meas | m_Mid_Basilic_V | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | DT | meas | m_Mid_Basilic_V | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | PG | meas | m_Mid_Basilic_V | r_PG | bilateral | 125100\20247-3\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | MG | meas | m_Mid_Basilic_V | r_PGmean | bilateral | 125100\20256-4\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | TAMV | meas | m_Mid_Basilic_V | r_TAMV | bilateral | 125100\20352-1\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | VTI | meas | m_Mid_Basilic_V | r_VTI | bilateral | 125100\20354-7\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | HR | meas | m_Mid_Basilic_V | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilic V | Dist | meas | m_Mid_Basilic_V | r_Distance | bilateral | 125100\G-0364\G-A188\T-48052\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48052, Basilic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | IJV | ? | meas | m_Int_Jugular | r_DopAngle | bilateral | 125100\C12122-01\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | IJV | EDV | meas | m_Int_Jugular | r_EDV | bilateral | 125100\11653-3\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | IJV | MDV | meas | m_Int_Jugular | r_MDV | bilateral | 125100\11665-7\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | IJV | TAPV | meas | m_Int_Jugular | r_TAPV | bilateral | 125100\11692-1\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | IJV | PSV | meas | m_Int_Jugular | r_PSV | bilateral | 125100\11726-7\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | IJV | PI | meas | m_Int_Jugular | r_PI | bilateral | 125100\12008-9\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | IJV | RI | meas | m_Int_Jugular | r_RI | bilateral | 125100\12023-8\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | IJV | S/D | meas | m_Int_Jugular | r_SD | bilateral | 125100\12144-2\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | IJV | AI | meas | m_Int_Jugular | r_AccelIndex | bilateral | 125100\20167-3\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | IJV | AT | meas | m_Int_Jugular | r_AccelTime | bilateral | 125100\20168-1\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | IJV | DT | meas | m_Int_Jugular | r_DecelTime | bilateral | 125100\20217-6\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | IJV | PG | meas | m_Int_Jugular | r_PG | bilateral | 125100\20247-3\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | IJV | MG | meas | m_Int_Jugular | r_PGmean | bilateral | 125100\20256-4\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | IJV | TAMV | meas | m_Int_Jugular | r_TAMV | bilateral | 125100\20352-1\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | IJV | VTI | meas | m_Int_Jugular | r_VTI | bilateral | 125100\20354-7\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | IJV | HR | meas | m_Int_Jugular | r_HeartRate | bilateral | 125100\8867-4\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | IJV | Dist | meas | m_Int_Jugular | r_Distance | bilateral | 125100\G-0364\T-48170\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48170, Internal Jugular vein
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | ? | meas | m_Prox_SCL_V | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | EDV | meas | m_Prox_SCL_V | r_EDV | bilateral | 125100\11653-3\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | MDV | meas | m_Prox_SCL_V | r_MDV | bilateral | 125100\11665-7\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | TAPV | meas | m_Prox_SCL_V | r_TAPV | bilateral | 125100\11692-1\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | PSV | meas | m_Prox_SCL_V | r_PSV | bilateral | 125100\11726-7\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | PI | meas | m_Prox_SCL_V | r_PI | bilateral | 125100\12008-9\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | RI | meas | m_Prox_SCL_V | r_RI | bilateral | 125100\12023-8\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | S/D | meas | m_Prox_SCL_V | r_SD | bilateral | 125100\12144-2\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | AI | meas | m_Prox_SCL_V | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | AT | meas | m_Prox_SCL_V | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | DT | meas | m_Prox_SCL_V | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | PG | meas | m_Prox_SCL_V | r_PG | bilateral | 125100\20247-3\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | MG | meas | m_Prox_SCL_V | r_PGmean | bilateral | 125100\20256-4\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | TAMV | meas | m_Prox_SCL_V | r_TAMV | bilateral | 125100\20352-1\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | VTI | meas | m_Prox_SCL_V | r_VTI | bilateral | 125100\20354-7\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | HR | meas | m_Prox_SCL_V | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL V | Dist | meas | m_Prox_SCL_V | r_Distance | bilateral | 125100\G-0364\G-A118\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | ? | meas | m_Mid_SCL_V | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | EDV | meas | m_Mid_SCL_V | r_EDV | bilateral | 125100\11653-3\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | MDV | meas | m_Mid_SCL_V | r_MDV | bilateral | 125100\11665-7\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | TAPV | meas | m_Mid_SCL_V | r_TAPV | bilateral | 125100\11692-1\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | PSV | meas | m_Mid_SCL_V | r_PSV | bilateral | 125100\11726-7\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | PI | meas | m_Mid_SCL_V | r_PI | bilateral | 125100\12008-9\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | RI | meas | m_Mid_SCL_V | r_RI | bilateral | 125100\12023-8\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | S/D | meas | m_Mid_SCL_V | r_SD | bilateral | 125100\12144-2\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | AI | meas | m_Mid_SCL_V | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | AT | meas | m_Mid_SCL_V | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | DT | meas | m_Mid_SCL_V | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | PG | meas | m_Mid_SCL_V | r_PG | bilateral | 125100\20247-3\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | MG | meas | m_Mid_SCL_V | r_PGmean | bilateral | 125100\20256-4\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | TAMV | meas | m_Mid_SCL_V | r_TAMV | bilateral | 125100\20352-1\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | VTI | meas | m_Mid_SCL_V | r_VTI | bilateral | 125100\20354-7\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | HR | meas | m_Mid_SCL_V | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL V | Dist | meas | m_Mid_SCL_V | r_Distance | bilateral | 125100\G-0364\G-A188\T-48330\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48330, Subclavian vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | ? | meas | m_Brach_Ceph_V | r_DopAngle | bilateral | 125100\C12122-01\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | EDV | meas | m_Brach_Ceph_V | r_EDV | bilateral | 125100\11653-3\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | MDV | meas | m_Brach_Ceph_V | r_MDV | bilateral | 125100\11665-7\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | TAPV | meas | m_Brach_Ceph_V | r_TAPV | bilateral | 125100\11692-1\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | PSV | meas | m_Brach_Ceph_V | r_PSV | bilateral | 125100\11726-7\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | PI | meas | m_Brach_Ceph_V | r_PI | bilateral | 125100\12008-9\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | RI | meas | m_Brach_Ceph_V | r_RI | bilateral | 125100\12023-8\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | S/D | meas | m_Brach_Ceph_V | r_SD | bilateral | 125100\12144-2\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | AI | meas | m_Brach_Ceph_V | r_AccelIndex | bilateral | 125100\20167-3\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | AT | meas | m_Brach_Ceph_V | r_AccelTime | bilateral | 125100\20168-1\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | DT | meas | m_Brach_Ceph_V | r_DecelTime | bilateral | 125100\20217-6\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | PG | meas | m_Brach_Ceph_V | r_PG | bilateral | 125100\20247-3\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | MG | meas | m_Brach_Ceph_V | r_PGmean | bilateral | 125100\20256-4\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | TAMV | meas | m_Brach_Ceph_V | r_TAMV | bilateral | 125100\20352-1\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | VTI | meas | m_Brach_Ceph_V | r_VTI | bilateral | 125100\20354-7\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | HR | meas | m_Brach_Ceph_V | r_HeartRate | bilateral | 125100\8867-4\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Brachioceph V | Dist | meas | m_Brach_Ceph_V | r_Distance | bilateral | 125100\G-0364\T-48620\T-49103 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48620, Innominate vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | ? | meas | m_Axillary_V | r_DopAngle | bilateral | 125100\C12122-01\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | EDV | meas | m_Axillary_V | r_EDV | bilateral | 125100\11653-3\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | MDV | meas | m_Axillary_V | r_MDV | bilateral | 125100\11665-7\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | TAPV | meas | m_Axillary_V | r_TAPV | bilateral | 125100\11692-1\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | PSV | meas | m_Axillary_V | r_PSV | bilateral | 125100\11726-7\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | PI | meas | m_Axillary_V | r_PI | bilateral | 125100\12008-9\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | RI | meas | m_Axillary_V | r_RI | bilateral | 125100\12023-8\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | S/D | meas | m_Axillary_V | r_SD | bilateral | 125100\12144-2\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | AI | meas | m_Axillary_V | r_AccelIndex | bilateral | 125100\20167-3\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | AT | meas | m_Axillary_V | r_AccelTime | bilateral | 125100\20168-1\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | DT | meas | m_Axillary_V | r_DecelTime | bilateral | 125100\20217-6\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | PG | meas | m_Axillary_V | r_PG | bilateral | 125100\20247-3\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | MG | meas | m_Axillary_V | r_PGmean | bilateral | 125100\20256-4\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | TAMV | meas | m_Axillary_V | r_TAMV | bilateral | 125100\20352-1\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | VTI | meas | m_Axillary_V | r_VTI | bilateral | 125100\20354-7\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | HR | meas | m_Axillary_V | r_HeartRate | bilateral | 125100\8867-4\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Ax V | Dist | meas | m_Axillary_V | r_Distance | bilateral | 125100\G-0364\T-49103\T-49110 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49110, Axillary vein
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | ? | meas | m_Prox_Cephalic_V | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | EDV | meas | m_Prox_Cephalic_V | r_EDV | bilateral | 125100\11653-3\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | MDV | meas | m_Prox_Cephalic_V | r_MDV | bilateral | 125100\11665-7\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | TAPV | meas | m_Prox_Cephalic_V | r_TAPV | bilateral | 125100\11692-1\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | PSV | meas | m_Prox_Cephalic_V | r_PSV | bilateral | 125100\11726-7\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | PI | meas | m_Prox_Cephalic_V | r_PI | bilateral | 125100\12008-9\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | RI | meas | m_Prox_Cephalic_V | r_RI | bilateral | 125100\12023-8\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | S/D | meas | m_Prox_Cephalic_V | r_SD | bilateral | 125100\12144-2\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | AI | meas | m_Prox_Cephalic_V | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | AT | meas | m_Prox_Cephalic_V | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | DT | meas | m_Prox_Cephalic_V | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | PG | meas | m_Prox_Cephalic_V | r_PG | bilateral | 125100\20247-3\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | MG | meas | m_Prox_Cephalic_V | r_PGmean | bilateral | 125100\20256-4\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | TAMV | meas | m_Prox_Cephalic_V | r_TAMV | bilateral | 125100\20352-1\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | VTI | meas | m_Prox_Cephalic_V | r_VTI | bilateral | 125100\20354-7\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | HR | meas | m_Prox_Cephalic_V | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ceph V | Dist | meas | m_Prox_Cephalic_V | r_Distance | bilateral | 125100\G-0364\G-A118\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | ? | meas | m_Dist_Cephalic_V | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | EDV | meas | m_Dist_Cephalic_V | r_EDV | bilateral | 125100\11653-3\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | MDV | meas | m_Dist_Cephalic_V | r_MDV | bilateral | 125100\11665-7\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | TAPV | meas | m_Dist_Cephalic_V | r_TAPV | bilateral | 125100\11692-1\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | PSV | meas | m_Dist_Cephalic_V | r_PSV | bilateral | 125100\11726-7\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | PI | meas | m_Dist_Cephalic_V | r_PI | bilateral | 125100\12008-9\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | RI | meas | m_Dist_Cephalic_V | r_RI | bilateral | 125100\12023-8\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | S/D | meas | m_Dist_Cephalic_V | r_SD | bilateral | 125100\12144-2\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | AI | meas | m_Dist_Cephalic_V | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | AT | meas | m_Dist_Cephalic_V | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | DT | meas | m_Dist_Cephalic_V | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | PG | meas | m_Dist_Cephalic_V | r_PG | bilateral | 125100\20247-3\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | MG | meas | m_Dist_Cephalic_V | r_PGmean | bilateral | 125100\20256-4\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | TAMV | meas | m_Dist_Cephalic_V | r_TAMV | bilateral | 125100\20352-1\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | VTI | meas | m_Dist_Cephalic_V | r_VTI | bilateral | 125100\20354-7\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | HR | meas | m_Dist_Cephalic_V | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ceph V | Dist | meas | m_Dist_Cephalic_V | r_Distance | bilateral | 125100\G-0364\G-A119\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | ? | meas | m_Mid_Cephalic_V | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | EDV | meas | m_Mid_Cephalic_V | r_EDV | bilateral | 125100\11653-3\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | MDV | meas | m_Mid_Cephalic_V | r_MDV | bilateral | 125100\11665-7\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | TAPV | meas | m_Mid_Cephalic_V | r_TAPV | bilateral | 125100\11692-1\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | PSV | meas | m_Mid_Cephalic_V | r_PSV | bilateral | 125100\11726-7\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | PI | meas | m_Mid_Cephalic_V | r_PI | bilateral | 125100\12008-9\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | RI | meas | m_Mid_Cephalic_V | r_RI | bilateral | 125100\12023-8\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | S/D | meas | m_Mid_Cephalic_V | r_SD | bilateral | 125100\12144-2\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | AI | meas | m_Mid_Cephalic_V | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | AT | meas | m_Mid_Cephalic_V | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | DT | meas | m_Mid_Cephalic_V | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | PG | meas | m_Mid_Cephalic_V | r_PG | bilateral | 125100\20247-3\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | MG | meas | m_Mid_Cephalic_V | r_PGmean | bilateral | 125100\20256-4\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | TAMV | meas | m_Mid_Cephalic_V | r_TAMV | bilateral | 125100\20352-1\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | VTI | meas | m_Mid_Cephalic_V | r_VTI | bilateral | 125100\20354-7\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | HR | meas | m_Mid_Cephalic_V | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ceph V | Dist | meas | m_Mid_Cephalic_V | r_Distance | bilateral | 125100\G-0364\G-A188\T-49103\T-49240 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49240, Cephalic vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | ? | meas | m_Prox_Ulnar_V | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | EDV | meas | m_Prox_Ulnar_V | r_EDV | bilateral | 125100\11653-3\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | MDV | meas | m_Prox_Ulnar_V | r_MDV | bilateral | 125100\11665-7\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | TAPV | meas | m_Prox_Ulnar_V | r_TAPV | bilateral | 125100\11692-1\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | PSV | meas | m_Prox_Ulnar_V | r_PSV | bilateral | 125100\11726-7\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | PI | meas | m_Prox_Ulnar_V | r_PI | bilateral | 125100\12008-9\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | RI | meas | m_Prox_Ulnar_V | r_RI | bilateral | 125100\12023-8\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | S/D | meas | m_Prox_Ulnar_V | r_SD | bilateral | 125100\12144-2\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | AI | meas | m_Prox_Ulnar_V | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | AT | meas | m_Prox_Ulnar_V | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | DT | meas | m_Prox_Ulnar_V | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | PG | meas | m_Prox_Ulnar_V | r_PG | bilateral | 125100\20247-3\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | MG | meas | m_Prox_Ulnar_V | r_PGmean | bilateral | 125100\20256-4\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | TAMV | meas | m_Prox_Ulnar_V | r_TAMV | bilateral | 125100\20352-1\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | VTI | meas | m_Prox_Ulnar_V | r_VTI | bilateral | 125100\20354-7\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | HR | meas | m_Prox_Ulnar_V | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar V | Dist | meas | m_Prox_Ulnar_V | r_Distance | bilateral | 125100\G-0364\G-A118\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | ? | meas | m_Dist_Ulnar_V | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | EDV | meas | m_Dist_Ulnar_V | r_EDV | bilateral | 125100\11653-3\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | MDV | meas | m_Dist_Ulnar_V | r_MDV | bilateral | 125100\11665-7\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | TAPV | meas | m_Dist_Ulnar_V | r_TAPV | bilateral | 125100\11692-1\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | PSV | meas | m_Dist_Ulnar_V | r_PSV | bilateral | 125100\11726-7\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | PI | meas | m_Dist_Ulnar_V | r_PI | bilateral | 125100\12008-9\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | RI | meas | m_Dist_Ulnar_V | r_RI | bilateral | 125100\12023-8\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | S/D | meas | m_Dist_Ulnar_V | r_SD | bilateral | 125100\12144-2\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | AI | meas | m_Dist_Ulnar_V | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | AT | meas | m_Dist_Ulnar_V | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | DT | meas | m_Dist_Ulnar_V | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | PG | meas | m_Dist_Ulnar_V | r_PG | bilateral | 125100\20247-3\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | MG | meas | m_Dist_Ulnar_V | r_PGmean | bilateral | 125100\20256-4\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | TAMV | meas | m_Dist_Ulnar_V | r_TAMV | bilateral | 125100\20352-1\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | VTI | meas | m_Dist_Ulnar_V | r_VTI | bilateral | 125100\20354-7\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | HR | meas | m_Dist_Ulnar_V | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar V | Dist | meas | m_Dist_Ulnar_V | r_Distance | bilateral | 125100\G-0364\G-A119\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | ? | meas | m_Mid_Ulnar_V | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | EDV | meas | m_Mid_Ulnar_V | r_EDV | bilateral | 125100\11653-3\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | MDV | meas | m_Mid_Ulnar_V | r_MDV | bilateral | 125100\11665-7\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | TAPV | meas | m_Mid_Ulnar_V | r_TAPV | bilateral | 125100\11692-1\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | PSV | meas | m_Mid_Ulnar_V | r_PSV | bilateral | 125100\11726-7\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | PI | meas | m_Mid_Ulnar_V | r_PI | bilateral | 125100\12008-9\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | RI | meas | m_Mid_Ulnar_V | r_RI | bilateral | 125100\12023-8\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | S/D | meas | m_Mid_Ulnar_V | r_SD | bilateral | 125100\12144-2\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | AI | meas | m_Mid_Ulnar_V | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | AT | meas | m_Mid_Ulnar_V | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | DT | meas | m_Mid_Ulnar_V | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | PG | meas | m_Mid_Ulnar_V | r_PG | bilateral | 125100\20247-3\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | MG | meas | m_Mid_Ulnar_V | r_PGmean | bilateral | 125100\20256-4\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | TAMV | meas | m_Mid_Ulnar_V | r_TAMV | bilateral | 125100\20352-1\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | VTI | meas | m_Mid_Ulnar_V | r_VTI | bilateral | 125100\20354-7\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | HR | meas | m_Mid_Ulnar_V | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar V | Dist | meas | m_Mid_Ulnar_V | r_Distance | bilateral | 125100\G-0364\G-A188\T-49103\T-49330 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49330, Ulnar vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | ? | meas | m_Prox_Radial_V | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | EDV | meas | m_Prox_Radial_V | r_EDV | bilateral | 125100\11653-3\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | MDV | meas | m_Prox_Radial_V | r_MDV | bilateral | 125100\11665-7\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | TAPV | meas | m_Prox_Radial_V | r_TAPV | bilateral | 125100\11692-1\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | PSV | meas | m_Prox_Radial_V | r_PSV | bilateral | 125100\11726-7\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | PI | meas | m_Prox_Radial_V | r_PI | bilateral | 125100\12008-9\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | RI | meas | m_Prox_Radial_V | r_RI | bilateral | 125100\12023-8\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | S/D | meas | m_Prox_Radial_V | r_SD | bilateral | 125100\12144-2\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | AI | meas | m_Prox_Radial_V | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | AT | meas | m_Prox_Radial_V | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | DT | meas | m_Prox_Radial_V | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | PG | meas | m_Prox_Radial_V | r_PG | bilateral | 125100\20247-3\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | MG | meas | m_Prox_Radial_V | r_PGmean | bilateral | 125100\20256-4\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | TAMV | meas | m_Prox_Radial_V | r_TAMV | bilateral | 125100\20352-1\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | VTI | meas | m_Prox_Radial_V | r_VTI | bilateral | 125100\20354-7\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | HR | meas | m_Prox_Radial_V | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad V | Dist | meas | m_Prox_Radial_V | r_Distance | bilateral | 125100\G-0364\G-A118\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | ? | meas | m_Dist_Radial_V | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | EDV | meas | m_Dist_Radial_V | r_EDV | bilateral | 125100\11653-3\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | MDV | meas | m_Dist_Radial_V | r_MDV | bilateral | 125100\11665-7\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | TAPV | meas | m_Dist_Radial_V | r_TAPV | bilateral | 125100\11692-1\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | PSV | meas | m_Dist_Radial_V | r_PSV | bilateral | 125100\11726-7\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | PI | meas | m_Dist_Radial_V | r_PI | bilateral | 125100\12008-9\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | RI | meas | m_Dist_Radial_V | r_RI | bilateral | 125100\12023-8\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | S/D | meas | m_Dist_Radial_V | r_SD | bilateral | 125100\12144-2\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | AI | meas | m_Dist_Radial_V | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | AT | meas | m_Dist_Radial_V | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | DT | meas | m_Dist_Radial_V | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | PG | meas | m_Dist_Radial_V | r_PG | bilateral | 125100\20247-3\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | MG | meas | m_Dist_Radial_V | r_PGmean | bilateral | 125100\20256-4\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | TAMV | meas | m_Dist_Radial_V | r_TAMV | bilateral | 125100\20352-1\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | VTI | meas | m_Dist_Radial_V | r_VTI | bilateral | 125100\20354-7\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | HR | meas | m_Dist_Radial_V | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad V | Dist | meas | m_Dist_Radial_V | r_Distance | bilateral | 125100\G-0364\G-A119\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | ? | meas | m_Mid_Radial_V | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | EDV | meas | m_Mid_Radial_V | r_EDV | bilateral | 125100\11653-3\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | MDV | meas | m_Mid_Radial_V | r_MDV | bilateral | 125100\11665-7\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | TAPV | meas | m_Mid_Radial_V | r_TAPV | bilateral | 125100\11692-1\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | PSV | meas | m_Mid_Radial_V | r_PSV | bilateral | 125100\11726-7\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | PI | meas | m_Mid_Radial_V | r_PI | bilateral | 125100\12008-9\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | RI | meas | m_Mid_Radial_V | r_RI | bilateral | 125100\12023-8\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | S/D | meas | m_Mid_Radial_V | r_SD | bilateral | 125100\12144-2\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | AI | meas | m_Mid_Radial_V | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | AT | meas | m_Mid_Radial_V | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | DT | meas | m_Mid_Radial_V | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | PG | meas | m_Mid_Radial_V | r_PG | bilateral | 125100\20247-3\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | MG | meas | m_Mid_Radial_V | r_PGmean | bilateral | 125100\20256-4\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | TAMV | meas | m_Mid_Radial_V | r_TAMV | bilateral | 125100\20352-1\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | VTI | meas | m_Mid_Radial_V | r_VTI | bilateral | 125100\20354-7\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | HR | meas | m_Mid_Radial_V | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad V | Dist | meas | m_Mid_Radial_V | r_Distance | bilateral | 125100\G-0364\G-A188\T-49103\T-49340 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49340, Radial vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | ? | meas | m_Prox_Brachial_V | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | EDV | meas | m_Prox_Brachial_V | r_EDV | bilateral | 125100\11653-3\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | MDV | meas | m_Prox_Brachial_V | r_MDV | bilateral | 125100\11665-7\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | TAPV | meas | m_Prox_Brachial_V | r_TAPV | bilateral | 125100\11692-1\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | PSV | meas | m_Prox_Brachial_V | r_PSV | bilateral | 125100\11726-7\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | PI | meas | m_Prox_Brachial_V | r_PI | bilateral | 125100\12008-9\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | RI | meas | m_Prox_Brachial_V | r_RI | bilateral | 125100\12023-8\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | S/D | meas | m_Prox_Brachial_V | r_SD | bilateral | 125100\12144-2\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | AI | meas | m_Prox_Brachial_V | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | AT | meas | m_Prox_Brachial_V | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | DT | meas | m_Prox_Brachial_V | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | PG | meas | m_Prox_Brachial_V | r_PG | bilateral | 125100\20247-3\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | MG | meas | m_Prox_Brachial_V | r_PGmean | bilateral | 125100\20256-4\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | TAMV | meas | m_Prox_Brachial_V | r_TAMV | bilateral | 125100\20352-1\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | VTI | meas | m_Prox_Brachial_V | r_VTI | bilateral | 125100\20354-7\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | HR | meas | m_Prox_Brachial_V | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach V | Dist | meas | m_Prox_Brachial_V | r_Distance | bilateral | 125100\G-0364\G-A118\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | ? | meas | m_Dist_Brachial_V | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | EDV | meas | m_Dist_Brachial_V | r_EDV | bilateral | 125100\11653-3\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | MDV | meas | m_Dist_Brachial_V | r_MDV | bilateral | 125100\11665-7\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | TAPV | meas | m_Dist_Brachial_V | r_TAPV | bilateral | 125100\11692-1\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | PSV | meas | m_Dist_Brachial_V | r_PSV | bilateral | 125100\11726-7\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | PI | meas | m_Dist_Brachial_V | r_PI | bilateral | 125100\12008-9\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | RI | meas | m_Dist_Brachial_V | r_RI | bilateral | 125100\12023-8\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | S/D | meas | m_Dist_Brachial_V | r_SD | bilateral | 125100\12144-2\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | AI | meas | m_Dist_Brachial_V | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | AT | meas | m_Dist_Brachial_V | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | DT | meas | m_Dist_Brachial_V | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | PG | meas | m_Dist_Brachial_V | r_PG | bilateral | 125100\20247-3\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | MG | meas | m_Dist_Brachial_V | r_PGmean | bilateral | 125100\20256-4\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | TAMV | meas | m_Dist_Brachial_V | r_TAMV | bilateral | 125100\20352-1\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | VTI | meas | m_Dist_Brachial_V | r_VTI | bilateral | 125100\20354-7\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | HR | meas | m_Dist_Brachial_V | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach V | Dist | meas | m_Dist_Brachial_V | r_Distance | bilateral | 125100\G-0364\G-A119\T-49103\T-49350 | SRT, T-49103, Vein Of Upper Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49350, Brachial vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | ? | meas | m_Prox_SFV | r_DopAngle | bilateral | 125100\C12122-01\G-035A\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | EDV | meas | m_Prox_SFV | r_EDV | bilateral | 125100\11653-3\G-035A\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | MDV | meas | m_Prox_SFV | r_MDV | bilateral | 125100\11665-7\G-035A\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | TAPV | meas | m_Prox_SFV | r_TAPV | bilateral | 125100\11692-1\G-035A\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | PSV | meas | m_Prox_SFV | r_PSV | bilateral | 125100\11726-7\G-035A\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | PI | meas | m_Prox_SFV | r_PI | bilateral | 125100\12008-9\G-035A\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | RI | meas | m_Prox_SFV | r_RI | bilateral | 125100\12023-8\G-035A\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | S/D | meas | m_Prox_SFV | r_SD | bilateral | 125100\12144-2\G-035A\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | AI | meas | m_Prox_SFV | r_AccelIndex | bilateral | 125100\20167-3\G-035A\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | AT | meas | m_Prox_SFV | r_AccelTime | bilateral | 125100\20168-1\G-035A\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | DT | meas | m_Prox_SFV | r_DecelTime | bilateral | 125100\20217-6\G-035A\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | PG | meas | m_Prox_SFV | r_PG | bilateral | 125100\20247-3\G-035A\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | MG | meas | m_Prox_SFV | r_PGmean | bilateral | 125100\20256-4\G-035A\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | TAMV | meas | m_Prox_SFV | r_TAMV | bilateral | 125100\20352-1\G-035A\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | VTI | meas | m_Prox_SFV | r_VTI | bilateral | 125100\20354-7\G-035A\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | HR | meas | m_Prox_SFV | r_HeartRate | bilateral | 125100\8867-4\G-035A\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFV | Dist | meas | m_Prox_SFV | r_Distance | bilateral | 125100\G-035A\G-0364\G-A118\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | ? | meas | m_Dist_SFV | r_DopAngle | bilateral | 125100\C12122-01\G-035A\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | EDV | meas | m_Dist_SFV | r_EDV | bilateral | 125100\11653-3\G-035A\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | MDV | meas | m_Dist_SFV | r_MDV | bilateral | 125100\11665-7\G-035A\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | TAPV | meas | m_Dist_SFV | r_TAPV | bilateral | 125100\11692-1\G-035A\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | PSV | meas | m_Dist_SFV | r_PSV | bilateral | 125100\11726-7\G-035A\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | PI | meas | m_Dist_SFV | r_PI | bilateral | 125100\12008-9\G-035A\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | RI | meas | m_Dist_SFV | r_RI | bilateral | 125100\12023-8\G-035A\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | S/D | meas | m_Dist_SFV | r_SD | bilateral | 125100\12144-2\G-035A\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | AI | meas | m_Dist_SFV | r_AccelIndex | bilateral | 125100\20167-3\G-035A\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | AT | meas | m_Dist_SFV | r_AccelTime | bilateral | 125100\20168-1\G-035A\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | DT | meas | m_Dist_SFV | r_DecelTime | bilateral | 125100\20217-6\G-035A\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | PG | meas | m_Dist_SFV | r_PG | bilateral | 125100\20247-3\G-035A\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | MG | meas | m_Dist_SFV | r_PGmean | bilateral | 125100\20256-4\G-035A\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | TAMV | meas | m_Dist_SFV | r_TAMV | bilateral | 125100\20352-1\G-035A\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | VTI | meas | m_Dist_SFV | r_VTI | bilateral | 125100\20354-7\G-035A\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | HR | meas | m_Dist_SFV | r_HeartRate | bilateral | 125100\8867-4\G-035A\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFV | Dist | meas | m_Dist_SFV | r_Distance | bilateral | 125100\G-035A\G-0364\G-A119\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | ? | meas | m_Mid_SFV | r_DopAngle | bilateral | 125100\C12122-01\G-035A\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | EDV | meas | m_Mid_SFV | r_EDV | bilateral | 125100\11653-3\G-035A\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | MDV | meas | m_Mid_SFV | r_MDV | bilateral | 125100\11665-7\G-035A\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | TAPV | meas | m_Mid_SFV | r_TAPV | bilateral | 125100\11692-1\G-035A\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | PSV | meas | m_Mid_SFV | r_PSV | bilateral | 125100\11726-7\G-035A\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | PI | meas | m_Mid_SFV | r_PI | bilateral | 125100\12008-9\G-035A\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | RI | meas | m_Mid_SFV | r_RI | bilateral | 125100\12023-8\G-035A\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | S/D | meas | m_Mid_SFV | r_SD | bilateral | 125100\12144-2\G-035A\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | AI | meas | m_Mid_SFV | r_AccelIndex | bilateral | 125100\20167-3\G-035A\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | AT | meas | m_Mid_SFV | r_AccelTime | bilateral | 125100\20168-1\G-035A\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | DT | meas | m_Mid_SFV | r_DecelTime | bilateral | 125100\20217-6\G-035A\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | PG | meas | m_Mid_SFV | r_PG | bilateral | 125100\20247-3\G-035A\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | MG | meas | m_Mid_SFV | r_PGmean | bilateral | 125100\20256-4\G-035A\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | TAMV | meas | m_Mid_SFV | r_TAMV | bilateral | 125100\20352-1\G-035A\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | VTI | meas | m_Mid_SFV | r_VTI | bilateral | 125100\20354-7\G-035A\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | HR | meas | m_Mid_SFV | r_HeartRate | bilateral | 125100\8867-4\G-035A\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFV | Dist | meas | m_Mid_SFV | r_Distance | bilateral | 125100\G-035A\G-0364\G-A188\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, G-035A, Superficial Femoral Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | CFV | ? | meas | m_CFV | r_DopAngle | bilateral | 125100\C12122-01\G-035B\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | CFV | EDV | meas | m_CFV | r_EDV | bilateral | 125100\11653-3\G-035B\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | CFV | MDV | meas | m_CFV | r_MDV | bilateral | 125100\11665-7\G-035B\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | CFV | TAPV | meas | m_CFV | r_TAPV | bilateral | 125100\11692-1\G-035B\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | CFV | PSV | meas | m_CFV | r_PSV | bilateral | 125100\11726-7\G-035B\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | CFV | PI | meas | m_CFV | r_PI | bilateral | 125100\12008-9\G-035B\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | CFV | RI | meas | m_CFV | r_RI | bilateral | 125100\12023-8\G-035B\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | CFV | S/D | meas | m_CFV | r_SD | bilateral | 125100\12144-2\G-035B\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | CFV | AI | meas | m_CFV | r_AccelIndex | bilateral | 125100\20167-3\G-035B\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | CFV | AT | meas | m_CFV | r_AccelTime | bilateral | 125100\20168-1\G-035B\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | CFV | DT | meas | m_CFV | r_DecelTime | bilateral | 125100\20217-6\G-035B\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | CFV | PG | meas | m_CFV | r_PG | bilateral | 125100\20247-3\G-035B\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | CFV | MG | meas | m_CFV | r_PGmean | bilateral | 125100\20256-4\G-035B\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | CFV | TAMV | meas | m_CFV | r_TAMV | bilateral | 125100\20352-1\G-035B\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | CFV | VTI | meas | m_CFV | r_VTI | bilateral | 125100\20354-7\G-035B\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | CFV | HR | meas | m_CFV | r_HeartRate | bilateral | 125100\8867-4\G-035B\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | CFV | Dist | meas | m_CFV | r_Distance | bilateral | 125100\G-035B\G-0364\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, G-035B, Common Femoral Vein
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | ? | meas | m_Prox_GSV_Thigh | r_DopAngle | bilateral | 125100\C12122-01\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | EDV | meas | m_Prox_GSV_Thigh | r_EDV | bilateral | 125100\11653-3\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | MDV | meas | m_Prox_GSV_Thigh | r_MDV | bilateral | 125100\11665-7\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | TAPV | meas | m_Prox_GSV_Thigh | r_TAPV | bilateral | 125100\11692-1\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | PSV | meas | m_Prox_GSV_Thigh | r_PSV | bilateral | 125100\11726-7\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | PI | meas | m_Prox_GSV_Thigh | r_PI | bilateral | 125100\12008-9\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | RI | meas | m_Prox_GSV_Thigh | r_RI | bilateral | 125100\12023-8\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | S/D | meas | m_Prox_GSV_Thigh | r_SD | bilateral | 125100\12144-2\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | AI | meas | m_Prox_GSV_Thigh | r_AccelIndex | bilateral | 125100\20167-3\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | AT | meas | m_Prox_GSV_Thigh | r_AccelTime | bilateral | 125100\20168-1\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | DT | meas | m_Prox_GSV_Thigh | r_DecelTime | bilateral | 125100\20217-6\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | PG | meas | m_Prox_GSV_Thigh | r_PG | bilateral | 125100\20247-3\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | MG | meas | m_Prox_GSV_Thigh | r_PGmean | bilateral | 125100\20256-4\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | TAMV | meas | m_Prox_GSV_Thigh | r_TAMV | bilateral | 125100\20352-1\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | VTI | meas | m_Prox_GSV_Thigh | r_VTI | bilateral | 125100\20354-7\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | HR | meas | m_Prox_GSV_Thigh | r_HeartRate | bilateral | 125100\8867-4\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Thigh | Dist | meas | m_Prox_GSV_Thigh | r_Distance | bilateral | 125100\G-0364\G-A118\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | ? | meas | m_Dist_GSV_Thigh | r_DopAngle | bilateral | 125100\C12122-01\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | EDV | meas | m_Dist_GSV_Thigh | r_EDV | bilateral | 125100\11653-3\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | MDV | meas | m_Dist_GSV_Thigh | r_MDV | bilateral | 125100\11665-7\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | TAPV | meas | m_Dist_GSV_Thigh | r_TAPV | bilateral | 125100\11692-1\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | PSV | meas | m_Dist_GSV_Thigh | r_PSV | bilateral | 125100\11726-7\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | PI | meas | m_Dist_GSV_Thigh | r_PI | bilateral | 125100\12008-9\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | RI | meas | m_Dist_GSV_Thigh | r_RI | bilateral | 125100\12023-8\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | S/D | meas | m_Dist_GSV_Thigh | r_SD | bilateral | 125100\12144-2\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | AI | meas | m_Dist_GSV_Thigh | r_AccelIndex | bilateral | 125100\20167-3\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | AT | meas | m_Dist_GSV_Thigh | r_AccelTime | bilateral | 125100\20168-1\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | DT | meas | m_Dist_GSV_Thigh | r_DecelTime | bilateral | 125100\20217-6\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | PG | meas | m_Dist_GSV_Thigh | r_PG | bilateral | 125100\20247-3\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | MG | meas | m_Dist_GSV_Thigh | r_PGmean | bilateral | 125100\20256-4\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | TAMV | meas | m_Dist_GSV_Thigh | r_TAMV | bilateral | 125100\20352-1\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | VTI | meas | m_Dist_GSV_Thigh | r_VTI | bilateral | 125100\20354-7\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | HR | meas | m_Dist_GSV_Thigh | r_HeartRate | bilateral | 125100\8867-4\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Thigh | Dist | meas | m_Dist_GSV_Thigh | r_Distance | bilateral | 125100\G-0364\G-A119\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | ? | meas | m_Mid_GSV_Thigh | r_DopAngle | bilateral | 125100\C12122-01\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | EDV | meas | m_Mid_GSV_Thigh | r_EDV | bilateral | 125100\11653-3\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | MDV | meas | m_Mid_GSV_Thigh | r_MDV | bilateral | 125100\11665-7\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | TAPV | meas | m_Mid_GSV_Thigh | r_TAPV | bilateral | 125100\11692-1\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | PSV | meas | m_Mid_GSV_Thigh | r_PSV | bilateral | 125100\11726-7\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | PI | meas | m_Mid_GSV_Thigh | r_PI | bilateral | 125100\12008-9\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | RI | meas | m_Mid_GSV_Thigh | r_RI | bilateral | 125100\12023-8\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | S/D | meas | m_Mid_GSV_Thigh | r_SD | bilateral | 125100\12144-2\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | AI | meas | m_Mid_GSV_Thigh | r_AccelIndex | bilateral | 125100\20167-3\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | AT | meas | m_Mid_GSV_Thigh | r_AccelTime | bilateral | 125100\20168-1\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | DT | meas | m_Mid_GSV_Thigh | r_DecelTime | bilateral | 125100\20217-6\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | PG | meas | m_Mid_GSV_Thigh | r_PG | bilateral | 125100\20247-3\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | MG | meas | m_Mid_GSV_Thigh | r_PGmean | bilateral | 125100\20256-4\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | TAMV | meas | m_Mid_GSV_Thigh | r_TAMV | bilateral | 125100\20352-1\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | VTI | meas | m_Mid_GSV_Thigh | r_VTI | bilateral | 125100\20354-7\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | HR | meas | m_Mid_GSV_Thigh | r_HeartRate | bilateral | 125100\8867-4\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Thigh | Dist | meas | m_Mid_GSV_Thigh | r_Distance | bilateral | 125100\G-0364\G-A188\R-10259\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, R-10259, Great Saphenous Vein of Thigh |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | ? | meas | m_Prox_GSV_Calf | r_DopAngle | bilateral | 125100\C12122-01\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | EDV | meas | m_Prox_GSV_Calf | r_EDV | bilateral | 125100\11653-3\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | MDV | meas | m_Prox_GSV_Calf | r_MDV | bilateral | 125100\11665-7\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | TAPV | meas | m_Prox_GSV_Calf | r_TAPV | bilateral | 125100\11692-1\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | PSV | meas | m_Prox_GSV_Calf | r_PSV | bilateral | 125100\11726-7\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | PI | meas | m_Prox_GSV_Calf | r_PI | bilateral | 125100\12008-9\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | RI | meas | m_Prox_GSV_Calf | r_RI | bilateral | 125100\12023-8\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | S/D | meas | m_Prox_GSV_Calf | r_SD | bilateral | 125100\12144-2\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | AI | meas | m_Prox_GSV_Calf | r_AccelIndex | bilateral | 125100\20167-3\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | AT | meas | m_Prox_GSV_Calf | r_AccelTime | bilateral | 125100\20168-1\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | DT | meas | m_Prox_GSV_Calf | r_DecelTime | bilateral | 125100\20217-6\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | PG | meas | m_Prox_GSV_Calf | r_PG | bilateral | 125100\20247-3\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | MG | meas | m_Prox_GSV_Calf | r_PGmean | bilateral | 125100\20256-4\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | TAMV | meas | m_Prox_GSV_Calf | r_TAMV | bilateral | 125100\20352-1\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | VTI | meas | m_Prox_GSV_Calf | r_VTI | bilateral | 125100\20354-7\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | HR | meas | m_Prox_GSV_Calf | r_HeartRate | bilateral | 125100\8867-4\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox GSV Calf | Dist | meas | m_Prox_GSV_Calf | r_Distance | bilateral | 125100\G-0364\G-A118\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | ? | meas | m_Dist_GSV_Calf | r_DopAngle | bilateral | 125100\C12122-01\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | EDV | meas | m_Dist_GSV_Calf | r_EDV | bilateral | 125100\11653-3\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | MDV | meas | m_Dist_GSV_Calf | r_MDV | bilateral | 125100\11665-7\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | TAPV | meas | m_Dist_GSV_Calf | r_TAPV | bilateral | 125100\11692-1\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | PSV | meas | m_Dist_GSV_Calf | r_PSV | bilateral | 125100\11726-7\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | PI | meas | m_Dist_GSV_Calf | r_PI | bilateral | 125100\12008-9\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | RI | meas | m_Dist_GSV_Calf | r_RI | bilateral | 125100\12023-8\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | S/D | meas | m_Dist_GSV_Calf | r_SD | bilateral | 125100\12144-2\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | AI | meas | m_Dist_GSV_Calf | r_AccelIndex | bilateral | 125100\20167-3\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | AT | meas | m_Dist_GSV_Calf | r_AccelTime | bilateral | 125100\20168-1\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | DT | meas | m_Dist_GSV_Calf | r_DecelTime | bilateral | 125100\20217-6\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | PG | meas | m_Dist_GSV_Calf | r_PG | bilateral | 125100\20247-3\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | MG | meas | m_Dist_GSV_Calf | r_PGmean | bilateral | 125100\20256-4\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | TAMV | meas | m_Dist_GSV_Calf | r_TAMV | bilateral | 125100\20352-1\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | VTI | meas | m_Dist_GSV_Calf | r_VTI | bilateral | 125100\20354-7\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | HR | meas | m_Dist_GSV_Calf | r_HeartRate | bilateral | 125100\8867-4\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist GSV Calf | Dist | meas | m_Dist_GSV_Calf | r_Distance | bilateral | 125100\G-0364\G-A119\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | ? | meas | m_Mid_GSV_Calf | r_DopAngle | bilateral | 125100\C12122-01\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | EDV | meas | m_Mid_GSV_Calf | r_EDV | bilateral | 125100\11653-3\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | MDV | meas | m_Mid_GSV_Calf | r_MDV | bilateral | 125100\11665-7\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | TAPV | meas | m_Mid_GSV_Calf | r_TAPV | bilateral | 125100\11692-1\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | PSV | meas | m_Mid_GSV_Calf | r_PSV | bilateral | 125100\11726-7\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | PI | meas | m_Mid_GSV_Calf | r_PI | bilateral | 125100\12008-9\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | RI | meas | m_Mid_GSV_Calf | r_RI | bilateral | 125100\12023-8\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | S/D | meas | m_Mid_GSV_Calf | r_SD | bilateral | 125100\12144-2\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | AI | meas | m_Mid_GSV_Calf | r_AccelIndex | bilateral | 125100\20167-3\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | AT | meas | m_Mid_GSV_Calf | r_AccelTime | bilateral | 125100\20168-1\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | DT | meas | m_Mid_GSV_Calf | r_DecelTime | bilateral | 125100\20217-6\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | PG | meas | m_Mid_GSV_Calf | r_PG | bilateral | 125100\20247-3\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | MG | meas | m_Mid_GSV_Calf | r_PGmean | bilateral | 125100\20256-4\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | TAMV | meas | m_Mid_GSV_Calf | r_TAMV | bilateral | 125100\20352-1\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | VTI | meas | m_Mid_GSV_Calf | r_VTI | bilateral | 125100\20354-7\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | HR | meas | m_Mid_GSV_Calf | r_HeartRate | bilateral | 125100\8867-4\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid GSV Calf | Dist | meas | m_Mid_GSV_Calf | r_Distance | bilateral | 125100\G-0364\G-A188\R-1025A\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, R-1025A, Great Saphenous Vein of Calf |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | ? | meas | m_Com_Iliac_V | r_DopAngle | bilateral | 125100\C12122-01\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | EDV | meas | m_Com_Iliac_V | r_EDV | bilateral | 125100\11653-3\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | MDV | meas | m_Com_Iliac_V | r_MDV | bilateral | 125100\11665-7\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | TAPV | meas | m_Com_Iliac_V | r_TAPV | bilateral | 125100\11692-1\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | PSV | meas | m_Com_Iliac_V | r_PSV | bilateral | 125100\11726-7\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | PI | meas | m_Com_Iliac_V | r_PI | bilateral | 125100\12008-9\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | RI | meas | m_Com_Iliac_V | r_RI | bilateral | 125100\12023-8\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | S/D | meas | m_Com_Iliac_V | r_SD | bilateral | 125100\12144-2\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | AI | meas | m_Com_Iliac_V | r_AccelIndex | bilateral | 125100\20167-3\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | AT | meas | m_Com_Iliac_V | r_AccelTime | bilateral | 125100\20168-1\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | DT | meas | m_Com_Iliac_V | r_DecelTime | bilateral | 125100\20217-6\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | PG | meas | m_Com_Iliac_V | r_PG | bilateral | 125100\20247-3\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | MG | meas | m_Com_Iliac_V | r_PGmean | bilateral | 125100\20256-4\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | TAMV | meas | m_Com_Iliac_V | r_TAMV | bilateral | 125100\20352-1\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | VTI | meas | m_Com_Iliac_V | r_VTI | bilateral | 125100\20354-7\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | HR | meas | m_Com_Iliac_V | r_HeartRate | bilateral | 125100\8867-4\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Com Iliac V | Dist | meas | m_Com_Iliac_V | r_Distance | bilateral | 125100\G-0364\T-48920\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48920, Common Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | ? | meas | m_Ext_Iliac_V | r_DopAngle | bilateral | 125100\C12122-01\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | EDV | meas | m_Ext_Iliac_V | r_EDV | bilateral | 125100\11653-3\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | MDV | meas | m_Ext_Iliac_V | r_MDV | bilateral | 125100\11665-7\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | TAPV | meas | m_Ext_Iliac_V | r_TAPV | bilateral | 125100\11692-1\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | PSV | meas | m_Ext_Iliac_V | r_PSV | bilateral | 125100\11726-7\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | PI | meas | m_Ext_Iliac_V | r_PI | bilateral | 125100\12008-9\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | RI | meas | m_Ext_Iliac_V | r_RI | bilateral | 125100\12023-8\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | S/D | meas | m_Ext_Iliac_V | r_SD | bilateral | 125100\12144-2\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | AI | meas | m_Ext_Iliac_V | r_AccelIndex | bilateral | 125100\20167-3\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | AT | meas | m_Ext_Iliac_V | r_AccelTime | bilateral | 125100\20168-1\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | DT | meas | m_Ext_Iliac_V | r_DecelTime | bilateral | 125100\20217-6\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | PG | meas | m_Ext_Iliac_V | r_PG | bilateral | 125100\20247-3\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | MG | meas | m_Ext_Iliac_V | r_PGmean | bilateral | 125100\20256-4\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | TAMV | meas | m_Ext_Iliac_V | r_TAMV | bilateral | 125100\20352-1\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | VTI | meas | m_Ext_Iliac_V | r_VTI | bilateral | 125100\20354-7\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | HR | meas | m_Ext_Iliac_V | r_HeartRate | bilateral | 125100\8867-4\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac V | Dist | meas | m_Ext_Iliac_V | r_Distance | bilateral | 125100\G-0364\T-48930\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48930, External Iliac Vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | ? | meas | m_Int_Iliac_V | r_DopAngle | bilateral | 125100\C12122-01\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | EDV | meas | m_Int_Iliac_V | r_EDV | bilateral | 125100\11653-3\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | MDV | meas | m_Int_Iliac_V | r_MDV | bilateral | 125100\11665-7\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | TAPV | meas | m_Int_Iliac_V | r_TAPV | bilateral | 125100\11692-1\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | PSV | meas | m_Int_Iliac_V | r_PSV | bilateral | 125100\11726-7\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | PI | meas | m_Int_Iliac_V | r_PI | bilateral | 125100\12008-9\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | RI | meas | m_Int_Iliac_V | r_RI | bilateral | 125100\12023-8\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | S/D | meas | m_Int_Iliac_V | r_SD | bilateral | 125100\12144-2\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | AI | meas | m_Int_Iliac_V | r_AccelIndex | bilateral | 125100\20167-3\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | AT | meas | m_Int_Iliac_V | r_AccelTime | bilateral | 125100\20168-1\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | DT | meas | m_Int_Iliac_V | r_DecelTime | bilateral | 125100\20217-6\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | PG | meas | m_Int_Iliac_V | r_PG | bilateral | 125100\20247-3\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | MG | meas | m_Int_Iliac_V | r_PGmean | bilateral | 125100\20256-4\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | TAMV | meas | m_Int_Iliac_V | r_TAMV | bilateral | 125100\20352-1\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | VTI | meas | m_Int_Iliac_V | r_VTI | bilateral | 125100\20354-7\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | HR | meas | m_Int_Iliac_V | r_HeartRate | bilateral | 125100\8867-4\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Int Iliac V | Dist | meas | m_Int_Iliac_V | r_Distance | bilateral | 125100\G-0364\T-48940\T-49403 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48940, Internal iliac vein
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | ? | meas | m_Lat_Sural_V | r_DopAngle | bilateral | 125100\C12122-01\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | EDV | meas | m_Lat_Sural_V | r_EDV | bilateral | 125100\11653-3\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | MDV | meas | m_Lat_Sural_V | r_MDV | bilateral | 125100\11665-7\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | TAPV | meas | m_Lat_Sural_V | r_TAPV | bilateral | 125100\11692-1\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | PSV | meas | m_Lat_Sural_V | r_PSV | bilateral | 125100\11726-7\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | PI | meas | m_Lat_Sural_V | r_PI | bilateral | 125100\12008-9\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | RI | meas | m_Lat_Sural_V | r_RI | bilateral | 125100\12023-8\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | S/D | meas | m_Lat_Sural_V | r_SD | bilateral | 125100\12144-2\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | AI | meas | m_Lat_Sural_V | r_AccelIndex | bilateral | 125100\20167-3\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | AT | meas | m_Lat_Sural_V | r_AccelTime | bilateral | 125100\20168-1\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | DT | meas | m_Lat_Sural_V | r_DecelTime | bilateral | 125100\20217-6\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | PG | meas | m_Lat_Sural_V | r_PG | bilateral | 125100\20247-3\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | MG | meas | m_Lat_Sural_V | r_PGmean | bilateral | 125100\20256-4\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | TAMV | meas | m_Lat_Sural_V | r_TAMV | bilateral | 125100\20352-1\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | VTI | meas | m_Lat_Sural_V | r_VTI | bilateral | 125100\20354-7\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | HR | meas | m_Lat_Sural_V | r_HeartRate | bilateral | 125100\8867-4\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Lat Sural V | Dist | meas | m_Lat_Sural_V | r_Distance | bilateral | 125100\G-0364\G-A104\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A104, Lateral
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | ? | meas | m_Med_Sural_V | r_DopAngle | bilateral | 125100\C12122-01\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | EDV | meas | m_Med_Sural_V | r_EDV | bilateral | 125100\11653-3\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | MDV | meas | m_Med_Sural_V | r_MDV | bilateral | 125100\11665-7\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | TAPV | meas | m_Med_Sural_V | r_TAPV | bilateral | 125100\11692-1\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | PSV | meas | m_Med_Sural_V | r_PSV | bilateral | 125100\11726-7\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | PI | meas | m_Med_Sural_V | r_PI | bilateral | 125100\12008-9\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | RI | meas | m_Med_Sural_V | r_RI | bilateral | 125100\12023-8\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | S/D | meas | m_Med_Sural_V | r_SD | bilateral | 125100\12144-2\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | AI | meas | m_Med_Sural_V | r_AccelIndex | bilateral | 125100\20167-3\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | AT | meas | m_Med_Sural_V | r_AccelTime | bilateral | 125100\20168-1\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | DT | meas | m_Med_Sural_V | r_DecelTime | bilateral | 125100\20217-6\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | PG | meas | m_Med_Sural_V | r_PG | bilateral | 125100\20247-3\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | MG | meas | m_Med_Sural_V | r_PGmean | bilateral | 125100\20256-4\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | TAMV | meas | m_Med_Sural_V | r_TAMV | bilateral | 125100\20352-1\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | VTI | meas | m_Med_Sural_V | r_VTI | bilateral | 125100\20354-7\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | HR | meas | m_Med_Sural_V | r_HeartRate | bilateral | 125100\8867-4\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Med Sural V | Dist | meas | m_Med_Sural_V | r_Distance | bilateral | 125100\G-0364\G-A109\T-49403\T-4942D | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-4942D, Gastrocnemius vein | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | ? | meas | m_Prox_LSV | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | EDV | meas | m_Prox_LSV | r_EDV | bilateral | 125100\11653-3\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | MDV | meas | m_Prox_LSV | r_MDV | bilateral | 125100\11665-7\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | TAPV | meas | m_Prox_LSV | r_TAPV | bilateral | 125100\11692-1\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | PSV | meas | m_Prox_LSV | r_PSV | bilateral | 125100\11726-7\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | PI | meas | m_Prox_LSV | r_PI | bilateral | 125100\12008-9\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | RI | meas | m_Prox_LSV | r_RI | bilateral | 125100\12023-8\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | S/D | meas | m_Prox_LSV | r_SD | bilateral | 125100\12144-2\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | AI | meas | m_Prox_LSV | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | AT | meas | m_Prox_LSV | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | DT | meas | m_Prox_LSV | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | PG | meas | m_Prox_LSV | r_PG | bilateral | 125100\20247-3\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | MG | meas | m_Prox_LSV | r_PGmean | bilateral | 125100\20256-4\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | TAMV | meas | m_Prox_LSV | r_TAMV | bilateral | 125100\20352-1\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | VTI | meas | m_Prox_LSV | r_VTI | bilateral | 125100\20354-7\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | HR | meas | m_Prox_LSV | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox LSV | Dist | meas | m_Prox_LSV | r_Distance | bilateral | 125100\G-0364\G-A118\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | ? | meas | m_Dist_LSV | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | EDV | meas | m_Dist_LSV | r_EDV | bilateral | 125100\11653-3\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | MDV | meas | m_Dist_LSV | r_MDV | bilateral | 125100\11665-7\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | TAPV | meas | m_Dist_LSV | r_TAPV | bilateral | 125100\11692-1\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | PSV | meas | m_Dist_LSV | r_PSV | bilateral | 125100\11726-7\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | PI | meas | m_Dist_LSV | r_PI | bilateral | 125100\12008-9\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | RI | meas | m_Dist_LSV | r_RI | bilateral | 125100\12023-8\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | S/D | meas | m_Dist_LSV | r_SD | bilateral | 125100\12144-2\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | AI | meas | m_Dist_LSV | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | AT | meas | m_Dist_LSV | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | DT | meas | m_Dist_LSV | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | PG | meas | m_Dist_LSV | r_PG | bilateral | 125100\20247-3\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | MG | meas | m_Dist_LSV | r_PGmean | bilateral | 125100\20256-4\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | TAMV | meas | m_Dist_LSV | r_TAMV | bilateral | 125100\20352-1\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | VTI | meas | m_Dist_LSV | r_VTI | bilateral | 125100\20354-7\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | HR | meas | m_Dist_LSV | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist LSV | Dist | meas | m_Dist_LSV | r_Distance | bilateral | 125100\G-0364\G-A119\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | ? | meas | m_Mid_LSV | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | EDV | meas | m_Mid_LSV | r_EDV | bilateral | 125100\11653-3\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | MDV | meas | m_Mid_LSV | r_MDV | bilateral | 125100\11665-7\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | TAPV | meas | m_Mid_LSV | r_TAPV | bilateral | 125100\11692-1\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | PSV | meas | m_Mid_LSV | r_PSV | bilateral | 125100\11726-7\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | PI | meas | m_Mid_LSV | r_PI | bilateral | 125100\12008-9\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | RI | meas | m_Mid_LSV | r_RI | bilateral | 125100\12023-8\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | S/D | meas | m_Mid_LSV | r_SD | bilateral | 125100\12144-2\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | AI | meas | m_Mid_LSV | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | AT | meas | m_Mid_LSV | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | DT | meas | m_Mid_LSV | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | PG | meas | m_Mid_LSV | r_PG | bilateral | 125100\20247-3\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | MG | meas | m_Mid_LSV | r_PGmean | bilateral | 125100\20256-4\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | TAMV | meas | m_Mid_LSV | r_TAMV | bilateral | 125100\20352-1\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | VTI | meas | m_Mid_LSV | r_VTI | bilateral | 125100\20354-7\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | HR | meas | m_Mid_LSV | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid LSV | Dist | meas | m_Mid_LSV | r_Distance | bilateral | 125100\G-0364\G-A188\T-49403\T-49550 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49550, Lesser Saphenous Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | ? | meas | m_Prox_PTV | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | EDV | meas | m_Prox_PTV | r_EDV | bilateral | 125100\11653-3\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | MDV | meas | m_Prox_PTV | r_MDV | bilateral | 125100\11665-7\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | TAPV | meas | m_Prox_PTV | r_TAPV | bilateral | 125100\11692-1\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | PSV | meas | m_Prox_PTV | r_PSV | bilateral | 125100\11726-7\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | PI | meas | m_Prox_PTV | r_PI | bilateral | 125100\12008-9\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | RI | meas | m_Prox_PTV | r_RI | bilateral | 125100\12023-8\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | S/D | meas | m_Prox_PTV | r_SD | bilateral | 125100\12144-2\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | AI | meas | m_Prox_PTV | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | AT | meas | m_Prox_PTV | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | DT | meas | m_Prox_PTV | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | PG | meas | m_Prox_PTV | r_PG | bilateral | 125100\20247-3\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | MG | meas | m_Prox_PTV | r_PGmean | bilateral | 125100\20256-4\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | TAMV | meas | m_Prox_PTV | r_TAMV | bilateral | 125100\20352-1\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | VTI | meas | m_Prox_PTV | r_VTI | bilateral | 125100\20354-7\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | HR | meas | m_Prox_PTV | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTV | Dist | meas | m_Prox_PTV | r_Distance | bilateral | 125100\G-0364\G-A118\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | ? | meas | m_Dist_PTV | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | EDV | meas | m_Dist_PTV | r_EDV | bilateral | 125100\11653-3\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | MDV | meas | m_Dist_PTV | r_MDV | bilateral | 125100\11665-7\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | TAPV | meas | m_Dist_PTV | r_TAPV | bilateral | 125100\11692-1\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | PSV | meas | m_Dist_PTV | r_PSV | bilateral | 125100\11726-7\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | PI | meas | m_Dist_PTV | r_PI | bilateral | 125100\12008-9\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | RI | meas | m_Dist_PTV | r_RI | bilateral | 125100\12023-8\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | S/D | meas | m_Dist_PTV | r_SD | bilateral | 125100\12144-2\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | AI | meas | m_Dist_PTV | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | AT | meas | m_Dist_PTV | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | DT | meas | m_Dist_PTV | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | PG | meas | m_Dist_PTV | r_PG | bilateral | 125100\20247-3\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | MG | meas | m_Dist_PTV | r_PGmean | bilateral | 125100\20256-4\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | TAMV | meas | m_Dist_PTV | r_TAMV | bilateral | 125100\20352-1\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | VTI | meas | m_Dist_PTV | r_VTI | bilateral | 125100\20354-7\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | HR | meas | m_Dist_PTV | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTV | Dist | meas | m_Dist_PTV | r_Distance | bilateral | 125100\G-0364\G-A119\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | ? | meas | m_Mid_PTV | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | EDV | meas | m_Mid_PTV | r_EDV | bilateral | 125100\11653-3\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | MDV | meas | m_Mid_PTV | r_MDV | bilateral | 125100\11665-7\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | TAPV | meas | m_Mid_PTV | r_TAPV | bilateral | 125100\11692-1\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | PSV | meas | m_Mid_PTV | r_PSV | bilateral | 125100\11726-7\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | PI | meas | m_Mid_PTV | r_PI | bilateral | 125100\12008-9\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | RI | meas | m_Mid_PTV | r_RI | bilateral | 125100\12023-8\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | S/D | meas | m_Mid_PTV | r_SD | bilateral | 125100\12144-2\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | AI | meas | m_Mid_PTV | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | AT | meas | m_Mid_PTV | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | DT | meas | m_Mid_PTV | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | PG | meas | m_Mid_PTV | r_PG | bilateral | 125100\20247-3\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | MG | meas | m_Mid_PTV | r_PGmean | bilateral | 125100\20256-4\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | TAMV | meas | m_Mid_PTV | r_TAMV | bilateral | 125100\20352-1\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | VTI | meas | m_Mid_PTV | r_VTI | bilateral | 125100\20354-7\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | HR | meas | m_Mid_PTV | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTV | Dist | meas | m_Mid_PTV | r_Distance | bilateral | 125100\G-0364\G-A188\T-49403\T-49620 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49620, Posterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | ? | meas | m_Prox_ATV | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | EDV | meas | m_Prox_ATV | r_EDV | bilateral | 125100\11653-3\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | MDV | meas | m_Prox_ATV | r_MDV | bilateral | 125100\11665-7\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | TAPV | meas | m_Prox_ATV | r_TAPV | bilateral | 125100\11692-1\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | PSV | meas | m_Prox_ATV | r_PSV | bilateral | 125100\11726-7\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | PI | meas | m_Prox_ATV | r_PI | bilateral | 125100\12008-9\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | RI | meas | m_Prox_ATV | r_RI | bilateral | 125100\12023-8\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | S/D | meas | m_Prox_ATV | r_SD | bilateral | 125100\12144-2\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | AI | meas | m_Prox_ATV | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | AT | meas | m_Prox_ATV | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | DT | meas | m_Prox_ATV | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | PG | meas | m_Prox_ATV | r_PG | bilateral | 125100\20247-3\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | MG | meas | m_Prox_ATV | r_PGmean | bilateral | 125100\20256-4\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | TAMV | meas | m_Prox_ATV | r_TAMV | bilateral | 125100\20352-1\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | VTI | meas | m_Prox_ATV | r_VTI | bilateral | 125100\20354-7\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | HR | meas | m_Prox_ATV | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATV | Dist | meas | m_Prox_ATV | r_Distance | bilateral | 125100\G-0364\G-A118\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | ? | meas | m_Dist_ATV | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | EDV | meas | m_Dist_ATV | r_EDV | bilateral | 125100\11653-3\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | MDV | meas | m_Dist_ATV | r_MDV | bilateral | 125100\11665-7\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | TAPV | meas | m_Dist_ATV | r_TAPV | bilateral | 125100\11692-1\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | PSV | meas | m_Dist_ATV | r_PSV | bilateral | 125100\11726-7\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | PI | meas | m_Dist_ATV | r_PI | bilateral | 125100\12008-9\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | RI | meas | m_Dist_ATV | r_RI | bilateral | 125100\12023-8\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | S/D | meas | m_Dist_ATV | r_SD | bilateral | 125100\12144-2\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | AI | meas | m_Dist_ATV | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | AT | meas | m_Dist_ATV | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | DT | meas | m_Dist_ATV | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | PG | meas | m_Dist_ATV | r_PG | bilateral | 125100\20247-3\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | MG | meas | m_Dist_ATV | r_PGmean | bilateral | 125100\20256-4\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | TAMV | meas | m_Dist_ATV | r_TAMV | bilateral | 125100\20352-1\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | VTI | meas | m_Dist_ATV | r_VTI | bilateral | 125100\20354-7\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | HR | meas | m_Dist_ATV | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATV | Dist | meas | m_Dist_ATV | r_Distance | bilateral | 125100\G-0364\G-A119\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | ? | meas | m_Mid_ATV | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | EDV | meas | m_Mid_ATV | r_EDV | bilateral | 125100\11653-3\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | MDV | meas | m_Mid_ATV | r_MDV | bilateral | 125100\11665-7\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | TAPV | meas | m_Mid_ATV | r_TAPV | bilateral | 125100\11692-1\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | PSV | meas | m_Mid_ATV | r_PSV | bilateral | 125100\11726-7\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | PI | meas | m_Mid_ATV | r_PI | bilateral | 125100\12008-9\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | RI | meas | m_Mid_ATV | r_RI | bilateral | 125100\12023-8\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | S/D | meas | m_Mid_ATV | r_SD | bilateral | 125100\12144-2\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | AI | meas | m_Mid_ATV | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | AT | meas | m_Mid_ATV | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | DT | meas | m_Mid_ATV | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | PG | meas | m_Mid_ATV | r_PG | bilateral | 125100\20247-3\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | MG | meas | m_Mid_ATV | r_PGmean | bilateral | 125100\20256-4\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | TAMV | meas | m_Mid_ATV | r_TAMV | bilateral | 125100\20352-1\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | VTI | meas | m_Mid_ATV | r_VTI | bilateral | 125100\20354-7\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | HR | meas | m_Mid_ATV | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATV | Dist | meas | m_Mid_ATV | r_Distance | bilateral | 125100\G-0364\G-A188\T-49403\T-49630 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49630, Anterior Tibial Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | ? | meas | m_Prox_POPV | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | EDV | meas | m_Prox_POPV | r_EDV | bilateral | 125100\11653-3\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | MDV | meas | m_Prox_POPV | r_MDV | bilateral | 125100\11665-7\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | TAPV | meas | m_Prox_POPV | r_TAPV | bilateral | 125100\11692-1\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | PSV | meas | m_Prox_POPV | r_PSV | bilateral | 125100\11726-7\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | PI | meas | m_Prox_POPV | r_PI | bilateral | 125100\12008-9\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | RI | meas | m_Prox_POPV | r_RI | bilateral | 125100\12023-8\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | S/D | meas | m_Prox_POPV | r_SD | bilateral | 125100\12144-2\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | AI | meas | m_Prox_POPV | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | AT | meas | m_Prox_POPV | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | DT | meas | m_Prox_POPV | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | PG | meas | m_Prox_POPV | r_PG | bilateral | 125100\20247-3\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | MG | meas | m_Prox_POPV | r_PGmean | bilateral | 125100\20256-4\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | TAMV | meas | m_Prox_POPV | r_TAMV | bilateral | 125100\20352-1\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | VTI | meas | m_Prox_POPV | r_VTI | bilateral | 125100\20354-7\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | HR | meas | m_Prox_POPV | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop V | Dist | meas | m_Prox_POPV | r_Distance | bilateral | 125100\G-0364\G-A118\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | ? | meas | m_Dist_POPV | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | EDV | meas | m_Dist_POPV | r_EDV | bilateral | 125100\11653-3\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | MDV | meas | m_Dist_POPV | r_MDV | bilateral | 125100\11665-7\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | TAPV | meas | m_Dist_POPV | r_TAPV | bilateral | 125100\11692-1\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | PSV | meas | m_Dist_POPV | r_PSV | bilateral | 125100\11726-7\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | PI | meas | m_Dist_POPV | r_PI | bilateral | 125100\12008-9\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | RI | meas | m_Dist_POPV | r_RI | bilateral | 125100\12023-8\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | S/D | meas | m_Dist_POPV | r_SD | bilateral | 125100\12144-2\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | AI | meas | m_Dist_POPV | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | AT | meas | m_Dist_POPV | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | DT | meas | m_Dist_POPV | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | PG | meas | m_Dist_POPV | r_PG | bilateral | 125100\20247-3\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | MG | meas | m_Dist_POPV | r_PGmean | bilateral | 125100\20256-4\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | TAMV | meas | m_Dist_POPV | r_TAMV | bilateral | 125100\20352-1\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | VTI | meas | m_Dist_POPV | r_VTI | bilateral | 125100\20354-7\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | HR | meas | m_Dist_POPV | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop V | Dist | meas | m_Dist_POPV | r_Distance | bilateral | 125100\G-0364\G-A119\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | ? | meas | m_Mid_POPV | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | EDV | meas | m_Mid_POPV | r_EDV | bilateral | 125100\11653-3\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | MDV | meas | m_Mid_POPV | r_MDV | bilateral | 125100\11665-7\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | TAPV | meas | m_Mid_POPV | r_TAPV | bilateral | 125100\11692-1\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | PSV | meas | m_Mid_POPV | r_PSV | bilateral | 125100\11726-7\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | PI | meas | m_Mid_POPV | r_PI | bilateral | 125100\12008-9\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | RI | meas | m_Mid_POPV | r_RI | bilateral | 125100\12023-8\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | S/D | meas | m_Mid_POPV | r_SD | bilateral | 125100\12144-2\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | AI | meas | m_Mid_POPV | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | AT | meas | m_Mid_POPV | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | DT | meas | m_Mid_POPV | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | PG | meas | m_Mid_POPV | r_PG | bilateral | 125100\20247-3\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | MG | meas | m_Mid_POPV | r_PGmean | bilateral | 125100\20256-4\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | TAMV | meas | m_Mid_POPV | r_TAMV | bilateral | 125100\20352-1\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | VTI | meas | m_Mid_POPV | r_VTI | bilateral | 125100\20354-7\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | HR | meas | m_Mid_POPV | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pop V | Dist | meas | m_Mid_POPV | r_Distance | bilateral | 125100\G-0364\G-A188\T-49403\T-49640 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49640, Popliteal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | ? | meas | m_Prox_Pero_V | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | EDV | meas | m_Prox_Pero_V | r_EDV | bilateral | 125100\11653-3\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | MDV | meas | m_Prox_Pero_V | r_MDV | bilateral | 125100\11665-7\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | TAPV | meas | m_Prox_Pero_V | r_TAPV | bilateral | 125100\11692-1\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | PSV | meas | m_Prox_Pero_V | r_PSV | bilateral | 125100\11726-7\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | PI | meas | m_Prox_Pero_V | r_PI | bilateral | 125100\12008-9\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | RI | meas | m_Prox_Pero_V | r_RI | bilateral | 125100\12023-8\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | S/D | meas | m_Prox_Pero_V | r_SD | bilateral | 125100\12144-2\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | AI | meas | m_Prox_Pero_V | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | AT | meas | m_Prox_Pero_V | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | DT | meas | m_Prox_Pero_V | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | PG | meas | m_Prox_Pero_V | r_PG | bilateral | 125100\20247-3\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | MG | meas | m_Prox_Pero_V | r_PGmean | bilateral | 125100\20256-4\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | TAMV | meas | m_Prox_Pero_V | r_TAMV | bilateral | 125100\20352-1\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | VTI | meas | m_Prox_Pero_V | r_VTI | bilateral | 125100\20354-7\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | HR | meas | m_Prox_Pero_V | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero V | Dist | meas | m_Prox_Pero_V | r_Distance | bilateral | 125100\G-0364\G-A118\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | ? | meas | m_Dist_Pero_V | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | EDV | meas | m_Dist_Pero_V | r_EDV | bilateral | 125100\11653-3\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | MDV | meas | m_Dist_Pero_V | r_MDV | bilateral | 125100\11665-7\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | TAPV | meas | m_Dist_Pero_V | r_TAPV | bilateral | 125100\11692-1\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | PSV | meas | m_Dist_Pero_V | r_PSV | bilateral | 125100\11726-7\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | PI | meas | m_Dist_Pero_V | r_PI | bilateral | 125100\12008-9\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | RI | meas | m_Dist_Pero_V | r_RI | bilateral | 125100\12023-8\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | S/D | meas | m_Dist_Pero_V | r_SD | bilateral | 125100\12144-2\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | AI | meas | m_Dist_Pero_V | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | AT | meas | m_Dist_Pero_V | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | DT | meas | m_Dist_Pero_V | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | PG | meas | m_Dist_Pero_V | r_PG | bilateral | 125100\20247-3\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | MG | meas | m_Dist_Pero_V | r_PGmean | bilateral | 125100\20256-4\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | TAMV | meas | m_Dist_Pero_V | r_TAMV | bilateral | 125100\20352-1\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | VTI | meas | m_Dist_Pero_V | r_VTI | bilateral | 125100\20354-7\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | HR | meas | m_Dist_Pero_V | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero V | Dist | meas | m_Dist_Pero_V | r_Distance | bilateral | 125100\G-0364\G-A119\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | ? | meas | m_Mid_Pero_V | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | EDV | meas | m_Mid_Pero_V | r_EDV | bilateral | 125100\11653-3\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | MDV | meas | m_Mid_Pero_V | r_MDV | bilateral | 125100\11665-7\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | TAPV | meas | m_Mid_Pero_V | r_TAPV | bilateral | 125100\11692-1\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | PSV | meas | m_Mid_Pero_V | r_PSV | bilateral | 125100\11726-7\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | PI | meas | m_Mid_Pero_V | r_PI | bilateral | 125100\12008-9\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | RI | meas | m_Mid_Pero_V | r_RI | bilateral | 125100\12023-8\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | S/D | meas | m_Mid_Pero_V | r_SD | bilateral | 125100\12144-2\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | AI | meas | m_Mid_Pero_V | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | AT | meas | m_Mid_Pero_V | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | DT | meas | m_Mid_Pero_V | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | PG | meas | m_Mid_Pero_V | r_PG | bilateral | 125100\20247-3\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | MG | meas | m_Mid_Pero_V | r_PGmean | bilateral | 125100\20256-4\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | TAMV | meas | m_Mid_Pero_V | r_TAMV | bilateral | 125100\20352-1\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | VTI | meas | m_Mid_Pero_V | r_VTI | bilateral | 125100\20354-7\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | HR | meas | m_Mid_Pero_V | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero V | Dist | meas | m_Mid_Pero_V | r_Distance | bilateral | 125100\G-0364\G-A188\T-49403\T-49650 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49650, Peroneal Vein |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | ? | meas | m_Prox_DPF | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | EDV | meas | m_Prox_DPF | r_EDV | bilateral | 125100\11653-3\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | MDV | meas | m_Prox_DPF | r_MDV | bilateral | 125100\11665-7\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | TAPV | meas | m_Prox_DPF | r_TAPV | bilateral | 125100\11692-1\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | PSV | meas | m_Prox_DPF | r_PSV | bilateral | 125100\11726-7\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | PI | meas | m_Prox_DPF | r_PI | bilateral | 125100\12008-9\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | RI | meas | m_Prox_DPF | r_RI | bilateral | 125100\12023-8\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | S/D | meas | m_Prox_DPF | r_SD | bilateral | 125100\12144-2\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | AI | meas | m_Prox_DPF | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | AT | meas | m_Prox_DPF | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | DT | meas | m_Prox_DPF | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | PG | meas | m_Prox_DPF | r_PG | bilateral | 125100\20247-3\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | MG | meas | m_Prox_DPF | r_PGmean | bilateral | 125100\20256-4\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | TAMV | meas | m_Prox_DPF | r_TAMV | bilateral | 125100\20352-1\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | VTI | meas | m_Prox_DPF | r_VTI | bilateral | 125100\20354-7\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | HR | meas | m_Prox_DPF | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox DPF V | Dist | meas | m_Prox_DPF | r_Distance | bilateral | 125100\G-0364\G-A118\T-49403\T-49660 | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-49660, Profunda Femoris Vein |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | SFJ | ? | meas | m_SFJ | r_DopAngle | bilateral | 125100\C12122-01\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Vascular | SFJ | EDV | meas | m_SFJ | r_EDV | bilateral | 125100\11653-3\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Vascular | SFJ | MDV | meas | m_SFJ | r_MDV | bilateral | 125100\11665-7\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Vascular | SFJ | TAPV | meas | m_SFJ | r_TAPV | bilateral | 125100\11692-1\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Vascular | SFJ | PSV | meas | m_SFJ | r_PSV | bilateral | 125100\11726-7\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Vascular | SFJ | PI | meas | m_SFJ | r_PI | bilateral | 125100\12008-9\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Vascular | SFJ | RI | meas | m_SFJ | r_RI | bilateral | 125100\12023-8\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Vascular | SFJ | S/D | meas | m_SFJ | r_SD | bilateral | 125100\12144-2\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Vascular | SFJ | AI | meas | m_SFJ | r_AccelIndex | bilateral | 125100\20167-3\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Vascular | SFJ | AT | meas | m_SFJ | r_AccelTime | bilateral | 125100\20168-1\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Vascular | SFJ | DT | meas | m_SFJ | r_DecelTime | bilateral | 125100\20217-6\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Vascular | SFJ | PG | meas | m_SFJ | r_PG | bilateral | 125100\20247-3\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Vascular | SFJ | MG | meas | m_SFJ | r_PGmean | bilateral | 125100\20256-4\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Vascular | SFJ | TAMV | meas | m_SFJ | r_TAMV | bilateral | 125100\20352-1\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Vascular | SFJ | VTI | meas | m_SFJ | r_VTI | bilateral | 125100\20354-7\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Vascular | SFJ | HR | meas | m_SFJ | r_HeartRate | bilateral | 125100\8867-4\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Vascular | SFJ | Dist | meas | m_SFJ | r_Distance | bilateral | 125100\G-0364\T-49403\T-D930A | SRT, T-49403, Vein of Lower Extremity |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-D930A, Saphenofemoral Junction
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | ? | meas | m_hilarart | r_DopAngle | bilateral | 125100\C12122-01\G-035C\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | EDV | meas | m_hilarart | r_EDV | bilateral | 125100\11653-3\G-035C\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | MDV | meas | m_hilarart | r_MDV | bilateral | 125100\11665-7\G-035C\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | TAPV | meas | m_hilarart | r_TAPV | bilateral | 125100\11692-1\G-035C\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | PSV | meas | m_hilarart | r_PSV | bilateral | 125100\11726-7\G-035C\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | PI | meas | m_hilarart | r_PI | bilateral | 125100\12008-9\G-035C\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | RI | meas | m_hilarart | r_RI | bilateral | 125100\12023-8\G-035C\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | S/D | meas | m_hilarart | r_SD | bilateral | 125100\12144-2\G-035C\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | AI | meas | m_hilarart | r_AccelIndex | bilateral | 125100\20167-3\G-035C\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | AT | meas | m_hilarart | r_AccelTime | bilateral | 125100\20168-1\G-035C\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | DT | meas | m_hilarart | r_DecelTime | bilateral | 125100\20217-6\G-035C\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | PG | meas | m_hilarart | r_PG | bilateral | 125100\20247-3\G-035C\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | MG | meas | m_hilarart | r_PGmean | bilateral | 125100\20256-4\G-035C\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | TAMV | meas | m_hilarart | r_TAMV | bilateral | 125100\20352-1\G-035C\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | VTI | meas | m_hilarart | r_VTI | bilateral | 125100\20354-7\G-035C\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | HR | meas | m_hilarart | r_HeartRate | bilateral | 125100\8867-4\G-035C\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Hilar A | Dist | meas | m_hilarart | r_Distance | bilateral | 125100\G-035C\G-0364\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, G-035C, Hilar Artery
Vascular Ultrasound Procedure Report | Abdominal | Renal PSV | ? | meas | m_renal_PSV | r_DopAngle | bilateral | 125100\C12122-01\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46600, Renal Artery
Vascular Ultrasound Procedure Report | Abdominal | Renal PSV | PSV | meas | m_renal_PSV | r_PSV | bilateral | 125100\11726-7\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | ? | meas | m_renaortorg | r_DopAngle | bilateral | 125100\C12122-01\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | EDV | meas | m_renaortorg | r_EDV | bilateral | 125100\11653-3\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | MDV | meas | m_renaortorg | r_MDV | bilateral | 125100\11665-7\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | TAPV | meas | m_renaortorg | r_TAPV | bilateral | 125100\11692-1\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | PSV | meas | m_renaortorg | r_PSV | bilateral | 125100\11726-7\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | PI | meas | m_renaortorg | r_PI | bilateral | 125100\12008-9\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | RI | meas | m_renaortorg | r_RI | bilateral | 125100\12023-8\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | S/D | meas | m_renaortorg | r_SD | bilateral | 125100\12144-2\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | AI | meas | m_renaortorg | r_AccelIndex | bilateral | 125100\20167-3\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | AT | meas | m_renaortorg | r_AccelTime | bilateral | 125100\20168-1\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | DT | meas | m_renaortorg | r_DecelTime | bilateral | 125100\20217-6\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | PG | meas | m_renaortorg | r_PG | bilateral | 125100\20247-3\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | MG | meas | m_renaortorg | r_PGmean | bilateral | 125100\20256-4\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | TAMV | meas | m_renaortorg | r_TAMV | bilateral | 125100\20352-1\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | VTI | meas | m_renaortorg | r_VTI | bilateral | 125100\20354-7\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | HR | meas | m_renaortorg | r_HeartRate | bilateral | 125100\8867-4\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Ren A Org | Dist | meas | m_renaortorg | r_Distance | bilateral | 125100\G-0364\G-036A\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | ? | meas | m_proxrenart | r_DopAngle | bilateral | 125100\C12122-01\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | EDV | meas | m_proxrenart | r_EDV | bilateral | 125100\11653-3\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | MDV | meas | m_proxrenart | r_MDV | bilateral | 125100\11665-7\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | TAPV | meas | m_proxrenart | r_TAPV | bilateral | 125100\11692-1\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | PSV | meas | m_proxrenart | r_PSV | bilateral | 125100\11726-7\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | PI | meas | m_proxrenart | r_PI | bilateral | 125100\12008-9\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | RI | meas | m_proxrenart | r_RI | bilateral | 125100\12023-8\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | S/D | meas | m_proxrenart | r_SD | bilateral | 125100\12144-2\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | AI | meas | m_proxrenart | r_AccelIndex | bilateral | 125100\20167-3\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | AT | meas | m_proxrenart | r_AccelTime | bilateral | 125100\20168-1\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | DT | meas | m_proxrenart | r_DecelTime | bilateral | 125100\20217-6\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | PG | meas | m_proxrenart | r_PG | bilateral | 125100\20247-3\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | MG | meas | m_proxrenart | r_PGmean | bilateral | 125100\20256-4\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | TAMV | meas | m_proxrenart | r_TAMV | bilateral | 125100\20352-1\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | VTI | meas | m_proxrenart | r_VTI | bilateral | 125100\20354-7\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | HR | meas | m_proxrenart | r_HeartRate | bilateral | 125100\8867-4\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Prox Ren A | Dist | meas | m_proxrenart | r_Distance | bilateral | 125100\G-0364\G-A118\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | ? | meas | m_distrenart | r_DopAngle | bilateral | 125100\C12122-01\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | EDV | meas | m_distrenart | r_EDV | bilateral | 125100\11653-3\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | MDV | meas | m_distrenart | r_MDV | bilateral | 125100\11665-7\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | TAPV | meas | m_distrenart | r_TAPV | bilateral | 125100\11692-1\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | PSV | meas | m_distrenart | r_PSV | bilateral | 125100\11726-7\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | PI | meas | m_distrenart | r_PI | bilateral | 125100\12008-9\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | RI | meas | m_distrenart | r_RI | bilateral | 125100\12023-8\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | S/D | meas | m_distrenart | r_SD | bilateral | 125100\12144-2\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | AI | meas | m_distrenart | r_AccelIndex | bilateral | 125100\20167-3\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | AT | meas | m_distrenart | r_AccelTime | bilateral | 125100\20168-1\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | DT | meas | m_distrenart | r_DecelTime | bilateral | 125100\20217-6\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | PG | meas | m_distrenart | r_PG | bilateral | 125100\20247-3\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | MG | meas | m_distrenart | r_PGmean | bilateral | 125100\20256-4\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | TAMV | meas | m_distrenart | r_TAMV | bilateral | 125100\20352-1\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | VTI | meas | m_distrenart | r_VTI | bilateral | 125100\20354-7\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | HR | meas | m_distrenart | r_HeartRate | bilateral | 125100\8867-4\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Dist Ren A | Dist | meas | m_distrenart | r_Distance | bilateral | 125100\G-0364\G-A119\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | ? | meas | m_midrenart | r_DopAngle | bilateral | 125100\C12122-01\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | EDV | meas | m_midrenart | r_EDV | bilateral | 125100\11653-3\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | MDV | meas | m_midrenart | r_MDV | bilateral | 125100\11665-7\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | TAPV | meas | m_midrenart | r_TAPV | bilateral | 125100\11692-1\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | PSV | meas | m_midrenart | r_PSV | bilateral | 125100\11726-7\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | PI | meas | m_midrenart | r_PI | bilateral | 125100\12008-9\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | RI | meas | m_midrenart | r_RI | bilateral | 125100\12023-8\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | S/D | meas | m_midrenart | r_SD | bilateral | 125100\12144-2\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | AI | meas | m_midrenart | r_AccelIndex | bilateral | 125100\20167-3\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | AT | meas | m_midrenart | r_AccelTime | bilateral | 125100\20168-1\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | DT | meas | m_midrenart | r_DecelTime | bilateral | 125100\20217-6\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | PG | meas | m_midrenart | r_PG | bilateral | 125100\20247-3\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | MG | meas | m_midrenart | r_PGmean | bilateral | 125100\20256-4\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | TAMV | meas | m_midrenart | r_TAMV | bilateral | 125100\20352-1\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | VTI | meas | m_midrenart | r_VTI | bilateral | 125100\20354-7\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | HR | meas | m_midrenart | r_HeartRate | bilateral | 125100\8867-4\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Ren A | Dist | meas | m_midrenart | r_Distance | bilateral | 125100\G-0364\G-A188\T-46600\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46600, Renal Artery |  |  |  |  | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | ? | meas | m_medseg | r_DopAngle | bilateral | 125100\C12122-01\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | EDV | meas | m_medseg | r_EDV | bilateral | 125100\11653-3\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | MDV | meas | m_medseg | r_MDV | bilateral | 125100\11665-7\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | TAPV | meas | m_medseg | r_TAPV | bilateral | 125100\11692-1\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | PSV | meas | m_medseg | r_PSV | bilateral | 125100\11726-7\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | PI | meas | m_medseg | r_PI | bilateral | 125100\12008-9\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | RI | meas | m_medseg | r_RI | bilateral | 125100\12023-8\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | S/D | meas | m_medseg | r_SD | bilateral | 125100\12144-2\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | AI | meas | m_medseg | r_AccelIndex | bilateral | 125100\20167-3\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | AT | meas | m_medseg | r_AccelTime | bilateral | 125100\20168-1\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | DT | meas | m_medseg | r_DecelTime | bilateral | 125100\20217-6\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | PG | meas | m_medseg | r_PG | bilateral | 125100\20247-3\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | MG | meas | m_medseg | r_PGmean | bilateral | 125100\20256-4\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | TAMV | meas | m_medseg | r_TAMV | bilateral | 125100\20352-1\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | VTI | meas | m_medseg | r_VTI | bilateral | 125100\20354-7\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | HR | meas | m_medseg | r_HeartRate | bilateral | 125100\8867-4\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Seg | Dist | meas | m_medseg | r_Distance | bilateral | 125100\G-0364\G-A109\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | ? | meas | m_infseg | r_DopAngle | bilateral | 125100\C12122-01\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | EDV | meas | m_infseg | r_EDV | bilateral | 125100\11653-3\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | MDV | meas | m_infseg | r_MDV | bilateral | 125100\11665-7\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | TAPV | meas | m_infseg | r_TAPV | bilateral | 125100\11692-1\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | PSV | meas | m_infseg | r_PSV | bilateral | 125100\11726-7\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | PI | meas | m_infseg | r_PI | bilateral | 125100\12008-9\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | RI | meas | m_infseg | r_RI | bilateral | 125100\12023-8\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | S/D | meas | m_infseg | r_SD | bilateral | 125100\12144-2\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | AI | meas | m_infseg | r_AccelIndex | bilateral | 125100\20167-3\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | AT | meas | m_infseg | r_AccelTime | bilateral | 125100\20168-1\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | DT | meas | m_infseg | r_DecelTime | bilateral | 125100\20217-6\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | PG | meas | m_infseg | r_PG | bilateral | 125100\20247-3\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | MG | meas | m_infseg | r_PGmean | bilateral | 125100\20256-4\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | TAMV | meas | m_infseg | r_TAMV | bilateral | 125100\20352-1\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | VTI | meas | m_infseg | r_VTI | bilateral | 125100\20354-7\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | HR | meas | m_infseg | r_HeartRate | bilateral | 125100\8867-4\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Seg | Dist | meas | m_infseg | r_Distance | bilateral | 125100\G-0364\G-A115\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | ? | meas | m_supseg | r_DopAngle | bilateral | 125100\C12122-01\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | EDV | meas | m_supseg | r_EDV | bilateral | 125100\11653-3\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | MDV | meas | m_supseg | r_MDV | bilateral | 125100\11665-7\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | TAPV | meas | m_supseg | r_TAPV | bilateral | 125100\11692-1\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | PSV | meas | m_supseg | r_PSV | bilateral | 125100\11726-7\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | PI | meas | m_supseg | r_PI | bilateral | 125100\12008-9\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | RI | meas | m_supseg | r_RI | bilateral | 125100\12023-8\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | S/D | meas | m_supseg | r_SD | bilateral | 125100\12144-2\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | AI | meas | m_supseg | r_AccelIndex | bilateral | 125100\20167-3\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | AT | meas | m_supseg | r_AccelTime | bilateral | 125100\20168-1\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | DT | meas | m_supseg | r_DecelTime | bilateral | 125100\20217-6\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | PG | meas | m_supseg | r_PG | bilateral | 125100\20247-3\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | MG | meas | m_supseg | r_PGmean | bilateral | 125100\20256-4\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | TAMV | meas | m_supseg | r_TAMV | bilateral | 125100\20352-1\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | VTI | meas | m_supseg | r_VTI | bilateral | 125100\20354-7\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | HR | meas | m_supseg | r_HeartRate | bilateral | 125100\8867-4\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Seg | Dist | meas | m_supseg | r_Distance | bilateral | 125100\G-0364\G-A116\T-46659\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | ? | meas | m_medarc | r_DopAngle | bilateral | 125100\C12122-01\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | EDV | meas | m_medarc | r_EDV | bilateral | 125100\11653-3\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | MDV | meas | m_medarc | r_MDV | bilateral | 125100\11665-7\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | TAPV | meas | m_medarc | r_TAPV | bilateral | 125100\11692-1\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | PSV | meas | m_medarc | r_PSV | bilateral | 125100\11726-7\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | PI | meas | m_medarc | r_PI | bilateral | 125100\12008-9\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | RI | meas | m_medarc | r_RI | bilateral | 125100\12023-8\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | S/D | meas | m_medarc | r_SD | bilateral | 125100\12144-2\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | AI | meas | m_medarc | r_AccelIndex | bilateral | 125100\20167-3\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | AT | meas | m_medarc | r_AccelTime | bilateral | 125100\20168-1\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | DT | meas | m_medarc | r_DecelTime | bilateral | 125100\20217-6\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | PG | meas | m_medarc | r_PG | bilateral | 125100\20247-3\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | MG | meas | m_medarc | r_PGmean | bilateral | 125100\20256-4\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | TAMV | meas | m_medarc | r_TAMV | bilateral | 125100\20352-1\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | VTI | meas | m_medarc | r_VTI | bilateral | 125100\20354-7\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | HR | meas | m_medarc | r_HeartRate | bilateral | 125100\8867-4\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Mid Arc | Dist | meas | m_medarc | r_Distance | bilateral | 125100\G-0364\G-A109\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | ? | meas | m_infarc | r_DopAngle | bilateral | 125100\C12122-01\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | EDV | meas | m_infarc | r_EDV | bilateral | 125100\11653-3\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | MDV | meas | m_infarc | r_MDV | bilateral | 125100\11665-7\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | TAPV | meas | m_infarc | r_TAPV | bilateral | 125100\11692-1\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | PSV | meas | m_infarc | r_PSV | bilateral | 125100\11726-7\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | PI | meas | m_infarc | r_PI | bilateral | 125100\12008-9\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | RI | meas | m_infarc | r_RI | bilateral | 125100\12023-8\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | S/D | meas | m_infarc | r_SD | bilateral | 125100\12144-2\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | AI | meas | m_infarc | r_AccelIndex | bilateral | 125100\20167-3\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | AT | meas | m_infarc | r_AccelTime | bilateral | 125100\20168-1\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | DT | meas | m_infarc | r_DecelTime | bilateral | 125100\20217-6\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | PG | meas | m_infarc | r_PG | bilateral | 125100\20247-3\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | MG | meas | m_infarc | r_PGmean | bilateral | 125100\20256-4\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | TAMV | meas | m_infarc | r_TAMV | bilateral | 125100\20352-1\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | VTI | meas | m_infarc | r_VTI | bilateral | 125100\20354-7\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | HR | meas | m_infarc | r_HeartRate | bilateral | 125100\8867-4\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Inf Arc | Dist | meas | m_infarc | r_Distance | bilateral | 125100\G-0364\G-A115\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | ? | meas | m_suparc | r_DopAngle | bilateral | 125100\C12122-01\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | EDV | meas | m_suparc | r_EDV | bilateral | 125100\11653-3\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | MDV | meas | m_suparc | r_MDV | bilateral | 125100\11665-7\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | TAPV | meas | m_suparc | r_TAPV | bilateral | 125100\11692-1\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | PSV | meas | m_suparc | r_PSV | bilateral | 125100\11726-7\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | PI | meas | m_suparc | r_PI | bilateral | 125100\12008-9\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | RI | meas | m_suparc | r_RI | bilateral | 125100\12023-8\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | S/D | meas | m_suparc | r_SD | bilateral | 125100\12144-2\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | AI | meas | m_suparc | r_AccelIndex | bilateral | 125100\20167-3\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | AT | meas | m_suparc | r_AccelTime | bilateral | 125100\20168-1\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | DT | meas | m_suparc | r_DecelTime | bilateral | 125100\20217-6\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | PG | meas | m_suparc | r_PG | bilateral | 125100\20247-3\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | MG | meas | m_suparc | r_PGmean | bilateral | 125100\20256-4\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | TAMV | meas | m_suparc | r_TAMV | bilateral | 125100\20352-1\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | VTI | meas | m_suparc | r_VTI | bilateral | 125100\20354-7\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | HR | meas | m_suparc | r_HeartRate | bilateral | 125100\8867-4\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Sup Arc | Dist | meas | m_suparc | r_Distance | bilateral | 125100\G-0364\G-A116\T-4668A\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior
Vascular Ultrasound Procedure Report | Abdominal | Ren V | ? | meas | m_renvein | r_DopAngle | bilateral | 125100\C12122-01\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | 99PMSBLUS, C12122-01, Doppler Correction Angle |  | UCUM, deg, deg | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Ren V | EDV | meas | m_renvein | r_EDV | bilateral | 125100\11653-3\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Ren V | MDV | meas | m_renvein | r_MDV | bilateral | 125100\11665-7\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Ren V | TAPV | meas | m_renvein | r_TAPV | bilateral | 125100\11692-1\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Ren V | PSV | meas | m_renvein | r_PSV | bilateral | 125100\11726-7\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Ren V | PI | meas | m_renvein | r_PI | bilateral | 125100\12008-9\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Ren V | RI | meas | m_renvein | r_RI | bilateral | 125100\12023-8\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, no units | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Ren V | S/D | meas | m_renvein | r_SD | bilateral | 125100\12144-2\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Ren V | AI | meas | m_renvein | r_AccelIndex | bilateral | 125100\20167-3\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20167-3, Acceleration Index |  | UCUM, mm/s2, mm/s2 | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Ren V | AT | meas | m_renvein | r_AccelTime | bilateral | 125100\20168-1\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20168-1, Acceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Ren V | DT | meas | m_renvein | r_DecelTime | bilateral | 125100\20217-6\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20217-6, Deceleration Time |  | UCUM, ms, ms | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Ren V | PG | meas | m_renvein | r_PG | bilateral | 125100\20247-3\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20247-3, Peak Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Ren V | MG | meas | m_renvein | r_PGmean | bilateral | 125100\20256-4\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20256-4, Mean Gradient |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Ren V | TAMV | meas | m_renvein | r_TAMV | bilateral | 125100\20352-1\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Ren V | VTI | meas | m_renvein | r_VTI | bilateral | 125100\20354-7\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 20354-7, Velocity Time Integral |  | UCUM, mm, mm | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Ren V | HR | meas | m_renvein | r_HeartRate | bilateral | 125100\8867-4\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, bpm | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Ren V | Dist | meas | m_renvein | r_Distance | bilateral | 125100\G-0364\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, mm, mm | [7 empty] | SRT, T-48740, Renal Vein
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI Avg | none | calc | c_Kid_EQI_Avg_p | none | bilateral | 125100\ C99-5100-021 \T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-021 , Average of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI Std | none | calc | c_Kid_EQI_Std_p | none | bilateral | 125100\ C99-5100-022\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-022, Standard Deviation of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI Med | none | calc | c_Kid_EQI_Med_p | none | bilateral | 125100\ C99-5100-023\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-023, Median of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI IQR | none | calc | c_Kid_EQI_Iqr_p | none | bilateral | 125100\ C99-5100-024\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-024, Interquartile Range of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI IQR/Med | none | calc | c_Kid_EQI_Iqrmed_p | none | bilateral | 125100\ C99-5100-025\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-025, InterquartileRange over Median of ShearWave Elastography samples |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI Avg Vel | none | calc | c_Kid_EQI_Avg_v | none | bilateral | 125100\ C99-5100-026\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-026, Average of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI Std Vel | none | calc | c_Kid_EQI_Std_v | none | bilateral | 125100\ C99-5100-027 \T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-027 , Standard Deviation of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI Med Vel | none | calc | c_Kid_EQI_Med_v | none | bilateral | 125100\ C99-5100-028\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-028, Median of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI IQR Vel | none | calc | c_Kid_EQI_Iqr_v | none | bilateral | 125100\ C99-5100-029\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-029, Interquartile Range of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kidney EQI IQR/Med Vel | none | calc | c_Kid_EQI_Iqrmed_v | none | bilateral | 125100\ C99-5100-030 \T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | 99PMSBLUS, C99-5100-030 , Interquartile Range over Median of ShearWave Velocity samples |  | UCUM, %, Percent | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | Kid Volume | none | calc | c_kidvol | none | bilateral | 125100\121221\T5100-01\T-71000 | 99PMSBLUS, T5100-01, Anatomic Structures |  | DCM, 121221, Volume of ellipsoid |  | UCUM, mm3, mm3 | [7 empty] | SNM3, T-71000, Kidney
Vascular Ultrasound Procedure Report | Abdominal | PV Bladder Vol | none | calc | c_pvbladder_vol | none | unilateral | 125100\C12011-08\C4-01\T5100-02 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C12011-08, Post Void Bladder Volume |  | UCUM, mm3, mm3 | [7 empty] | 99PMSBLUS, C4-01, Post Void Bladder
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI Avg | none | calc | c_Liver_EQI_Avg_p | none | unilateral | 125100\ C99-5100-021 \T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-021 , Average of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI Std | none | calc | c_Liver_EQI_Std_p | none | unilateral | 125100\ C99-5100-022\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-022, Standard Deviation of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI Med | none | calc | c_Liver_EQI_Med_p | none | unilateral | 125100\ C99-5100-023\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-023, Median of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI IQR | none | calc | c_Liver_EQI_Iqr_p | none | unilateral | 125100\ C99-5100-024\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-024, Interquartile Range of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI IQR/Med | none | calc | c_Liver_EQI_Iqrmed_p | none | unilateral | 125100\ C99-5100-025\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-025, InterquartileRange over Median of ShearWave Elastography samples |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI Avg Vel | none | calc | c_Liver_EQI_Avg_v | none | unilateral | 125100\ C99-5100-026\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-026, Average of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI Std Vel | none | calc | c_Liver_EQI_Std_v | none | unilateral | 125100\ C99-5100-027 \T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-027 , Standard Deviation of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI Med Vel | none | calc | c_Liver_EQI_Med_v | none | unilateral | 125100\ C99-5100-028\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-028, Median of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI IQR Vel | none | calc | c_Liver_EQI_Iqr_v | none | unilateral | 125100\ C99-5100-029\T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-029, Interquartile Range of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Liver EQI IQR/Med Vel | none | calc | c_Liver_EQI_Iqrmed_v | none | unilateral | 125100\ C99-5100-030 \T5100-02\T-62000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-030 , Interquartile Range over Median of ShearWave Velocity samples |  | UCUM, %, Percent | [7 empty] | SNM3, T-62000, Liver
Vascular Ultrasound Procedure Report | Abdominal | Bladder Vol | none | calc | c_bladder_vol | none | unilateral | 125100\C12011-04\T5100-02\T-74000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C12011-04, Bladder Volume |  | UCUM, mm3, mm3 | [7 empty] | SNM3, T-74000, Bladder
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI Avg | none | calc | c_Spl_EQI_Avg_p | none | unilateral | 125100\ C99-5100-021 \T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-021 , Average of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI Std | none | calc | c_Spl_EQI_Std_p | none | unilateral | 125100\ C99-5100-022\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-022, Standard Deviation of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI Med | none | calc | c_Spl_EQI_Med_p | none | unilateral | 125100\ C99-5100-023\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-023, Median of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI IQR | none | calc | c_Spl_EQI_Iqr_p | none | unilateral | 125100\ C99-5100-024\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-024, Interquartile Range of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI IQR/Med | none | calc | c_Spl_EQI_Iqrmed_p | none | unilateral | 125100\ C99-5100-025\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-025, InterquartileRange over Median of ShearWave Elastography samples |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI Avg Vel | none | calc | c_Spl_EQI_Avg_v | none | unilateral | 125100\ C99-5100-026\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-026, Average of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI Std Vel | none | calc | c_Spl_EQI_Std_v | none | unilateral | 125100\ C99-5100-027 \T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-027 , Standard Deviation of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI Med Vel | none | calc | c_Spl_EQI_Med_v | none | unilateral | 125100\ C99-5100-028\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-028, Median of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI IQR Vel | none | calc | c_Spl_EQI_Iqr_v | none | unilateral | 125100\ C99-5100-029\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-029, Interquartile Range of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen EQI IQR/Med Vel | none | calc | c_Spl_EQI_Iqrmed_v | none | unilateral | 125100\ C99-5100-030 \T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-030 , Interquartile Range over Median of ShearWave Velocity samples |  | UCUM, %, Percent | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Spleen Vol | none | calc | c_splnvol | none | unilateral | 125100\121221\T5100-02\T-C3000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | DCM, 121221, Volume of ellipsoid |  | UCUM, mm3, mm3 | [7 empty] | SNM3, T-C3000, Spleen
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI Avg | none | calc | c_Pan_EQI_Avg_p | none | unilateral | 125100\ C99-5100-021 \T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-021 , Average of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI Std | none | calc | c_Pan_EQI_Std_p | none | unilateral | 125100\ C99-5100-022\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-022, Standard Deviation of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI Med | none | calc | c_Pan_EQI_Med_p | none | unilateral | 125100\ C99-5100-023\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-023, Median of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI IQR | none | calc | c_Pan_EQI_Iqr_p | none | unilateral | 125100\ C99-5100-024\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-024, Interquartile Range of ShearWave Elastography samples |  | UCUM, mmHg, mmHg | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI IQR/Med | none | calc | c_Pan_EQI_Iqrmed_p | none | unilateral | 125100\ C99-5100-025\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-025, InterquartileRange over Median of ShearWave Elastography samples |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI Avg Vel | none | calc | c_Pan_EQI_Avg_v | none | unilateral | 125100\ C99-5100-026\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-026, Average of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI Std Vel | none | calc | c_Pan_EQI_Std_v | none | unilateral | 125100\ C99-5100-027 \T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-027 , Standard Deviation of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI Med Vel | none | calc | c_Pan_EQI_Med_v | none | unilateral | 125100\ C99-5100-028\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-028, Median of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI IQR Vel | none | calc | c_Pan_EQI_Iqr_v | none | unilateral | 125100\ C99-5100-029\T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-029, Interquartile Range of ShearWave Elastography Velocity samples |  | UCUM, mm/s, mm/s | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | Abdominal | Pancreas EQI IQR/Med Vel | none | calc | c_Pan_EQI_Iqrmed_v | none | unilateral | 125100\ C99-5100-030 \T5100-02\T-65000 | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral) |  | 99PMSBLUS, C99-5100-030 , Interquartile Range over Median of ShearWave Velocity samples |  | UCUM, %, Percent | [7 empty] | SRT, T-65000, Pancreas
Vascular Ultrasound Procedure Report | TCD | MCA/ICA | none | calc | c_tcd_MCAICA | none | bilateral | 125100\C12105-03\T-40501 | SRT, T-40501, Blood Vessel of Head |  | 99PMSBLUS, C12105-03, MCA/Dist ICA Ratio |  | UCUM, 1, no units
Vascular Ultrasound Procedure Report | Vascular | ICA/CCA | none | calc | c_ICA_CCA_Ratio | none | bilateral | 125100\33868-1\T-45005 | SRT, T-45005, Artery of neck |  | LN, 33868-1, ICA/CCA velocity ratio |  | UCUM, 1, no units
Vascular Ultrasound Procedure Report | Abdominal | RAR | none | calc | c_RAR | none | bilateral | 125100\33869-9\T-48740\T-71019 | SRT, T-71019, Vascular Structure Of Kidney |  | LN, 33869-9, Renal Artery/Aorta velocity ratio |  | UCUM, 1, no units | [7 empty] | SRT, T-48740, Renal Vein